## matrix42 export data
Ticket ID |  | Type | Summary | Category | Priority | Reaction | Solutiontime | Status | User | Responsible | Location | Responsible Role | Owner | Created | Closed | Reaction Time Escalated | Solution Time Escalated | Reminder Date | Solution
TCK77425 | User | Service Request | WG: Bracco - 2026 Spare parts Prices | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-02-03 09:55 | 2026-02-03 11:01 | 0 | 0 |  | it's this code: [Image]
TCK77411 | User | Service Request | TR: Demande de prix + délai | Process Instrumentation > Sales Technical Support | Low |  |  | On Hold | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-02-03 09:05 |  | 0 | 0 | 2026-02-04 09:23
TCK77370 | User | Service Request | INFOS GAZ | Process Instrumentation > Sales Technical Support | Low |  |  | On Hold | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-02-02 15:32 |  | 0 | 0 | 2026-02-03 16:13
TCK77301 | User | Service Request | certificates of CE and PED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-02-02 08:26 | 2026-02-02 10:26 | 0 | 0 |  | the PED doesn't apply to thermometers as they are not under pressure equipment. and as the PED doesn't apply there is not much to put on a CE certificate, sorry.
TCK77258 | User | Service Request | DEFINITION PRODUIT YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-30 15:31 | 2026-01-30 15:58 | 0 | 0 |  | je t'aurais bien dis de proposer un PFMN sur séparateur mais on ne passera pas avec les connexions elec ATEX, l'IP et la précision. Apres je ne sais pas si on peut répondre avec de la tierce partie montée sur séparateur, pour l'instant il me semble pas qu'on l'ai fait a part pour de grands projets, sinon il faudra faire une demande de définition a la tierce partie choisie (YOKO ?) pour le transmetteur et valider le montage sur notre séparateur (D438 +OR)
TCK77257 | User | Service Request | TR: TR: Pressostat différentiel : Naval Group | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-30 15:12 | 2026-02-03 11:47 | 0 | 0 |  | Je ferme, je pense que ça va prendre un peu de temps pour qu'ils revoient déjà toutes nos réponses, on se redis quand il y aura un retour.
TCK77173 | User | Service Request | TED avec option 0597 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-29 15:32 | 2026-01-29 15:42 | 0 | 0 |  | je t'avais répondu, dans le dernier ticket (0597 127V EDF) c'est bon pour moi, on peut les faire, maintenant il faut que tu demande le code et le prix.
TCK77137 | User | Service Request | TR: Accessories for DADL03 (DN 32 DIN 11851)/ Lalita | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-29 10:21 | 2026-01-29 10:58 | 0 | 0 |  | j'ai dis a Ang qu'on ne faisait plus ces pièce, c'est du piping ça doit être fait par le tuyauteur.
TCK77136 | User | Service Request | FW: PFMN-64.BC3R.DAIL16.2.L23.01 (11167784) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-29 10:10 | 2026-01-29 11:03 | 0 | 0 |  | we don't have IDF (ISO 2853 the BS being obsolete). the closest we have for Dairy industry will be the SMS and DIN (DAEx and DADx)
TCK77085 | User | Service Request | plan TMH-AMS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-28 15:11 | 2026-01-29 09:32 | 1 | 0 |  | Et voici
TCK77044 | User | Service Request | MMD3-G21.E20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-28 09:40 | 2026-01-28 15:04 | 0 | 0 |  | yes you can technically go with this MMD3, it seems also the right thing for me, the creation of the dial might have a cost if not existing, so your options will be to include the creation in the price (on 30 units it'll not represent much), or offer it in bar. next possibility is to offer a MEX3G21E20
TCK77006 | User | Service Request | info sur les CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-27 14:50 | 2026-01-27 15:00 | 0 | 0 |  | c'est marrant c'est une impression type ticket qui doit être faite avec tous leurs produit après leur passage au banc. c'est une sorte de certificat de calibration light (pas de norme ni de numéro du calibrateur de reference. chez nous le plus proche sera un certificat de conformité mais merci de l'info je vais pouvoir pousser plus for pour cette automatisation.
TCK77003 | User | Service Request | WG: DEMANDE DE PRIX N°: SACO26DP00890 / CCBL / DP 26CCB00450-25916 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-27 14:20 | 2026-01-27 15:04 | 0 | 0 |  | it's fine, just base it on MEX3 + DA in the P2F and then change MEX3 to MEX2 if everything is fine and ask the price to nathalie or build it separately (MEX2 on one hand and DA on the other + moutning price). you can also check the compatibility of your mounting in the seal navigator.
TCK76985 | User | Service Request | TR: Lead 30299126 ¦ CAF ¦ Opport 10106652 ¦ RFP : Pressure Gauges for CIVITY EMU ONCF (Morocco) Project | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-27 11:58 | 2026-01-29 12:12 | 0 | 0 |  | Premières actions faites, envoi des 3D, revue avec Dioni et le client pour clarif avant d'aller plus loin.
TCK76950 | User | Service Request | 1/4" API Thread | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-27 09:56 | 2026-01-27 11:37 | 0 | 0 |  | I'm surprised, the API is specific to only high stressed piping joint for underground / underwater application for the O&G industry. Also API is just the American Petroleum institute, and in their specific thread profiles, they based it on different norms, so NPT API could probably be compatible with standard NPT in the limits of usage of the standard NPT. I'm afraid it's a typo or an information coming from someone non specialized.
TCK76894 | User | Service Request | PFMN sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-26 14:29 | 2026-01-26 14:44 | 0 | 0 |  | tu peux utiliser le seal navigator pour vérifier ça il fait tous les calculs. [Image]
TCK76888 | User | Service Request | BASG / 10297002 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-26 13:51 | 2026-01-26 13:59 | 0 | 0 |  | je n'ai rien du tout dans le code, le mieux serait de retrouver une commande mais pour ça tu auras plus daccès que moi (je n'en trouve aucune). sinon, il faudra redéfinir avec le client.
TCK76875 | User | Service Request | Request for SEAL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-26 11:37 | 2026-01-26 11:45 | 0 | 0 |  | yes, I would offer something like this (see below). the cooling accessory could be integrated in the sensor (I think if you select a PBMN or PFMN you can have it directly on the sensor). [Image]
TCK76864 | User | Service Request | données de sécurité du TED6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-26 11:19 | 2026-01-26 11:39 | 0 | 0 |  | voici le MTTF(d) du TED
TCK76848 | User | Service Request | Gauge Identification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-26 10:01 | 2026-01-26 10:12 | 0 | 0 |  | from the Label, nothing oints toward a Bourdon product, the codes and product name are not Bourdon. can they access the QR code or send a picture of the dial ? otherwise you can offer to replace it with a MEX3.
TCK76752 | User | Service Request | 转发: 差异-9000019679-SO12-107882-ZPN201SEM/6/2237/SETP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-23 08:22 | 2026-01-23 09:40 | 0 | 0 |  | here is the detail of the code in the picture (label on the box and matching code visible on the scale plate (right picture): 11187285 ZPN 201SEM6/2237_SETP [Image] Option 1930 was the previous way of coding the 1/2NPT option before it was integrated into the main codification string as a 6.
TCK76663 | User | Service Request | version de TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 17:06 | 2026-01-22 09:25 | 1 | 0 |  | désolé, c'est une commande 2240, je n'ai pas les autorisations pour l'ouvrir mais l'ADV pourra.
TCK76658 | User | Service Request | WG: Her Meuli vom Paul Scherrer Institut AG erbittet Rückruf unter der Nummer 056/310 30 36 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 16:01 | 2026-01-21 16:40 | 0 | 0 |  | absolute measurement means that the reference is the vacuum, then we use a differential and actually vacuum one of the bellow, this is real absolute gauge (MA7) then you have the coompound gauges (standard gauge having part of the scale in vacuum) it's measuring vacuum but is not absolute, the 0 is still atmospheric pressure. and finally you have my least favorite, the pseudo absolute, you take a compount gauge and print 0 instead of -1 and simulate you have an absolute. I don't advise to go this way and we have to mark on the gauge "pseudo-absolute" to indicate it's not really measured against vaccum.
TCK76653 | User | Service Request | AROE- RTREGIONDUS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 15:12 | 2026-01-21 16:35 | 0 | 0 |  | c'est les mêmes 2 joints ci-dessous sur tes 2 ref: 10302601 JT CUIR 14X6,5X3 2 PC
TCK76643 | User | Service Request | WG: Anfrage Offerte | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 14:00 | 2026-01-21 14:25 | 0 | 0 |  | no problem, I miss most information but it seems like a simple bimetal on a welded thermowell like those: https://www.bourdon-instruments.com/fr/en/product-overview/accessories-temperature/thermowell-with-welded-connection/c/37884/products or those but it might be a bit short: https://www.bourdon-instruments.com/fr/en/product-overview/accessories-temperature/thermowell-with-welding-sleeve/c/37885/products
TCK76631 | User | Service Request | DROE100/211-133 /653 EK8001/D0A91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 12:57 | 2026-01-21 14:08 | 0 | 0 |  | this is th answer I sent to roberto Spricigo for the same request yesterday. even with the datasheet there is things that are not clear but I still attached it in the ticket. for the replacement I would offer as below but I still don't know what 653 (guessed 0..6bar from text) EK (I just took the 8001 switch) and DOA91 (no idea) means. Order code: MG5-2433.B20 MG old Code / SAP code:
TCK76628 | User | Service Request | ETTN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 11:44 | 2026-01-21 12:00 | 0 | 0 |  | pas de soucis c'est du standard.
TCK76589 | User | Service Request | replacement ¦ E913039B35R ¦ Proszę o ofertę | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 07:45 | 2026-01-21 09:31 | 0 | 0 |  | sure, we kept only the ATEX version of this sensor but it's still available, take a Y91339B35R
TCK76583 | User | Service Request | Inquiry on MEX5 Thread Options and Connection Accessories (260121) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-21 02:31 | 2026-01-21 09:29 | 0 | 0 |  | I guess something is not very clear in your request, we have NPT as standard in the catalogue (MEX5D60H21 for example, the "6" in the code is for 1/2" NPT). normally BSPT is not in use anymore and just PT doesn't makes much sense to me (it could be a norm I don't know but I doubt it). concerning the accessories and diaphragm seals, we have tons of it and for most usual connections in all industries, but I have no idea what a "1.5s ferrule" could be, sorry.
TCK76573 | User | Service Request | TVE - naval group- Pressostat différentiel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-20 16:52 | 2026-01-20 17:41 | 0 | 0 |  | Si ils monte les arrivées BP et HP a l'envers ce n'est pas problématique en matière de soufflets, ils sont robustes et résisteront 20 bar d'un coté et 16 de l'autre (max des specs selon la datasheet) a l'envers, par contre ça ne switchera pas du tout car ça éloignera le switch du contact. Pour la phrase en rouge ça vient de la DESP qui définie les accessoires de sécurité spécifiquement et nous ne sommes pas dans ce cas.
TCK76549 | User | Service Request | WG: Francare - Quotation - HAENNI Instruments AG - N° 202601292 - GMA14352 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-20 13:55 | 2026-01-20 17:51 | 1 | 0 |  | even with the datasheet there is things that are not clear but I still attached it in the ticket. for the replacement I would offer as below but I still don't know what 653 (guessed 0..6bar from text) EK (I just took the 8001 switch) and DOA91 (no idea) means. Order code: MG5-2433.B20 MG old Code / SAP code:
TCK76525 | User | Service Request | TED spécifique pour EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-20 10:35 | 2026-01-20 11:55 | 0 | 0 |  | coup de chance on a trouvé, c'est l'option 0597 (alim 127V) qui est utilisée par EDF et on a des cartes. ci-dessous des exemples de codes existants, mais comme tous les TED, il faudra demander l'ouverture du code, le prix etc...
TCK76507 | User | Service Request | TR: Offre Baumer Capillaire/suite visite de ce jour | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-20 09:49 | 2026-01-20 10:09 | 0 | 0 |  | non ça va être problématique a cette température car même si il résiste, toutes ses propriété mécanique notamment la torsion (qui est intéressante dans un capillaire) seront changées drastiquement. ils faut qu'ils fassent du rigide à mon avis.
TCK76451 | User | Service Request | IP protect class 转发: Fw:3100228447差异 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 13:03 | 2026-01-19 14:25 | 0 | 0 |  | I'm sorry but the result of our testings for IP ratings are for internal use only and cannot be shared. but know that it's an EN certification which means that it's verified during stndard ISO 9001/ EN 9100 yearly inspections and renewal.
TCK76437 | User | Service Request | BAIT / customer SAET S.p.A./ I: RICHIESTA OFFERTA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 11:38 | 2026-01-19 12:10 | 0 | 0 |  | the vales they choose will limit us as the connection is already defined (1/4NPT). for Hydrogen applications we reccomend metal metal tighness (double ferrule tube compression or cone and thread connections). but here it's not possible unless they change th valves. at 10bar max and ambiant temperature, 1/4NPT connection is mounted right and with adapted sealing (I hear they offered loctite specific for H2 mounting). it should be ok. except for the connection, the hydrogen will then be in contact with a thick tube and connection, it's all welded and all 316L and has been tested up to 200bar for 3 month on pure hydrogen with no impact (we tested it completely after and made micrographies of the tubes to se if there was traces of permeation or hydrogen stress induced cracks).
TCK76436 | User | Service Request | Liquid CO2 Gauges. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 11:35 | 2026-01-19 12:01 | 0 | 0 |  | CO2 liquid is usual, and the name is R134A to get the special refrigeration dial. but we should still respect the gauge limits and the min is -40C for the process fluid if the gauge is dry. in terms of block and bleed manifold, the AMFD02N will have the same limits.
TCK76435 | User | Service Request | MEX3's Rolled Cases | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 11:32 | 2026-01-19 11:58 | 0 | 0 |  | yes, it's all old codes from 2015, please prefer offering: MEX3-D20.B20/0133 6 pcs MEX3-D20.B26/0133 2 pcs MEX3-D20.B59/0133 2 pcs MEX3-D20.B22/0133
TCK76426 | User | Service Request | MIX7-D30.B20/2155 + D824-10L.42CK.B64.04 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 10:27 | 2026-01-19 11:46 | 0 | 0 |  | it makes sense now, HF fluid combines properties of strong acids, H+ ionisations that prevent from using some materials and coatings and complete incompatibility with glycerine (react with organic polyhydroxyl compound) and destroy the silicons atom chains. the only solution is a fluid saturated with fluor because then the fluor doesn't react with it. then the mounting makes sense, I would just change the pseudo absolute with a compound one (0...10bar abs (B22/2155) becomes -1...9bar R (B79)) it's more usual now to do it like this.
TCK76423 | User | Service Request | TR: Rif. 436617 - Richiesta Offerta BOURDON 15/01 1341 16261 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 10:17 | 2026-01-19 11:03 | 0 | 0 |  | j'en tire pas grand chose de plus que toi je pense, le premier code n'est pas correct, je ne trouve rien de correspondant à 420. mais en se basant sur le texte, je proposerais un: DRO 100 311 111 545 (R134A) pour le 2eme, j'ai trouvé: [Image]
TCK76421 | User | Service Request | WG: 2453619/ WG: 1007 SwissRe Lake ¦ Bestellung Tauchhülse | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 10:12 | 2026-01-19 10:53 | 0 | 0 |  | V4A is a stainless steel (similar to 316L) for thermowells, we can offer the small (T0x series for HVAC, general industry), the AGW (industry, welded) or the AGF (Heavy duty applications, bored) then you chose your length, diameters and connection to adapt it to your thermometer or specifications.
TCK76419 | User | Service Request | replacement ¦ TE1 SSSS1042000*0 ¦ INQUIRY | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 10:01 | 2026-01-19 10:30 | 0 | 0 |  | you better ask Denmark for it, I think BADK made it and I don't find the order number and all the SSS digits means "customer specific" I attach the datasheet for reference and the replacement will probably also be from sensor Denmark, maybe a TFRN.
TCK76404 | User | Service Request | MEP5(R)-462.F22.AORP0382/SETP.ARIN2606.D820-10N.2237.R20.09/0693_0678_0763_9941_9030 / LSP | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 08:05 | 2026-01-19 09:32 | 0 | 0 |  | we do not advise to do this because first of all we cannot confirm the mounting blocks at 13bar really (we can read it on a reference gauge in production but not on the mounting itself as it's only a 10bar). Also the gauge will already be in safety befor it blocks, at 12bar the movement will be protected mecanichally and so you force a bit on the gauge despite having a protection. the good practice is to protect the mounting starting at the end of scale (10bar).
TCK76403 | User | Service Request | [BAKR] Bourdon product inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lee, Jun Hyeong [ljun] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 08:01 | 2026-01-19 09:48 | 0 | 0 |  | to replace this as it seems they want personnalization, I suggest the MEX3. it's a bit more high end than what they currently have but we have better quantity and pricing so the price gap is not that high and a lot more option to cover their needs (especially for 2 units). example: Order code: MEX3-G51.B22/9695_9696 MEX old Code / SAP code:
TCK76398 | User | Service Request | MEP5-AORP-ARIN-D820 / Vietnam Project LSP | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-19 07:10 | 2026-01-19 09:27 | 0 | 0 |  | we didn't made test with the AORP (as it's not a temperature accessory). but we could based our reflection on a similar case, in terms of length, diameter and material, we could compare the AORP to 1/2 and ASIPD and we tested the temperature in the middle of an ASIPD. here are the results and as you can see at 200C process (blue) we have around 75C in the middle of the ASIPD (similar to at the gauge connection in your case) with 20C ambiant temperature at the gauge. I think it is conform to what your customer expect. [Image]
TCK76390 | User | Service Request | TR: AO 6000073873 | Service Desk > Non IT | Low |  |  | Closed | Pétré, Sandra [sape] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Mosberger, Marco [mmos] | 2026-01-16 16:31 | 2026-01-16 16:44 | 0 | 0 |  | c'est Stéphane qui pourra te dire ça, je ne prends pas de décisions concernant les prix et la vente.
TCK76385 | User | Service Request | replacement ¦ AL-H-2-4BM-0/6BAR-RZ-PY-PW-GH ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 16:00 | 2026-01-16 16:41 | 0 | 0 |  | the gauge below is a Baumer India reference: AL-H-2-4BM-0/6BAR-RZ-PY-PW-GH You can replace it with this similar reference from Bourdon: MIX7D31B20/0751/Q008/Q003
TCK76381 | User | Service Request | WG: Capteur pression TED5 société SFMCP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 14:50 | 2026-01-16 15:32 | 0 | 0 |  | yes, the TED5 has 2 galvanically isolated setpoint which means it takes already 4 wires. so it is an 8 pins M12 we use, you can find it on the datasheet with the exact wiring.
TCK76374 | User | Service Request | RE: IC Price Required MIX7-D30.B20/2155 + D824-10L.42CK.B64.04 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 14:04 | 2026-01-16 14:48 | 0 | 0 |  | Laisse tomber pour l'instant, la demande est bizarre j'ai demandé a Daniel de regarder le besoin de plus près.
TCK76369 | User | Service Request | RE: MEP8-430.B31.D415-2L20.4JE.B20.01/0078 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 13:53 | 2026-01-16 14:11 | 0 | 0 |  | yes pas de soucis : [Image]
TCK76347 | User | Service Request | MIX7 Pipe Mount option "P" | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 11:20 | 2026-01-21 10:08 | 0 | 0 |  | please find the step file enclosed.
TCK76325 | User | Service Request | TR: Product confirmation for diaphragm seals for pulp n paper industry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-16 09:09 | 2026-01-16 11:50 | 1 | 0 |  | étant donné qu'il n'y a pas de norme de référence, assurer la compatibilité n'est pas facile. pour moi il n'y a pas de différences autre que dans les tolérances potentiellement et 1mm sur une cote non critique. leur flasque est usiné, le notre est pressé (pas de problèmes à ces pressions).
TCK76299 | User | Service Request | MPG6 3D MODEL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-15 16:28 | 2026-01-15 17:18 | 0 | 0 |  | I added the 3D to the ticket.
TCK76196 | User | Service Request | DN25 PN100 Flush Diaphragm Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-14 15:41 | 2026-01-14 18:03 | 0 | 0 |  | No flush solution on my side, I would probably offer either the D415 or the D820 in another DN (most standard if they want not to be bothered anymore and get the best prices is DN50). but I think they will have trouble finding this mounting flush, they might accept the D415 option if their installation is already done.
TCK76171 | User | Service Request | Y91 avec vis frein | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-14 11:35 | 2026-01-14 12:07 | 0 | 0 |  | oui pas de problème mais ce sera une nouvelle ouverture de code, celui que tu as est ancien et la gamme contient les anciennes cellules. le nouveau sera fait avec les configurateurs.
TCK76102 | User | Service Request | TVE-SNEF-SIR SIT 20654990 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 15:00 | 2026-01-14 14:39 | 1 | 0 |  | il faudra être très prudent avec ces recommandations car je ne trouve que des informations parcellaires et sur des produits approchants mais sans forcement de concentration. Phosphate trisodique (solution aqueuse 0,1%) informations trouvées sur des utilisations dans le food and bev, la pharma et en appli FDA. base forte mais vu la dilution, le 316L me parait envisageable. si trop de doutes, voir pour du plastique. Hydrazine (solution aqueuse 5%) pas beaucoup d'infos, un fournisseur de matériaux mets le 316 en résistance limité mais sans définition de la concentration (solution saturée ?). du 316L en concentration limité resistera probablement mieux mais c'est dur de seprononcer. le plastique reste une solution. [Image] Ammoniac Normalement pas de problème c'est une application répandue et je trouve de multiples tables qui le mettent comme compatible. Ethanolamine 50% Ca semble ok à priori sur les tables que j'ai trouvé: [Image] Mix, on ne se prononcera pas, on a pas du tout les infos sur les interactions.
TCK76084 | User | Service Request | D030HL02 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 12:26 | 2026-01-13 13:56 | 0 | 0 |  | it's still available, we have to work on the supply to improve our prices but for now the codes are open in SAP.
TCK76082 | User | Service Request | rick-antargaz | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 12:14 | 2026-01-13 13:52 | 0 | 0 |  | le seul truc qui me vient a l'esprit c'est l'UL 94 V0 pour les MPG6 ou le certificat feu fumée des produits ferro.
TCK76055 | User | Service Request | MMA4 D30 B15 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 09:51 | 2026-01-13 10:49 | 0 | 0 |  | It's hard to find a concrete date but I have some clues that can help us. the latest datasheet I found for the product was from 2000. then it doesn't appear in any Baumer document (so after 2005). you can set the phase out date to 2005, I'm sure it was part of baumer rationalisation. a MEX5 will be a pretty much identical (slightly better) replacement.
TCK76052 | User | Service Request | RE: RICK - BABE - EOC evergem - surfactants - offertevraag 4 manometers (interne ref: WO 4275493) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 09:41 | 2026-01-13 11:46 | 0 | 0 |  | le D055 n'est pas top, létanchéité est faite par un joint polymère ce sera le point faible. en plus un adaptateur est toujours un 2ème point faible. est ce que leurs M16 n'aurait pas un fond conique par hasard ? si c'est ça on a lidéal: [Image] Sinon je conseillerais de souder un adaptateur M16 ou chiffrer une connexion M16 (ci-dessous), ça ne fera qu'une seule faiblesse et de sécuriser au maximum la connexion avec un joint le plus adapté possible et de le changer régulièrement. [Image]
TCK76049 | User | Service Request | TR: spare part of E913 COPCO 2065 10R (10263139) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 09:17 | 2026-01-13 11:31 | 0 | 0 |  | I would offer the same in ATEX + the cable: Y91-3.2.3.B29.R/2164 + 10303915 (M12*5 (1 pin not used) 10m shielded cable).
TCK76048 | User | Service Request | TR: item inquiry : 11152433 / 2 VALVE MANIFOLD HAL6M02V | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-13 09:14 | 2026-01-14 14:59 | 1 | 0 |  | Pour moi ce n'est pas exactement le même. les produits ont basiquement la même fonction (manifold 2 voies format carré) mais celui référencé chez Sonepar sous cette référence est le manifold Schneider. sur la photo vous avez un manifold 2V de chez Baumer Inde (BTIN).
TCK76015 | User | Service Request | MX35-D20.A38/0679 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-12 17:00 | 2026-01-13 09:59 | 1 | 0 |  | it makes sense when it's a low pressure and you require to lower the sensitivity to match a diaphragm that is too small. but on a 400bar, it first cannot be fitted with a DANC, but also it makes no sense, take a MEX5, I'm sure it's just a mistake on their side or a misunderstanding.
TCK76012 | User | Service Request | MCX7 Blow out disk | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-12 16:55 | 2026-01-13 09:56 | 1 | 0 |  | yes, there is the same blow out as the MEX5 and also a overpressure vent on the capsule.
TCK76005 | User | Service Request | Difference between option 0794, 0831 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Budzisz, Kamil [buk] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-12 16:00 | 2026-01-13 09:55 | 1 | 0 |  | so here are all the answers about the paint cleaning you should need. MIT and MAT low cost gauges are retail gauges from China and we don't master the cleaning process for those. I would suggest to go for the MEX5 (which seems also like a good industrial standard for when you handle clean processes). the difference between 0794 and 0831 is only a marking difference (blue: "paint label" on 0831 and black "SILIKONFREI" with 0794) otherwise they're extremely similar, the cleanings are the same and the values for testing the cleaning level with the cleanospector at the en also. I suggest the 0794 on both TBI and MEX, it's more international. concerning a certificate, there is no standard for this but the Q002 which contains a specific verification of the properties of the product will state the cleaning was checked.
TCK75974 | User | Service Request | MAT5, TB100 ¦ Silicon free | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Budzisz, Kamil [buk] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-12 12:15 | 2026-01-13 09:47 | 1 | 0 |  | This Ticket has been closed as it was merged to ticket TCK76005 - Difference between option 0794, 0831 "#" created by .
TCK75965 | User | Service Request | TR: UL approval - pressure sensor CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-12 11:20 | 2026-01-13 08:54 | 1 | 0 |  | The certificate is valid and you can find it on the UL website, it's just on our side it was choosen to use the C01 option for now to reflect the UL. Best, Nicolas
TCK75918 | User | Service Request | PER DIVISION IS 0.05BAR? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-11 23:54 | 2026-01-13 08:52 | 1 | 0 |  | please find hereafter the dial you requested:
TCK75863 | User | Service Request | EDF-TVE-MDX VS MX Z811BA5Y | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-09 11:14 | 2026-01-09 11:43 | 0 | 0 |  | Salut Yannick, tembête pas on va pas faire de pièces de rechanges sur des produits qui ont 20 ans. si ils ont un problème sur un MDX remplacement complet de l'appareil par un MX7. pour un MX7 au besoin on regardera si il est opportun de proposer des composants pour la maintenance, décision vente. Best, Nicolas
TCK75858 | User | Service Request | WG: demande de renseignement | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-09 11:03 | 2026-01-09 11:38 | 0 | 0 |  | liquid nitrogen is not a problem in itself, 316L will be fine but the temperature might be the issue, check you are within the temperature limits of the datasheet.
TCK75820 | User | Service Request | MEP8-460.B38/0679_0693_0843 / Theerasak | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2026-01-09 07:21 | 2026-01-09 09:05 | 0 | 0 |  | it is valid and can be offered. but 0.5% is not a class recognised by the EN so it cannot be marked EN837. take the option 0.6% instead and it'll be EN.
TCK75793 | User | Service Request | BAIT / customer: Demaclonko | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-08 14:12 | 2026-01-08 14:56 | 0 | 0 |  | the most standard is the MX7, it'll normally fit the application. the MCD7 is for very low static (max250mbar).
TCK75778 | User | Service Request | ADNOC Approval | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-08 12:13 | 2026-01-08 17:08 | 1 | 0 |  | Defined over teams, most simple mounting no specific option requested from ADNOC usually and we are referenced but there is no official certificate for this. MEX5D60B31 + D415 DN25 PN100 316L
TCK75723 | User | Service Request | 23101636 / Anfrage Aquarius/ Nr. 202504096 / AG 20665233 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-07 17:16 | 2026-01-07 17:45 | 0 | 0 |  | sure, but it's this code: RPPN-7A6.206 10297573 the one you had was probably closed because it was created as a double.
TCK75682 | User | Service Request | Indian DH Tri-Clamp Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-07 11:20 | 2026-01-07 11:49 | 0 | 0 |  | yes, they use american norms, not our EN and even less old DIN ;) you can offer a DANC012L2A, it'll fit as replacement.
TCK75675 | User | Service Request | TBHA limits | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-07 10:45 | 2026-01-07 10:54 | 0 | 0 |  | If they want to source a bigger spring to attach it, I guess it's fine but they'll be responsible for the good contact of the back of the thermometer with the pipe.
TCK75652 | User | Service Request | Technical Clarification / Tien | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2026-01-07 08:38 | 2026-01-07 10:51 | 0 | 0 |  | our tube will resist industrial vacuum (we test it ourself down to at least 10-3mbar absolute) the weak point will be more the diaphragm but we test it at the same level (10-3mbar absolute). but if they might go in vacuum, I would still prefer they take a -1...9 instead of a 0...10, both will resist the vacuum, but in the -1...9, the pointer will have room to move in the vacuum section, in the 10 it'll go below the 0 and be blocked mechanically by the under/overpressure safety bock in the movement. MEP5(R)-462.F79.AORP0382/SETP.ARIN2606.D820-10N.2237.R20.09/0693_0678_0763_9941_9030 and in the seals navigator, you can see the resistance of your mounting to the process conditions depending on the Class and rating, fillings, gauge, accessories etc... (it takes everything in account and gives you the limits
TCK75624 | User | Service Request | BABE Prijsaanvraag 3500013368 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-06 16:15 | 2026-01-07 10:07 | 1 | 0 |  | Comme discuté, changement pour LRS9 et: si remplacement sur existant (la même ref). si nouveau avec fourniture de partie a souder, proposer plutôt le D295.
TCK75620 | User | Service Request | Création code option CTX : 0772 Vis frein collée | Service Desk > Non IT | Low |  |  | Closed | Le Barze, Loic [llo] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-06 15:55 | 2026-01-07 10:05 | 1 | 0 |  | je pense que ça arrive par erreur chez moi, je ne peux pas changer les configs, par contre pour l'harmonisation cette option existe déjà chez nous sur les mêmes vis et raccords et avec le même process de collage je suppose (Loctite) et son numéro d'option est 0783
TCK75605 | User | Service Request | AQMO-DYA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-06 13:58 | 2026-01-06 14:15 | 0 | 0 |  | je pense qu'il y avait un 9 en trop mais je n'ai rien trouve avec 22901400048169. Désolé
TCK75581 | User | Service Request | BABE - BCC2404 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-06 09:36 | 2026-01-06 10:04 | 0 | 0 |  | c'est le code d'un ancien sous ensemble 07...2m qu'ils ont mais pour moi c'est ce standard qu'ils veulent. 10262332
TCK75577 | User | Service Request | RFQ of M/s. Aarti Industries - Instruments for MPP Hydrogenation Project | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-06 09:12 | 2026-01-07 17:54 | 0 | 0 |  | I would offer this, with the information we have it's the best option to ensure function and resistance and the lowest setp pressure possible with a diaphragm seal mounting. RP2-YM.354/SETP + D820-10L.2236.R2009 (step: 2bar falling, max stat: 40bar, max peak: class 600 (around 60bar at 180C)) RP2-YM.344/SETP + D435-2L20.216.G2009 (step: 0.5bar falling, max stat: 10bar) RP2-YM.344/SETP + D435-2L20.216.G2009 (step: 0.5bar falling, max stat: 10bar) RP2-YM.344/SETP + D435-2L20.216.G2009 (step: 0.5bar falling, max stat: 10bar)
TCK75510 | User | Service Request | RE: [BAKR]_Claim of MEP3-420.B35/0841 from customer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2026-01-05 15:50 | 2026-01-05 16:21 | 0 | 0 |  | Handled internally with claim
TCK75345 | User | Service Request | EF6 parameter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-30 15:18 | 2026-01-05 10:03 | 1 | 0 |  | You know we can't make specific test for each customer conditions but luckily there is a better solution which is norm compliance. here what will match your customer request is our compliance to the EN 50155 about electrical material used on railway applications. it specify in depth the conditions, electrical caracteristics, tests etc... that our product must respect, here an extract of the table (we're not authorized to disclose the full norm) containing the default requirement (you can see the altitude criteria used coming from the EN50125) the temperature criterias and the electrical caracteristics. Having one complete norm allow us to follow a clear protocol for design and qualification. And the customer to rely on a complete standard that include all characteristics for the use in his field of activity.
TCK75317 | User | Service Request | [BAKR]_Claim of MEP3-420.B35/0841 from customer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Noh, Ji Hoon [noj] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-24 08:26 | 2026-01-05 09:47 | 1 | 1 |  | Claim process seems started, thanks for the information Nicolas
TCK75299 | User | Service Request | TR: MIX7-460.D26.DT1-2N90.0.960.01/0751_9030 (11264241) / GTS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-23 16:55 | 2025-12-23 17:21 | 0 | 0 |  | oui ca existe c'est l'option 1999:
TCK75264 | User | Service Request | Anfrage Nr. BOURDON HAENNI(Nr.9266086) Bitte diese Anfrage Nummer im Angebot geben | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-23 07:23 | 2026-01-05 15:42 | 0 | 1 |  | Closed until and answer from the customer if any.
TCK75261 | User | Service Request | TR: UL approval - pressure sensor CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-22 18:28 | 2026-01-07 14:01 | 0 | 1 |  | pour l'instant l'option UL n'est pas complètement définie, ils utilisent l'option C01 comme dans cet exemple: 11280478
TCK75233 | User | Service Request | WG: Frigotechnik / Offer 20666541 / BAT / DR0 80 / Artikelnr. 11270522 | Service Desk > General | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Gerber, Matthias [gmat] | 2025-12-22 13:07 | 2026-01-06 10:41 | 1 | 1 |  | the code creation request should be sent to FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org>. if it's just for the dial creation I think you can ask your sales contact here (Nathalie ?) you can also inform them of the answer and picture I gave you it might help make sure the design stays on track.
TCK75229 | User | Service Request | TR: Quotation request - ADE0062 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-22 12:12 | 2025-12-22 14:22 | 1 | 0 |  | Jsuis pas sur a 100% il y en a 2 qui peuvent coller mais le plus vendu c'est celui la: 10262332 Apres pour etre sur il faut ouvrir et regarder le cadran en dessous.
TCK75202 | User | Service Request | DT2 Seal option 1803 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-22 09:46 | 2025-12-22 14:11 | 1 | 0 |  | It was an internal choice offering a high temperature DT2 option in the case of a low pressure instead of a low pressure DT8. the gasket was the limiting point so concerning the temperature, DT2 and 8 have the same characteristics.
TCK75193 | User | Service Request | replecement ¦ 11181585 ¦ UT6-2NM ¦ Zapytania ofertowe KM/4158/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-22 06:47 | 2025-12-22 14:07 | 1 | 0 |  | the UT is a switch from baumer india, it can be replaced by an RP2NR645. we don't have the 1/4NPTM connection as standard, but I offered with 1/2NPT and they'll have to check if the microswitch characteristic fits their application.
TCK75141 | User | Service Request | 11731487 ADAPTATEUR | Service Desk > Non IT | Low |  |  | Closed | Fontenne, Céline [cfon] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-19 11:10 | 2025-12-19 15:32 | 1 | 0 |  | Error destinataire
TCK75137 | User | Service Request | TR: Demande de pièce - anneau de rinçage en 1"1/2 - RTJ - CLASSE 600 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-19 10:34 | 2025-12-19 16:15 | 1 | 0 |  | ci-joint le plan des anneaux en RTJ a utiliser pour l'instant, le 1"1/2 600 n'existe pas (seulement en 900) mais tu peux le demander en te basant quand meme sur ce plan.
TCK75127 | User | Service Request | TR: Demande de prix - Z811BA64 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-19 09:34 | 2025-12-19 15:36 | 1 | 0 |  | La voici: 10308966 STD BLACK POINTER 150MM S-G DIN Best, Nicolas
TCK75078 | User | Service Request | WG: Alesa Anfrage CHEQ-008380 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-18 14:19 | 2025-12-18 14:53 | 0 | 0 |  | it's a switch from united electric (datasheet enclosed). I would replace it with and RT2 but we don't have ranges that wide, you'll have to check with the customer to define the appropriate range if there is any.
TCK75071 | User | Service Request | ASME Approval | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-18 13:10 | 2025-12-18 17:13 | 0 | 0 |  | for gauges I suppose they mean the ANSI/ASME B40.1 which is the american definition of the content of the EN 837-1. but to me it's only a guideline and there is no real certification for it. our gauges if you take it in psi respects it but I don't think you can provide certification (same as for the EN837-1, we don't provide a certificate). if you order a certificate of calibration though on a psi or psi based scale, the ASME B40 will be referenced as a norm of reference.
TCK75070 | User | Service Request | MEX3 3D CAD Files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-18 12:21 | 2025-12-18 13:56 | 0 | 0 |  | here you go, you cannot yet access it but maybe soon with onebaumer. I asked for a lot more information to be shared with trained people in sales units it's in the hands of Mr Imhof but the blocking point is the templates of trainings for now.
TCK75041 | User | Service Request | TED sur séparateur. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-17 16:52 | 2025-12-18 14:45 | 1 | 0 |  | oui, ajoute le revetement PTFE (pulverisé, c'est dommage que tu ai du vide, l'autre aurait été mieux mais sinon ca parait la bonne solution. je cherchais plus simple pour que ce ne soit pas trop cher mais je ne vois pas.
TCK75039 | User | Service Request | TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-17 16:39 | 2025-12-17 17:00 | 0 | 0 |  | J'aimerais bien te dépanner mais la qualité insiste sur le fait de ne rien dépanner hors du processus de claim.
TCK74986 | User | Service Request | BarFlex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-17 09:43 | 2025-12-17 09:59 | 0 | 0 |  | oui, cétait un 10270443 EA4-AF3.B31R qui a été vendu à AREMECA à Vendôme (pour une balance de calibration surement).
TCK74970 | User | Service Request | MODEL: LSP 050.050.1 S/N: 31662 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lim, Janice [ljan] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-17 08:50 | 2025-12-17 09:37 | 0 | 0 |  | this was a swiss product and is replaced by other Baumer products you can find on this page: https://www.baumer.com/int/en/product-overview/process-sensors/level-sensors/potentiometric-level-sensors/c/35486
TCK74952 | User | Service Request | Submersible Pressure gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-17 05:38 | 2025-12-17 09:31 | 0 | 0 |  | sorry this was the MAH3 but it was completely stopped and is not replaced
TCK74925 | User | Service Request | FLEX.S112..... | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-16 15:50 | 2025-12-16 16:00 | 0 | 0 |  | you can still order it, 10309326 FLEX.S112 G1/2 FEM. 1M. Inside it's a simple tube, the diameter is small, if they have particules, I would suggest to add a small filter before it, it should solve the issue.
TCK74786 | User | Service Request | FW: Request for Micro Switch details used in PS & DPS by BHFR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-15 13:51 | 2025-12-15 14:34 | 0 | 0 |  | I would love to help but it is sensitive information you request, it's up to the product manager to validate we can share it. I think it'll be ok but you need to ask Jorg Strüfing for those information.
TCK74773 | User | Service Request | VS: Frese A/S - PT100 sensor for Oil & Gas applications | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-15 11:13 | 2025-12-15 14:23 | 1 | 0 |  | if you want a drilled bar stock TW it'll be the AGF (enclosed the datasheet) it'll be something like this but I don't have the diameters I let you select it: AGFGX217Nxx300P we can do the calculation, it's called wake frequency calculation, we will send you a table with the necessary values for the customer to fill and we'll validate the thermowell is suitable according to the normalised calculation I think you request it as a technical calculation or document. for the hydrostatic test I think you request it with a Q002 notifying "hydrostatic test" but this is to check with sales when you'll request the price.
TCK74719 | User | Service Request | Y91 avec réglage personnalisé | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-12 16:46 | 2025-12-12 17:49 | 0 | 0 |  | Oui il faut prendre l'option de réglage et demander un réglage spécifique, on en a quelques uns d'ouvert comme ça. mais je conseillerais quand même de plus partir sur quelque chose de fait pour être réglé comme un PBSN ou un PP.
TCK74716 | User | Service Request | MEX5 + D421 - Baumer Singapour | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-12 15:56 | 2025-12-12 16:23 | 0 | 0 |  | Oui mais le code pour l'electropolissage du mano c'est 0156 (c'est sur les produits elec que c'est 9060) donc: MEX5-D62.D77/0751_0760_0156 + D421-2N20.229.G20.09
TCK74715 | User | Service Request | WG: *** urgent *** RFQ PR 122605293 / Item HAGT442270P0002 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-12 15:55 | 2025-12-12 16:17 | 0 | 0 |  | surely not MTA as its capsule gauges for low pressure, it makes no sense here. it seems like low cost HP application but I don't see which case material is requested, I would say MAT2/3 or MIT3 (63mm only)
TCK74692 | User | Service Request | Montage BAFR / PBMN + AKPL + D820 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-12 11:33 | 2025-12-12 14:01 | 1 | 0 |  | oui pas de soucis, il y aura juste en plus l'option de précision dégradée (0.5%) et l'option de montage. sinon c'est ok:
TCK74685 | User | Service Request | Frigotechnik / Offer 20666541 / BAT / DR0 80 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-12 10:29 | 2025-12-12 16:42 | 1 | 0 |  | I have the red one in -1...24bar and R449a but the other one I had it only in blue in stock but the red version should look the same.
TCK74642 | User | Service Request | TR: 11166832 / PBSN-1.X.B15.R.A1.14.06.2.X.C01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 16:18 | 2025-12-11 16:46 | 0 | 0 |  | Email request sent to Mario, follow up via email
TCK74627 | User | Service Request | D918MD13V9606M52Q0M0EE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 14:19 | 2025-12-11 15:04 | 0 | 0 |  | le joint est en 316L, par contre attention le DN15 est hors THR.
TCK74613 | User | Service Request | ETTN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 11:53 | 2025-12-11 14:34 | 1 | 0 |  | oui je pense avoir trouvé le premier code dont ils parlent par contre dans le 2eme le digit 6 n'existe pas donc je ne sais pas ce qu'ils veulent. peut être que c'est pour 60mm et dans ce ça la ce serait le même code. ETTN61X3A0/X002 11157705 [Image]
TCK74578 | User | Service Request | RE: BAKR_Specification Compliance Verification Request(251211) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 04:21 | 2025-12-11 09:29 | 0 | 0 |  | Answered on the previous ticket.
TCK74577 | User | Service Request | BAKR_Specification Compliance Verification Request(251211) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 03:57 | 2025-12-11 09:28 | 0 | 0 |  | here is what I would offer, it checks everything except the IP (65 instead of 67) but it seems weird to as for IP67 for this type of application, 65 should be enough. if they really need it you have to go to the safety gauge MEP3 to get this IP but the price is not the same and it feels like an overkill to me. Order code: MEX3-651.K22 MEX old Code / SAP code:
TCK74576 | User | Service Request | BAKR_Inquiry Regarding Availability of Bourdon Pressure Gauge Models(251210) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-11 01:58 | 2025-12-11 09:36 | 1 | 0 |  | we dropped this design a couple years ago but we still have the same product (only the case design changed). here are the corresponding new references: Order code: MR5-6D60.F20 MR old Code / SAP code: Order code: MR5-2D60.F22 MR old Code / SAP code: Order code: MR5-7D60.F22 MR old Code / SAP code:
TCK74538 | User | Service Request | capteur PBMN monté sur DAEL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-10 12:25 | 2025-12-10 14:45 | 0 | 0 |  | pas de soucis tu peux y aller comme ca, le montage sera precis a 0.5% et pas de reglage en dessous de l'echelle ou il y aura de la perte de précision.
TCK74505 | User | Service Request | TR: Prijsaanvraag Duhaso | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-10 09:13 | 2025-12-11 14:58 | 1 | 0 |  | Philippe n'a pas de plan pour celui ci, on ne risquait pas de le trouver, si tu en as vraiment besoin je dois demander a Philippe de le créer. je te transmet déjà la fiche produit, ça suffira peut être pour un remplacement. FICHE PRODUIT - MANOMETRE FERROVIAIRE / PRODUCT FILE - RAILWAY PRESSURE GAUGE Internal codeSAP code (MFE)DiameterCircuit numberPressure (bar)AccuracyConnectionMounting studLeft locating pin grooveRight locating pin grooveLighttingPointer beneath dial (H+6)Case colourDial colourLettering colourRoundel colourSpecifics markingsFront pointer colourFront pointer black heelRear pointer colourRear pointer black heelTrim colourTrim screw colourInternal drawing nBH parts list drawing nCustomer drawingCustomer part list drawingHomologationRemarksTechnical documentDial screen printingFront guard screen printing ResolutionInternal user's manualCustomer article code
TCK74486 | User | Service Request | FW: [한국플루텍] 견적문의 - HAENNI / BAUMER (251208-E) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Noh, Ji Hoon [noj] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-10 02:17 | 2025-12-16 14:40 | 1 | 0 |  | I don't provide complete project definition. Because technical clarifications, price revisions, and condition adjustments (documentation, packaging, inspection, etc.) will then be necessary, and I wouldn't be able to manage it. The best practice is for the SU to define everything they can and ask me questions or request assistance on specific points as needed. The project manager must have defined the products, certificates, conditions, etc., to be able to manage the rest of the project. Still I read the specs and can tell you what to look at: Bourdon tube pressure gauges: MEP8.461.Bxx Capsule gauges (no filling in capsule gauges) MCX7 460 Bxx Accessories: Gauge union: ARIN2N06 (necessary ? already in 1/2NPT M) Manifold: AMFD02N (F/F) + ARIN2606 (M/M) Level gauge: we don't have Thermometers: TBHI (100mm or 130mm but fillable) TBI (160mm but not fillable) Thermowell: AGFR all the details are still to define but it should at least guide you.
TCK74451 | User | Service Request | étalonnage sur TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 14:36 | 2025-12-09 14:56 | 0 | 0 |  | il faut demander au ventes c'est une erreur dans la livraison des produits (je suis sur qu'il vont les retrouver). par contre je pense qu'il vaut mieux donner les numéros d'OF ou de commande, je ne sais pas si les numéros que tu donne sur les TED sont suivis.
TCK74445 | User | Service Request | MMD5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 13:23 | 2025-12-09 14:10 | 0 | 0 |  | first thing to check is if it's existing already, I have 2 similar products existing, (SPIRAX SARCO and EQUANS SWITZERLAND logos). if it's one of those you just use the existing code. if it's not, then you just make the quotation with the option for the logo and the price for it's creation (defined with sales). then at the order stage, there will be the creation of the Lxxx number and the validation of the Dial look with the customer.
TCK74442 | User | Service Request | Doc reçu | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 12:35 | 2025-12-10 17:26 | 0 | 0 |  | Doc sent via wetransfer after bug on other communications
TCK74429 | User | Service Request | TR: NCF-2661 // MEP5 Manomètre - Plage de la bande rouge inversée | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 10:43 | 2025-12-09 15:22 | 1 | 0 |  | Ticket clôturé en attente de retour client, pour la suite Yannick tu pourras surement m'envoyer sur ma boite mail standard comme ça va probablement demander des partages et de la com externe ce qui n'est pas pratique dans les tickets.
TCK74414 | User | Service Request | TR: [한국플루텍] 견적문의 - HAENNI / BAUMER (251208-E) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 09:49 | 2025-12-09 15:27 | 1 | 0 |  | je ne fait pas de définition complète de projet normalement car il faudra ensuite faire les technical clarifications, les révisions de prix et de conditions (documentation, packing, inspection etc...) et je ne pourrais pas tout gérer pour eux. la bonne pratique est qu'ils fassent la définition mais peuvent me poser des questions ou me demander de l'aide sur certains points au besoin, mais ils doivent pouvoir ensuite avoir défini les produits certificats, conditions etc... pour être capables de gérer la suite du projet.
TCK74401 | User | Service Request | WG: Artikel 10264786 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-09 08:33 | 2025-12-09 09:19 | 0 | 0 |  | please find enclosed the drawing and the related norms when we refer to G1/2: [Image]
TCK74355 | User | Service Request | WG: Konformitätserklärungen / Einbauerklärungen (zu BE2400503) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-08 11:37 | 2025-12-08 15:29 | 1 | 0 |  | this question was answered from Jorg to Pia a couple weeks ago, here is the answer (in german): Von: Strüfing Jörg Gesendet: Donnerstag, 20. November 2025 09:15 An: Bogner Pia <PBogner@baumer.com "mailto:PBogner@baumer.com"> Cc: Sutter Matthias <MSutter@baumer.com "mailto:MSutter@baumer.com"> Betreff: AW: Konformitätserklärungen / Einbauerklärungen (zu BE2400503) allo Pia, für die Schutzhülse und das Ventil gibt es keine EU Konformitätserklärungen, da sie unter keine EU Richtlinie fallen. Diese Produkte haben auch keine CE Kennzeichnung. Das ZPX kenne ich nicht, ist kein PI Produkt und steht im SAP auf Labor 100 (BECH).
TCK74351 | User | Service Request | replacement ¦ 11053093 ¦ PBCN 13B20RA1440721 ¦ Zapytanie ofertowe APE Group | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-08 11:13 | 2025-12-08 11:25 | 0 | 0 |  | the main caracteristic of the PBCN was the robustness of the cell, so the best replacement I see is the Y91 that is also extremely robust: [Image]
TCK74335 | User | Service Request | NIH - TMH-AMS- AORP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-08 09:58 | 2025-12-08 10:24 | 0 | 0 |  | oui c'est ça, ils peuvent les changer mais il n'y aura plus de garantie chez nous mais ça parait normal. il y a 2 joints par AORP tel que ci-dessous (AORP, j'ai pas chécké les AORPB): Joint 1: Dimensions: 2.6 x 1.9mm Type: joint torique Matiere: FPM (Fluorocarbone) Dureté: 80 shore A Joint 2: Dimensions: 7.2 x 1.9mm Type: joint torique Matiere: FPM (Fluorocarbone) Dureté: 70 shore A
TCK74321 | User | Service Request | 转发: Dooo🌻 和 张伟 的聊天记录 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-08 07:22 | 2025-12-08 10:03 | 1 | 0 |  | all parts in contact with fluid (all parts of the sensing element assembly (see picture) are in 316L. I checked the 4 material certificates for the subcomponents we currently have in stock and it's all 316L. the marks are very light even if it'll be up to quality dept to state if it's important. if the customer deems it necessary, you can start a claim process with your sales contact.
TCK74305 | User | Service Request | Compatibility MFT5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-05 22:51 | 2025-12-09 11:59 | 0 | 0 |  | we had the offer from fantinelli, sales will communicate a price to you. the static pressure will be limited to 100bar, we cannot go higher.
TCK74214 | User | Service Request | MIX7-D30.B20/2155 assembled with a media separator D824-10L.42CK.B64.04 ¦ Zapytanie ofertowe Baumer - UP/2457/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-05 07:11 | 2025-12-05 09:30 | 0 | 0 |  | you're at the technical support here, I can't give price or lead time. but I take this opportunityto raise a question concerning your code, the LRS4 (04 at the end of the diaphragm seal code) is a specific filling dedicated to Oxygen application but I don't seethe oxygen cleaning requested, is that normal ? for the price and lead time, please contact my sales colleagues.
TCK74202 | User | Service Request | GRCH-ETE-FLOEX-2504911 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-04 17:58 | 2025-12-05 09:24 | 0 | 0 |  | c'est une gateway LoRa mais ce n'est pas de chez nous: https://multitech.com/all-products/cellular/cellular-gateways/conduit-ip67-base-station/
TCK74179 | User | Service Request | TR: Westinghouse ¦ Indicador de temperatura con contactos y capilar para ASC A-35363-2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-04 14:13 | 2025-12-04 15:15 | 0 | 0 |  | on a pas de valeur exacte, on sait seulement que lhystérésis est faible et inférieure à la précision de l'appareil.
TCK74168 | User | Service Request | R: Silos Revolution - BAUMER: Non conformità TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Giannatempo, Franco [fgi] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-04 12:25 | 2025-12-04 15:17 | 1 | 0 |  | This is no a technical question, you need to open a claim process. your quality contact at BHFR will be Noel Robin nrobin@baumer.com. Best, Nicolas
TCK74167 | User | Service Request | I: Silos Revolution - BAUMER: Non conformità TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-04 12:14 | 2025-12-04 15:17 | 1 | 0 |  | This is no a technical question, you need to open a claim process. your quality contact at BHFR will be Noel Robin nrobin@baumer.com. Best, Nicolas
TCK74114 | User | Service Request | BAKR_Inquiry Regarding TBI Configuration Options(251204) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-04 05:07 | 2025-12-04 09:15 | 0 | 0 |  | we always provide the option already mounted to the product as it's a traceability option (the TAG plate identifies the product). it would be a specific request to have it separated.
TCK74093 | User | Service Request | Montage PFMN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-03 14:04 | 2025-12-03 15:21 | 0 | 0 |  | pas de soucis tu peux y aller: [Image]
TCK74072 | User | Service Request | WG: 4700267814 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-03 10:53 | 2025-12-03 11:56 | 0 | 0 |  | I found old similar codes, and they seem pretty standard, the first one below just has an option to be G1/4 [Image] the exact code they request aslo exists (10319761) and can be replace by a MEX3D21B26 and you can add the option 0751 if you want but it was not in the original one, in the DRO it was a plexiglass so you coul also choose option 0753
TCK74057 | User | Service Request | TR: MZ7-A60.X12G/0842_9712 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-03 09:22 | 2025-12-03 09:42 | 0 | 0 |  | Comme discuté, c'est bien équivalent en termes des valeurs (vu que 0.59bar = 59 kPa). par contre dans le descriptif ils demandent des kPa et le cadran est en bar donc les unités sont differentes.
TCK73955 | User | Service Request | BAIT / CTX - PP10R | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-02 08:59 | 2025-12-02 10:17 | 0 | 0 |  | when I hear liquid + pump I hear water hammer, it would be good to protect the CTX with a restrictor screw or a dampener I think. it doesn't come from the CTX, it was tested and approved for multiple demanding applications like railway, UL water applications, autoclaves etc... the PP10R is also very good but I don't think the problem resides in the product. example of a code with restrictor screw: CTX-3.3.9.B79.0/0771 11274763
TCK73954 | User | Service Request | TR: Demande de devis manomètre à cellule différentielle | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-02 08:48 | 2025-12-02 10:12 | 0 | 0 |  | Option d'offres MFT5/7 definies par teams.
TCK73933 | User | Service Request | BAKR_For inquiries regarding TBI160, 9941 options (251202) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-02 01:36 | 2025-12-02 09:27 | 0 | 0 |  | the option 9941 stainless wired tag plate can be added with any mounting or connection (as long as there is one to attach the tag plate). for the price you can maybe just take the price from the P2Finder or ask sales they can probably just give you the price of the tag plate and you add it to the TBI you need.
TCK73901 | User | Service Request | 1/4" Face Seal Female Swivel 1/4" | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 15:34 | 2025-12-01 17:24 | 0 | 0 |  | They might mean this type of connection, please check with them: [Image]
TCK73900 | User | Service Request | BAKR_Inquiry for Recommended Accessories for TBI160 (251201) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 15:31 | 2025-12-01 17:16 | 0 | 0 |  | the metallic wired TAG plate is the option 9941, you should have it in product finder, SAP codes or configurators.
TCK73898 | User | Service Request | Baumer India DH seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 15:27 | 2025-12-01 17:14 | 0 | 0 |  | sorry we don't have an ASME BPE model yet, it's an ongoing discussion, I would like to push also to get it and france wants it. but for now you can't answer on this, sorry.
TCK73873 | User | Service Request | module NAEV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 11:27 | 2025-12-01 17:11 | 1 | 0 |  | oui c'est une barrière a mettre en amont du YTED en dehors de la zone ATEX pour convertir la puissance en signal conforme pour la zone. on ne fait plus ces barrières ni le YTED désolé, et j'ai vérifié on en a plus en stock.
TCK73870 | User | Service Request | MEX8 Safety Glass and Gaskets | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 11:04 | 2025-12-01 12:09 | 0 | 0 |  | the question of if we can sell it or not as spares and pricing is on the sales, you'll have to ask them. but I can give you the part number at least: 10308635 VOYANT 160 TRIPLEX
TCK73868 | User | Service Request | MEX5-AX1.X82/9712_9704_DC99 (11252164)/ Alex | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2025-12-01 10:55 | 2025-12-01 12:06 | 0 | 0 |  | sure, here is the drawing of your connection: [Image]
TCK73848 | User | Service Request | Leitenberger Order 2381305 / 6050756 / drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 09:20 | 2025-12-01 09:56 | 0 | 0 |  | Sorry but same answer as for the other, KHW didn't had 3D drawings, but you can order it as a technical document.
TCK73847 | User | Service Request | 55269266 / Offer 20663457 / Breuer / delivery time for Y91 with seal 1545/ DN25 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 09:19 | 2025-12-01 09:53 | 0 | 0 |  | The 1545 is DN20 and it's the 1542 that is DN25 and this will dictate the connection between the gauge and the seal (G1/4 or G1/2) but this is not a choice for the customer, it'll have no impact on his side as he cannot touch this connection that ensure filling liquid tightness. here are the 2 options to offer: Y91-3.2.3.B79.R.2165/0592/9010 + 1545 (DN20) + LRS1 Y91-3.3.3.B79.R.2165/0592/9010 + 1542 (DN25) + LRS1 0592 (Coonection with diaphragm seal) and 9010 (accuracy of the mounting xx%) are mandatory.
TCK73846 | User | Service Request | 2411717 / Groz-Beckert / 4513483657 / Step File Articel 10320147 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 09:19 | 2025-12-01 09:43 | 0 | 0 |  | sorry we don't have 3D of haenni gauges, also it should probably be transformed into a MEX5 with a diaphragm seal tha exists in a norm/standard. you can still order it though it's not blocked but for getting the 3D drawing you must order it with the order as it's not existing yet.
TCK73844 | User | Service Request | WG: Automatische Antwort: Druckschalter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-12-01 08:59 | 2025-12-01 09:22 | 0 | 0 |  | we have 2 certifications for contact pressure gauge and they come from the same certification with this difference of C/A. you have the wrong label, here is an extract of the current certificate stating for MG5 it should be a C.
TCK73752 | User | Service Request | TR: BAIT / customer Fluid processing / I: misura temperatura esterna | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-28 08:57 | 2025-11-28 09:16 | 0 | 0 |  | C'est un montage de l'ETTN on peut encore en faire, tu peux leur envoyer la datasheet et on verra si il est possible de le rouvrir.
TCK73740 | User | Service Request | BARK_Inquiry Regarding Overrange Protection for Bourdon Pressure Gauges (251128) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-28 04:30 | 2025-11-28 09:13 | 0 | 0 |  | Sure, I can give you some details about overrange protection and I'll send you the training I do on option you can find more details in slides 8 to 11. first a gauge as standard is protected for OverPressure un to 130% of the full range (0...10bar gauge resists 13bar OP) then we have 2 options that add internal safety mecanisms in the tube and movement to increase this safety (options 0191 and 0192), be careful, as we don't respect the EN max alowable pressure anymore it's not maked or certified EN anymore. the best protection is an external overrange protector (AORP) that'll close below the gauge at a set pressure and resist up to 400bars.
TCK73713 | User | Service Request | TR: VDAV -- AVF - N/demande DP2511-2079 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-27 14:52 | 2025-12-04 16:59 | 0 | 0 |  | Nathalie, voici la réponse du PM, c'est ok, c'est leur responsabilité et attention a ne pas le faire au detriment de nos instruments. [Image] Best, Nicolas
TCK73711 | User | Service Request | replacement for 10301650 Y91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-27 14:27 | 2025-11-27 15:03 | 0 | 0 |  | sorry, the very low pressure on the Y91 were phased out the closest you can offer is the Y91 with N08 (100mbar which is the lowest range). Y91-3.3.9.N08.R/2164_2151 or a PBMN also in 100mbar range but you can configure it to 60mbar.
TCK73696 | User | Service Request | TR: CAF ¦ AMLD Project ¦ manometers screw | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-27 12:11 | 2025-11-27 14:13 | 0 | 0 |  | a priori ils n'ont pas le système de montage alstom de toute façon donc il faut juste proposer létrier de fixation qui est sur cette photo. [Image]
TCK73695 | User | Service Request | BarFlex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-27 11:56 | 2025-11-27 15:15 | 1 | 0 |  | oui il y a des choses qui ne vont pas, la connexion ne peut pas être console sur un barflex 0.1%, ça doit être Fixe (F) ou portable (P). et le N72 sera trop bas, notre minimum sera de 10bar déchelle. je te joints la datasheet.
TCK73668 | User | Service Request | WG: Demande mise à Jour offre EIC SA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-27 09:17 | 2025-11-27 14:08 | 1 | 0 |  | according to me you need the code below, check the details of the SETP but it checks your request. the status is declared as PN (phase out defined) but you can request the PM's approval to reopen it (normally it'll be OK and the we can unblock it and you can order. PM is Mario Marolf. 11258923 CPX-B.3.B22.R.1/SETP (SETP details in the SAP text code)
TCK73645 | User | Service Request | MAT5 ¦ Silicon free option | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Budzisz, Kamil [buk] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-26 15:42 | 2025-11-26 17:06 | 0 | 0 |  | it'll not be available on low cost gauges, I suggest doing an SPR on a MEX5 with the option (0831) and the qty to request the best price.
TCK73632 | User | Service Request | MEX3 + DAVA - regulation 1935/2004 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Budzisz, Kamil [buk] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-26 14:37 | 2025-11-26 15:11 | 0 | 0 |  | here is the declaration of conformity to the EC 1935/2004 for the DAxx and for the compatible filling liquids in your mounting (BH1, LRS1). Also more used internationnally is the 3A certification (you can check compatibility in the seals navigator), I enclosed our certificate also.
TCK73594 | User | Service Request | TR: inquiry BBE-25V-0186-1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-26 08:41 | 2025-11-26 09:05 | 0 | 0 |  | c'est normalement celui ci sur cette application 10262332 SDI 439/131.211 0,7...2M si il a un doute le client peut verifier sous le premier que c'est bien un 0.7...2m.
TCK73565 | User | Service Request | 55251850 / Anfrage Nr. S18573 / AG / contact gauge / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-25 16:14 | 2025-11-25 18:02 | 0 | 0 |  | this is a haenni type number and I think Jorg can trace it. On my side I see it could be an order from BAFR but I can't access it. here is how to understand it: 1102879: order number 01: position in the order 004: 4th product on the required qty.
TCK73564 | User | Service Request | 243436 /Bestellung / Kobold / Ihrer Ref. 0135997/2100104/3 / Cross Section from Seal 1627 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-25 16:07 | 2025-11-27 15:59 | 0 | 0 |  | I cannot share details about the moving parts, volumes etc... that's only internal confidential information. here is the drawing for the external dimensions: [Image]
TCK73555 | User | Service Request | BABE-SDI- Quotation request - SPN0271 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-25 14:43 | 2025-11-25 15:13 | 0 | 0 |  | je n'ai pas avec le Logo gregersen mais c'est le meme que le Axflow qu'on fourni a tout le monde, je lui proposerais celui ci en remplacement. 10262332 SDI 439/131.211 0,7...2M 0..1/4..1/2..3/4..1 CUBIC/CYLINDRIC CONNECTION WITH TAPERED RING SCREW CONNECTION FOR TUBE 8 MM AXFLOW LOGO missing in the details ? 0.7-2m NG8 Lans Nivotherm B.V. AUDIN S.A Baumer (Singapore) Pte. Ltd. MGA Controls Ltd Sinmo GmbH Berg & Larsen A/S Maersk Tankers A/S Louwers Mechanical Parts DFDS A/S Imatech BV [Image]
TCK73534 | User | Service Request | MFT7 PURGES | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-25 11:29 | 2025-11-25 11:51 | 0 | 0 |  | le 1/4 c'est la grande vis et sur ce qu'on achète nous il n'y a pas de petite vis, juste le bouchon 1/4NPT.
TCK73524 | User | Service Request | MFT7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-25 10:17 | 2025-11-25 11:07 | 0 | 0 |  | on ne pourra pas proposer ces nettoyages a cause du remplissage OM de MFT. tu peux proposer notre gamme MX si la statique est faible.
TCK73492 | User | Service Request | PG with Dashpot can combine with accuracy class 0.5 option / Thanh | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2025-11-25 05:53 | 2025-11-25 09:51 | 0 | 0 |  | yes it is possible but I'll still suggest to take the 0.6% accuracy option because it's conform to the EN, with the 0.5% you loose the EN marking and conformity.
TCK73479 | User | Service Request | AVF - N/demande DP2511-2019 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-24 16:48 | 2025-11-25 11:35 | 1 | 0 |  | je ne vois pas de problème majeur mais pour évaluer la résistance du doigt de gant il faut faire le calcul de stress. avec le fluide (eau de mer je peux estimer une viscosité et prendre une température standard (10-20C) mais il va surtout nous manquer le diamètre du pipe (et quelques autres données, je vais récupérer un tableau macweld et on regardera).
TCK73473 | User | Service Request | BAKR_Bourdon Product Inquiry (251124) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-24 15:55 | 2025-11-25 10:07 | 1 | 0 |  | Pressure gauges: Our gauge union are the ARIN but a 1/2NPT M/F doesn't make much sense as it means it help you connect a 1/2NPTM to a 1/2NPTF which you can do without the union. Our manifolds are the AMFD serie, you can take an AMFD02N +ARIN2606 to make it M/F different pointers will allow 0 adjustment like the 0678 for example (not advised with filling and can be limited in low pressure). Temperature gauges: the option for stainless steel tag plate is 9941. it seems they fit pretty well with your request but we can give you details on it once the product is defi
TCK73472 | User | Service Request | TR: CTX-3.B.3.B22.A/4780 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-24 15:52 | 2025-11-25 09:56 | 1 | 0 |  | 4780 is the signal inversion, and it only makes sense to use it when you have a -1...0 as -1 is the vacuum you'll reach by increasing the vacuum (it's a lower value but the inversion is useful to keep your 0 as the 4mA) it's not at all used for standard ranges.
TCK73403 | User | Service Request | BAES -- CAF ¦ AMLD Project ¦ manometers screw | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-24 09:17 | 2025-11-25 10:53 | 1 | 0 |  | the 3 screws in red are provided but it's not panel mounting screws, it just fix the front flange to the case. the fixing is done by the grooves on the case (what you see in the last red square). or with a fixing flange if they don't have this system (I sent you a picture)
TCK73307 | User | Service Request | INFO MFT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-21 11:02 | 2025-11-25 09:46 | 0 | 0 |  | ce sont les vis de purge et elles sont en 1/4NPT, il ne sont rien sensé connecter dessus mais peuvent les ouvrir lors d'un démontage pour nettoyer/purger les flasques.
TCK73295 | User | Service Request | BAIT / Fontanelli / TSFE 100 with Analog 4-20 mA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-21 09:57 | 2025-11-21 10:13 | 0 | 0 |  | It'll be a TSFx (Stiko) and so it's to be requested to BHDE. but it exists in their catalog (see below):
TCK73249 | User | Service Request | Faisabilité technique | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-20 15:04 | 2025-11-20 16:17 | 0 | 0 |  | le code SAP est ouvert, le client peut recommander sans problème. 11213492 MX35-D30.B76.XDSE:53910-01/LRS1_0751++
TCK73152 | User | Service Request | TR: 2823 - DDP - Trémie Cendres | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-19 14:25 | 2025-11-19 15:05 | 0 | 0 |  | solved over teams
TCK73049 | User | Service Request | Manomètre avec séparateur à bride mariage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-18 13:46 | 2025-11-18 15:23 | 0 | 0 |  | I think you overcomplicate things here, if I understand we you need the 2 mountings below: MEX5-D30.B77 + D821-10L.42DD.B2001 (11011237 + 11201827 + LRS1) MEX5-D30.B79 + D821-10L.42DD.B2001 (11011238 + 11201827 + LRS1) DN20 or 25 to confirm (you selected 20 in the datasheet in yellow but 25 in the seals navigator). all those things have prices in SAP or if any doubt it can be confirmed by contacting sales.
TCK73037 | User | Service Request | Old ATEX Certificate LCIE 93.C6145X | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-18 12:46 | 2025-11-18 15:15 | 1 | 0 |  | I don't find any with the number you gave me (even for othe products). but here is the current one for the RP2E and I have the old ones also but nothing with the number you gave me.
TCK73027 | User | Service Request | TR: GDM/IN10/PRESSURE GAUGE/BAUMER: RFQ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-18 11:27 | 2025-11-18 15:39 | 1 | 0 |  | we cannot state compliance on our own as the 6.2 of their PMI specification require that the "contractor" (Shell ? Modec ?) to validate our PMI procedure. so what we can do is: request the PMI test procedure, last certificate for PMI gun and operator to David Lecomte (you need the approval as it's internal) Provide all this to the customer. upon approval, validate the number of tests (we do it on all parts in contact as standard but it'll define the number of shot depending on the product). SU will define the percentage of products to test depending on the product type and annexes of the PNI Specification.
TCK73001 | User | Service Request | Montage MEX5-D30.B26 +ASIPD2L3+ D820-10L.2227.Q2009 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-18 09:59 | 2025-11-18 12:02 | 0 | 0 |  | oui aucun soucis le montage est bon
TCK72947 | User | Service Request | 55185193 / Badura Ing. / RFQ RUWAREC Manomter / Offer 20661782 / + deliv. time / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-17 15:12 | 2025-11-27 16:00 | 1 | 1 |  | As Jorg confirmed you yesterday, all those parts from schneider are available with NACE.
TCK72946 | User | Service Request | RE: EIFFAGE-- VDAV --- MF2221 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-17 14:47 | 2025-11-17 17:17 | 1 | 0 |  | Salut Yannick, on regarde ça demain, ça m'a été renvoyé mais je n'ai pas les pièces jointe ni l'historique donc je ferme ce ticket mais on regarde ça ensemble.
TCK72943 | User | Service Request | ZPN204SHN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-17 14:35 | 2025-11-17 17:43 | 1 | 0 |  | you're unlucky, here it's neither... I would say probably a copy/paste error on a Nuclear SHM. but here is how to differentiate: old ZP/T have those switches: [Image] and Nuclear ZP/T have those: [Image] SHM is a very common one in the nuclear, it's probably this one.
TCK72938 | User | Service Request | RPPN-DA3-204 (new RPPN-7D3.204) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-17 14:01 | 2025-11-17 14:48 | 0 | 0 |  | no problem to offer the RPPN but the customer cannot replace the below or spring on it's own, he'll not be able to calibrate it properly afterward.
TCK72917 | User | Service Request | TR: DDP & Délai | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-17 10:46 | 2025-11-17 11:33 | 0 | 0 |  | Vu ensemble, tu rouvre le ticket quand on est sur si c'est du Diff.
TCK72831 | User | Service Request | FW: Inquiry MEX5 For Refrigerant R134A / Alex IPSH | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2025-11-14 10:05 | 2025-11-14 11:59 | 0 | 0 |  | They both exists, here are 2 examples that could replace what you need (I don't see the connection but it matches on what I can see). 11252165 11252164
TCK72797 | User | Service Request | TR: Hydro Funk Anfrage / Kontaktmano / Rü Spezfikation / Nr. from Dial / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-13 17:04 | 2025-11-14 11:40 | 0 | 0 |  | here is the best replacements we can offer with 2 options: DPCE100.816.133.426/4710/8011 (2 contacts 0...-250mbar) DPCE100.816.133.406/4710/8001 (1 contact 0...-100mbar)
TCK72765 | User | Service Request | NEMA 4x Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-13 12:29 | 2025-11-13 17:31 | 1 | 0 |  | we never mede the tests of NEMA and use only IP, usually NEMA4x is required on safety gauges and we would be compliant and even better with the IP67 of MEP. we don't have tables because it's not exact match so it's hard to officially state compliance. concerning the Duplex, you can always ask with an X as the digit of material and make a price request. but we usually offer first the 316L and the HC 276 as standard option (Duplex is a nice name but it's only a stainless steel very similar to the 316L).
TCK72755 | User | Service Request | AROE--Tauernmoos Project - RFQ for Valves & Auxiliary Equipment - BAUMER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-13 10:42 | 2025-11-13 17:25 | 1 | 0 |  | Si il te faut du G1/2 F/F il faudra forcément du spécial mais nos instructions sur les produits indiens ne sont pas claires. Aujourd'hui je te conseillerais l'existant en Schneider: 11236852 XMFD:SAFFSA-G4G4-A https://www.as-schneider.com/gb/en/Products-Services/Instrumentation-Valves-and-Manifolds/E-Series-Manifolds/Block-Bleed-Manifolds/Block-Bleed-Manifolds.html
TCK72749 | User | Service Request | Calcul de stress sur un AGFR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-13 09:58 | 2025-11-13 16:23 | 1 | 0 |  | oui c'est possible, il faut que tu contacte Jean Marie Moreau, il a le logiciel par contre il faut beaucoup plus de données process que ça, il te donnera la liste.
TCK72720 | User | Service Request | TR: erreur configuration site Bourdon TBI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-12 16:24 | 2025-11-14 09:24 | 1 | 0 |  | oui c'est une erreur, on pense qu'elle vient du TBI redesign qui devait être remplissable mais le projet est en pause, on va demander le retrait de l'option sur le site.
TCK72700 | User | Service Request | GOST K MEP5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-12 13:17 | 2025-11-13 09:37 | 1 | 0 |  | when Carlo worked on opening Russian market (GOST approval is for the Russian Area) we got it (enclosed). but with the difficulties to penetrate the market and then the war we never renewed it. In KZ our biggest and only market is the Railway and we have a specific alstom approval so no need for local ones. but our CE/EN/ISO approvals are more widely used of course and cover all the same things as the GOST and more.
TCK72693 | User | Service Request | vdav -CHU - DUFFAU | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-12 11:48 | 2025-11-12 15:46 | 1 | 0 |  | avec ces pouvoirs de coupures ce sera forcement un RPxx et le SIL et l'ISO 13849 ne s'appliquent qu'aux switches électroniques. on a la DESP mais pas les 2 autres. on a le K3 par contre qui est une norme de sécurité aussi (dans le nucléaire) mais on ne teste pas vraiment les mêmes paramètres (c'est plus lié a de la résistance mécanique/materiel).
TCK72663 | User | Service Request | Y91 With Hirschmann electrical connection | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-12 08:18 | 2025-11-12 10:00 | 0 | 0 |  | Y9 was a good call, but PBSN will be available with the 0...5V output. for the connector, Hirschmann is a brand name and we don't work with them anymore but I guess they mean the DIN 43650 and you can have it on both Y9 and PBSN.
TCK72563 | User | Service Request | Coated Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-10 15:00 | 2025-11-12 10:07 | 1 | 0 |  | Then it looks perfectly fine to me, the gasket will be on the raised face so all parts in contact are protected. Jorg might want to check if when we do the coating we reduce the height of the RF, we don't but it doesn't change anithing at the mounting because the bolts are planned with a much higher margin than the 2mm. Best, Nicolas
TCK72546 | User | Service Request | RP2E type N microswitch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-10 11:19 | 2025-11-12 09:45 | 1 | 0 |  | the closest is the standard R: Min 0.4A instead of 0.1A Max 5 instead of 10A between 24 and 48 Vdc otherwise it's pretty close. Also, the tropicalizations doesn't make much sense in an explosion proof enclosure (not in contact with external atmosphere).
TCK72465 | User | Service Request | MEX5-D30.B22.DADL05.2L2A.01/0228 / Tien | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2025-11-07 06:29 | 2025-11-07 09:04 | 0 | 0 |  | yes, it's perfectly valid.
TCK72459 | User | Service Request | 55186480 / Anfrage Linhorn /Nr. C2511040079 / AG 20660646 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-06 17:04 | 2025-11-06 17:14 | 0 | 0 |  | you're right, here is all details I could find: RP2-EL.353 2381229/10
TCK72434 | User | Service Request | Kindly advise model code for temperature gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-06 13:07 | 2025-11-06 15:46 | 1 | 0 |  | here is the datasheet of the TIX and you're right, it's now replaced by the TSFE.
TCK72428 | User | Service Request | PO3139 BAVE - MEX5.B90.K35 | Service Desk > Non IT | Low |  |  | Closed | Janvier, Marine [janma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-06 12:18 | 2025-11-06 16:00 | 1 | 0 |  | Comme vu ensemble, pour la codif (FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org>)
TCK72304 | User | Service Request | FW: RDO ns. rif. AC0922 - Termometro | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-05 13:38 | 2025-11-06 10:11 | 1 | 0 |  | I don't think it's for the thermometer they need it but probably for the thermowell and there is all the usual mentions of norms on our certificates. here is an extract of the page where the relevant norms are mentionned (this is for an 316L AGF, same material as on your customer's certificate).
TCK72303 | User | Service Request | DT2 Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-05 13:33 | 2025-11-05 14:26 | 0 | 0 |  | it really depends on what the customer needs there is 4 elements in contact with the fluid, the diaphragm, the upper part, lower part and gasket. the diaphragm is the first we protect anyway because it's thin so more prone to be damaged. to damage the massive parts it takes a lot more time.
TCK72293 | User | Service Request | TR: Demande de devis K1 FOS ST3800 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-05 12:08 | 2025-11-05 17:11 | 1 | 0 |  | au prix d'une membrane Hastelloy ça protégera en cas de température un peu plus élevée mais si c'est vraiment max 45C, le 316L suffirait en théorie.
TCK72260 | User | Service Request | MEX5+D421 with flushing ring/ Turcomp | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-05 09:49 | 2025-11-05 11:37 | 0 | 0 |  | the replacement for the D421 is the D425, and there will be possibilities for flushing ports. but you'll first need a bit more details like at least the PN/DN and norm of the flange connection.
TCK72230 | User | Service Request | Setting up of WEP5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-04 23:06 | 2025-11-06 09:06 | 1 | 0 |  | for setting up the WEP5 you send a Downlink from your network server software (internally we use chirpstack which is open source) the procedure and the structure are defined in the manual in paragraphs 9.5, 12.3 and 12.4. I attach the manual for reference.
TCK72174 | User | Service Request | 30308375 ¦ Camar Industrial ¦ Caudalimetro + Calibración certificacion // OFERTA Y PLAZO DE ENTREGA SDI 439 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-04 10:51 | 2025-11-05 17:12 | 0 | 0 |  | Défini par email avec Jorg, hardware standard et cadran a remplacer.
TCK72165 | User | Service Request | Y92 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-04 10:14 | 2025-11-04 11:40 | 0 | 0 |  | on a les joints FKM, ce n'est pas le problème, par contre il a été déterminé que le joint adapté pour les applications hydrogène était l'EPDM et l'option 7912 contient le revêtement or, le joint EPDM spécifiquement et le remplissage et dégraissage, le tout pour donner le meilleur résultat sur les applications hydrogène. si le client veut y déroger, on peut techniquement le faire mais ce ne sera pas l'option 7912 et il faudra passer la demande de code par le PM sensor.
TCK72076 | User | Service Request | TR: Tauernmoos Project - RFQ for Valves & Auxiliary Equipment - BAUMER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-11-03 09:54 | 2025-11-04 11:33 | 1 | 0 |  | Voici les plans, pour les mettre en MM03, j'ai proposé mais il faudra mettre en place un process standard je pense. pour le 2eme il n'y a pas le bouchon mais pour un plan d'ensemble ça ne devrait pas être gênant. Best, Nicolas
TCK72036 | User | Service Request | UL Certification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-31 18:22 | 2025-11-05 09:54 | 1 | 0 |  | we never qualified our pressure switches UL as US is not our primary market and they are both ATEX and IECEx. but the microswitch inside (only active electronic component) is UL LISTED - FILE E12252 (code L22).
TCK72031 | User | Service Request | Contact Pressure Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-31 15:31 | 2025-11-04 09:09 | 1 | 0 |  | yes sure, all switch will always be in one position over the setpoint and a different position below. also the 2 switches can be different. you can have a normally open (open below the setp and closed after) and a normally closed (closed below setp and open after). a normally open from the point of view of your customer will go from close to open with falling pressure (going from above the setp to below).
TCK72021 | User | Service Request | 0-3000 lb/in2 150mm Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-31 12:56 | 2025-11-03 18:08 | 1 | 0 |  | it's psi (pound force per square inch lb(f)/in2) so you'll probably not have it with this writing lb/in2 but surely it exists as psi.
TCK71981 | User | Service Request | Presentation wep5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-30 21:56 | 2025-11-03 18:02 | 1 | 0 |  | Presentation sent via teams
TCK71888 | User | Service Request | wep5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-29 23:20 | 2025-10-30 09:27 | 0 | 0 |  | the LoRa gateway are made to connect to multiple instruments as long as they use LoRa protocol. I don't know exactly the gateway they are using of course so it's hard to be very sure but that's how it normally works. in the factory here, we are testing multiple devices from diverse manufacturers and for different applications (temperature, pressure, humidity, level etc...) all connected to 1 gateway.
TCK71838 | User | Service Request | AROE-ABB-- B16116 - Consultation instrumentation de pression pour développement nouvelle gare raclée | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-29 10:15 | 2025-10-29 11:58 | 0 | 0 |  | On peut faire 3 points de soudure sur un montage en gaz sur le G3/4 ou le 3/4 NPT. pour faire un cordon de soudure complet il faudrait créer le D055 en NPT femelle donc plus compliqué et probablement SPR. pour le Ra, sur cette membrane il est <0.8microns.
TCK71836 | User | Service Request | TR: Demande de prix mano différentiel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-29 10:02 | 2025-10-29 11:12 | 0 | 0 |  | oui ça sera quand même assez différent mais tu peux définir un MFT et envoyer la demande au commerce pour un prix avec les contacts (précise bien la puissance, je pense que c'est la que ça va bloquer) mais on proposera au plus proche.
TCK71823 | User | Service Request | BAPL ¦ replacements ¦ E910 ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-29 08:51 | 2025-10-29 11:16 | 0 | 0 |  | the codes contains a lot of mistakes but I can at least say it was E91 which are replaced by Y91 and at the end is the pressure range. in the middle, the 35 could be the connection and gasket type. I would suggest to offer Y91335BxxR (add the different ranges).
TCK71738 | User | Service Request | RE: Demande devis TAITTINGER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 10:02 | 2025-10-28 11:54 | 0 | 0 |  | Quasiment sur que c'est un DAEL (les DADL sont coniques sur les cotés de la face). pour la taille je dirais a la louche un DAEL14 mais ça se vérifie.
TCK71733 | User | Service Request | TR: Affaire GAZECHIM - Certificat Euro Chlore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 09:38 | 2025-10-28 14:30 | 0 | 0 |  | comme tu te doute on a pas surtout qu'a ma connaissance Eurochlore n'est pas un organisme qui propose des certifications... Euro chlor par contre propose la certification GEST 17/492 mais a ma connaissance ça ne concerne que les vannes.
TCK71729 | User | Service Request | TR: [BAKR]_Y91 Pressure sensor inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 09:13 | 2025-10-28 11:48 | 0 | 0 |  | A transmitter doesn't display a pressure with a unit, it just send a 4...20mA, it'll be on his display that the customer will set the unit he wants.
TCK71728 | User | Service Request | TR: [BAKR] Product Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 09:10 | 2025-10-28 11:44 | 0 | 0 |  | Visuellement ça me parait être un: 10290325 MIX7-D61.D24 si ils veulent les 2 echelles ce sera un MIX7-D61.W24.
TCK71725 | User | Service Request | TR: 압력게이지 견적 요청 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 09:01 | 2025-10-28 11:36 | 0 | 0 |  | it was phased out years ago, it was a specific product for customer Baule. I attached the specification for reference but we won't replace it directly like this. I would offer the closest standard which would be : Y91333B25R/2151 + AKPL2F5302 + D0552L14.01/0592/9010 main differences are the connection and it's a 4...20mA not a 0...20mA (a bit outdated). but it would need to go through validation for the mounting as it's really on the limit and to determine the options of mounting and accuracy.
TCK71720 | User | Service Request | TR: FFKM Gasket | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-28 08:42 | 2025-10-28 15:02 | 0 | 0 |  | if they select the connection 09 and the gasket material 3 at the end of the code, the gasket will be this one (in FKM but we don't have an option in FFKM): 22x2 fluorocarbon (FKM) O'ring (hardness: 75 ; shoreA) (SAP component code: 10304594)
TCK71679 | User | Service Request | D853A5N2217320 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-27 15:19 | 2025-10-27 18:12 | 0 | 0 |  | you can have it, the datasheet still doesn't exist but the old one is still relevant (we only changed the way we manufacture it, not it's properties). be careful, the min range to mount is to check (probably minimum 1.6bar range with a MEX5)
TCK71669 | User | Service Request | software to configurate wep5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-27 14:17 | 2025-10-27 17:13 | 0 | 0 |  | please find enclosed the datasheet and manual. LoRa protocol is standardized so you need a gateway but then the softwares you can use as to receive store or display the data are multiples and some very good ones are open source. Internally for demo for example we use chirpstack to connect to the gateway and the gauge and to receive and decrypt the data.
TCK71668 | User | Service Request | TR: FFKM Gasket | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-27 14:11 | 2025-10-27 17:07 | 0 | 0 |  | Ça doit être des produits Suisses ou danois, on ne fait pas de produits avec une connexion G51 chez nous.
TCK71581 | User | Service Request | TR: demande de prix + délai | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-24 14:52 | 2025-10-28 12:00 | 0 | 0 |  | je leurs proposerais celui la, c'est la partie femelle d'un 1/4"HP (écrou 9/16UNF) et c'est le plus utilisé des 9/16UNF femelle chez nous (code 173, C015). [Image]
TCK71550 | User | Service Request | TR: BAKR / Bourdon Instruments UL certificatee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-24 09:02 | 2025-10-27 09:14 | 0 | 0 |  | When we usually talk about UL, we mean the electric safety norm and our products PG and TG are mechanical and not included in those norms (and we don't do TC and RTDs). there is some specific part of the UL where we comply, like for the fire resistance of the MPG6 case, it respects the UL94 V0 but otherwise I don't see what we could have. if they have more details about the UL norm they need us to respect I could check further.
TCK71520 | User | Service Request | PBMN Low Pressure sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-23 15:59 | 2025-10-23 16:28 | 0 | 0 |  | pas de soucis ça se monte sans problème: [Image]
TCK71518 | User | Service Request | MPG6 3D Cad Files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-23 15:25 | 2025-10-27 11:06 | 0 | 0 |  | there you go, I attached the 3D drawing.
TCK71491 | User | Service Request | VDAV --- BIO INOX --- 1095 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-23 11:05 | 2025-10-23 11:45 | 0 | 0 |  | on considère que le type de connexion n'a pas d'influence sur le montage (comparé aux volume et sensibilités des éléments sensibles) donc si c'est ok en type D c'est OK en type F.
TCK71442 | User | Service Request | FW: Suria : MEP6 | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-22 12:17 | 2025-10-22 15:22 | 0 | 0 |  | Answered to Ang, the replacement will be MPG6D61D31 (new design)
TCK71395 | User | Service Request | RP2-NL.345/SETP +DANC02.2L2A.01 / Alex | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2025-10-22 04:52 | 2025-10-22 10:54 | 0 | 0 |  | Answered to Ang this morning, see below: You want a 6bar switch to set at 0.9bar rising, then here are the calculations: Forbidden zones (below 10% and over 90% of the range: below 0.6bar and above 5.4bar). Nothing (set or reset) can fall outside of those limits. Then a setp of 0.9bar rising is possible if the resetpoint is also above 0.6bar. Setp being relay close to 10%, Ill consider the deadband at 10% which is 50mbar (exact calculation can be done with linear extrapolation or proportionality but I dont want to over complicate). So here is your final setting: For Model : RP2-NL.345/SETP Range : 0 to 6 barG SP: 0.9 bar rising, 0.85bar falling. Best, Nicolas
TCK71361 | User | Service Request | GRCH-MONEL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-21 14:09 | 2025-10-21 14:37 | 0 | 0 |  | oui j'ai bien les memes infos, MEM3 400bar et MEM5 600bar.
TCK71345 | User | Service Request | RP2NR346-0201 /IPS | Service Desk > General | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-21 11:44 | 2025-10-21 14:45 | 1 | 0 |  | 11241775 RP2-NR.346/9941_ED02_SETP SETP 4bar decreasing made in october 2023
TCK71287 | User | Service Request | RE: Demande de devis | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-20 15:34 | 2025-10-22 10:35 | 1 | 0 |  | Salut Yannick, je ne trouve rien c'est trop vieux mais on peut redefinir un MVE7 avec des marquages. Best, Nicolas
TCK71283 | User | Service Request | Request for STEP File – Item MEX5D63E20/0751/0078/0760 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Baric, Manuela [bmanu] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-20 14:59 | 2025-10-20 15:24 | 0 | 0 |  | no problem, please find the step file in the ticket attachements.
TCK71223 | User | Service Request | REMPLACEMENT PRODUIT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-20 09:59 | 2025-10-20 12:08 | 0 | 0 |  | non il doit toujours être dispo, par contre EDF n'en a pas commandé à ma connaissance depuis plus de 15ans. on a même du stock sur le micro, mais vérifie bien avec EDF que c'est ce qu'ils veulent je pensais qu'ils ne prenaient quasiment plus que du K3 ce qui serait le CHM dans notre cas (aussi hermétique mais pas ultrasensible, voir les différences d'utilisation ci-dessous). [Image]
TCK71220 | User | Service Request | AIRBUS CHIFFRAGE NOUVEAU PRODUIT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-20 09:43 | 2025-10-20 11:15 | 0 | 0 |  | pour un nouveau produit pour Airbus je laisse Philippe prendre le lead, je pense qu'il n'y a que lui qui a assez expérience avec eux pour répondre.
TCK71157 | User | Service Request | Création sous-traitance FÜD chez Aremeca | Service Desk > Non IT | Low |  |  | Closed | Firmin, Chloée [fchl] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-17 10:32 | 2025-10-17 11:19 | 0 | 0 |  | T'es sur que tu l'a envoyé sur la bonne boite ? Marco Mosberger me l'a réaffecté.
TCK71116 | User | Service Request | TR: CERTIFICAT ATEX MEX5 BAUMER BOURDON sales.fr@baumer.com MBS-0625-124 R2 2510C01078 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-16 16:14 | 2025-10-17 10:22 | 1 | 0 |  | Yannick, voici ou on se situe, et il n'y aura donc pas de certificat directement par l'organisme accrédité, mais notre notification de réception du dossier de qualif (que tu as donné) et notre déclaration de conformité. [Image] Best, Nicolas
TCK71111 | User | Service Request | TR: Demande de prix et délai. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-16 14:28 | 2025-10-16 16:40 | 1 | 0 |  | c'est trop vieux pour retrouver une trace SAP mais avec les photos j'ai la plu part des info, je proposerais en remplacement un Y91323B35R ce sera la même connexion de pression et électrique, toujours ATEX et sortie 4...20 mA, on passe juste sur un peu plus précis et tout inox 316L. de vu c'est un 1/4 Gaz, mais à confirmer ce n'est qu'une photo.
TCK71106 | User | Service Request | BAES- RTNA4-Switch G | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-16 14:16 | 2025-10-17 09:06 | 1 | 0 |  | Salut Yannick, la config ci-dessous est possible: RTNA4 en code 403 avec un micro G ? switch 1: 80C à la hausse switch 2: 90C à la hausse Best, Nicolas
TCK71052 | User | Service Request | Material creation inquiry (3) | Service Desk > General | Low |  |  | Closed | Lee, Na Young [lena] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Mao, Yihua [maoy] | 2025-10-16 01:58 | 2025-10-16 09:17 | 0 | 0 |  | ticketing system error
TCK71051 | User | Service Request | Material creation inquiry (1) | Service Desk > General | Low |  |  | Closed | Lee, Na Young [lena] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Mao, Yihua [maoy] | 2025-10-16 01:57 | 2025-10-16 09:17 | 0 | 0 |  | ticketing system error
TCK71050 | User | Service Request | Material creation inquiry (2) | Service Desk > General | Low |  |  | Closed | Lee, Na Young [lena] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Mao, Yihua [maoy] | 2025-10-16 01:57 | 2025-10-16 09:16 | 0 | 0 |  | I think you got a bit confused in the different ticketing systems here are some information about how it work: FUNC BG IT Servicedesk <servicedesk@baumer.com> is for IT requests FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> is for requirering code creation for mechanical (an a few electronic products). Nicolas Birrien <bourdontechnicalsupport@baumer.com> is for technical support and is normally used only by technical local referent. I let you forward your 3 code creation requests to func-frven-mat-dat@baumernet.org.
TCK71009 | User | Service Request | TR: Quotation request for BARFLEX®. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-15 12:00 | 2025-10-15 14:18 | 1 | 0 |  | flexible hose for barflex with G1/4Female on process side: 10309324 flexible hose for barflex with G1/2Female on process side: 10309326 for adaptors, it works only with the standard dedicated flexible: Standard flexible hose for barflex M16*20 female: 10309319 + adaptor connection with valve M16*20Female to G1/4Male: 10303759 or + adaptor connection with valve M16*20Female to 1/4NPT Male: 10303760
TCK70984 | User | Service Request | MPE6-D61.D31 / Kah Fei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-15 09:35 | 2025-10-22 14:07 | 0 | 0 |  | Bee Tin, here is the drawing of the dial of a MPG6D61D31 (new design): [Image] Best, Nicolas
TCK70982 | User | Service Request | john Cocheril // 80112593 // creation par anticiaption // PBMN-2.4.B22.R.A1. D295-2N20.22P. 01/9010 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-15 09:15 | 2025-10-15 11:29 | 1 | 0 |  | Salut Yannick, c'est pas pour moi ca doit etre pour la codif. Best, Nicolas
TCK70891 | User | Service Request | Recherche article | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-14 09:52 | 2025-10-14 15:43 | 1 | 0 |  | 10270455 vu aussi par teams.
TCK70888 | User | Service Request | doigt de gant AGFR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-14 09:42 | 2025-10-14 17:35 | 1 | 0 |  | oui on en a déjà fait plein, il faudra faire la demande de prix et le digit matière dans la codif sera le "9".
TCK70847 | User | Service Request | RDN8 Pressure Switch Setpoints | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 15:22 | 2025-10-14 17:31 | 1 | 0 |  | you're right, everything seems fine except the last one, solutions can be to rise the setp to what you calculated. or to take a more specific RD specialized on low diff and high stat like a RDN6AXX158
TCK70831 | User | Service Request | HADEN Systems - MEX3 with 1/4 OD tube connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 13:52 | 2025-10-14 12:16 | 1 | 0 |  | yes no problem, if they validatedthe C050 connection or requested it specifically you can just ask for the code creation.
TCK70805 | User | Service Request | MEP6 Over range and under range stops | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 11:18 | 2025-10-16 16:36 | 1 | 0 |  | Closed but do not hesitate to reopen if you have feedback and need more information.
TCK70798 | User | Service Request | MONTE SOUDE MANO +AKPL +SEPARATEUR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 10:47 | 2025-10-14 12:00 | 1 | 0 |  | nous on préfère monter en gaz, bonne étanchéité et démontable au besoin. mais par contre a cause du joint ce n'est pas soudable complètement (seulement 2 points de soudure option 1996). le NPT permet une soudure orbitale complète ce que certains clients préfèrent (O&G) avec option 9030
TCK70790 | User | Service Request | Aerosol Adapter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 10:22 | 2025-10-14 13:56 | 1 | 0 |  | After checking, neither India nor BHFR supplied those adaptors. we only supplied the gauge. Best, Nicolas
TCK70781 | User | Service Request | TR: D821 Seals / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 09:32 | 2025-10-14 11:49 | 1 | 0 |  | pour moi tout passe: les 4 bar en DN3/4: 1.MEX5-430.B19_0751+D821-10L.2213.G20.LRS1 = 3 2.MEX5-430.B19_0751+D821-10L.2213.G60.LRS1 = 2 et les 2.5bar en DN1": 3.MEX5-430.B18_0751+D821-10L.2214.G20.LRS1 = 7
TCK70761 | User | Service Request | RFQ NO 6000222055 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-13 07:50 | 2025-10-14 11:43 | 1 | 0 |  | please find enclosed the old datasheet (I didn't foud the one of the MDP7 but only it's replacement the MP7): Order code: MFT7-P13.F15Q.62 MFT old Code / SAP code:
TCK70714 | User | Service Request | WG: Calibration Certificates | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-10 09:40 | 2025-10-10 10:43 | 0 | 0 |  | here is the reference, both serial number reference the same order and product: 11279608 MEP5-450.B38(R) order: 2371280
TCK70690 | User | Service Request | anciens afficheurs AN44 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-09 15:15 | 2025-10-09 16:21 | 0 | 0 |  | Ci-joint la notice des AN44
TCK70663 | User | Service Request | FW: PROT 25-00042495 RFQ for OPP.TY 1062534880336 10324301 MAXIMUM TURBINE SUPPORT EUROPE S.R.L 3Y | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-09 10:24 | 2025-10-09 10:41 | 0 | 0 |  | I usually can find it because those old references were copied in SAP, but I don't have this specific one. I can just tell you that from the outside, the only option I see is the ATEX marking on the label, so I think it probably was a: MEX2-D50.B35/0751/0078
TCK70613 | User | Service Request | WG: 3D.stp DRAWINGS- AF/4192/2025- YR REF 2393548 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Baric, Manuela [bmanu] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-08 15:13 | 2025-10-08 17:30 | 1 | 0 |  | I can certainly help, this 3D exists and as only the range changes, 1 drawing covers all you configurations. you can find the 3D file enclosed in the ticket.
TCK70600 | User | Service Request | BAIT / Label for CTX / Y91 | Service Desk > Non IT | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-08 14:21 | 2025-10-09 10:27 | 1 | 0 |  | we have the technical possibility to change or add something ont the engraving of the Y91 or the label of the CTX. but we don't yet have a clear option and process to handle it. You'll need to fill an SPR to request the option and justify with some business, but I think it would be good to have that possibility in the future anyway.
TCK70530 | User | Service Request | grch- CHAUDROTECH_demande d'informations pour le montage d'un manomètre | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-07 17:09 | 2025-10-07 17:49 | 0 | 0 |  | pour moi le produit est un 11258836 MEX5-D31.B20.D055-2L1A.01/9941. ils leurs faut un manchon a souder taraudé G1"1/2 (BSP) et tu peux leurs donner la datasheet pour les détails.
TCK70484 | User | Service Request | TR: DA Dexis projet Mano SKF pour kit HP Thermodyne | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-07 11:22 | 2025-10-10 14:37 | 1 | 0 |  | passer d'abord en G1/2 comme discuté sinon on proposera des KHW.
TCK70475 | User | Service Request | AROE -- ARKEMA 6000073686 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-07 10:49 | 2025-10-10 14:36 | 1 | 0 |  | comme discuté sur teams, ton code parait une bonne alternative surtout après le mitraillage des spéciaux
TCK70438 | User | Service Request | TR: SP089944 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-06 17:04 | 2025-10-07 17:35 | 1 | 0 |  | on ne donne pas nos valeurs exacts mais si c'est pour un respect de réglementation en cas de fuite je peux te confirmer qu'il contient moins de 1200 mm3 de fluide de remplissage.
TCK70354 | User | Service Request | RFQ - QPS2538447 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-03 14:46 | 2025-10-07 16:02 | 1 | 0 |  | I would start with offering this, but max pressure of the complete mounting is 0.25bar diff and 10 bar stat is 10bar MX7-833.B10H/0751 + AKPLBF5303 + DT3-2L20.0.2N0.09 [Image]
TCK70349 | User | Service Request | Flushing Rings | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-03 13:08 | 2025-10-07 10:01 | 1 | 0 |  | A new datasheet for flushing rings was finished 6 month ago and is currently under review for launching at the PM. in the meantime, I have a drawing number you can request the price for this and it'll be integrated in the final code with the drawing number the same way as an accessory. 53192-61-AS-07-08
TCK70348 | User | Service Request | RE: Demande / Semhach | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-03 12:53 | 2025-10-07 15:12 | 1 | 0 |  | le numéro nous apprend des choses, K7224 veut dire qu'il a été fait en octobre d'une année finissant pas 7 et que cétait le 227ème produit. malheureusement, étant un RP2E (ATEX), les numéros de traca modernes (après les années 2000) devrait être K7A227. ça plus le logo Bourdon Sedeme veut dire qu'il a été fait en 1997 et nos archives commencent en 1998. malgrès tout, le reste du marquage nous permet de définir la base du produit mais on n'aura aucun moyen de savoir si des options ou personnalisations étaient incluses désolé. voici au moins la base du code produit: Order code: RP2-ER.355/SETP RP2 old Code / SAP code:
TCK70251 | User | Service Request | 3D drawing (STEP File) for M27-7463-B22 / Songful | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-02 08:59 | 2025-10-02 09:27 | 0 | 0 |  | it doesn't exist yet for the M27, but you can order it's creation to sales and send an order, they'll have it created
TCK70221 | User | Service Request | TR: Manomètres DP pour ventillation CEA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 16:42 | 2025-10-01 18:15 | 0 | 0 |  | on ne fournis pas en général ces application car les gens prennent de petits mano dédiés (connections pour filtres a air, faible dépression, compact et pas très cher). nous n'avons pas de produit dédié, en plus pour descendre à 10mbar de diff il faudra un MFT, en terme d'installation et de prix je doute que ca soit interessant mais c'est possible.
TCK70220 | User | Service Request | TR: TBHA Anliege Thermometer mit Sockel.. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 16:37 | 2025-10-02 10:17 | 1 | 0 |  | on a stoppé les étendues de mesure au dessus de 120C sur les TBHA quand on a récupéré la ligne de kirchberg car on trouvais trop dincertitudes sur la calibration au dessus. étant donné que le TBHA est le seul en montage externe on a pas de remplacement en mécanique. et en électronique, un ETTN6 avec un montage de surface pourra aller jusquà 150C. au dessus de ça on aura pas de solution équivalente, il faudra un bossage et l'insertion du plongeur pour faire des mesures au dessus.
TCK70211 | User | Service Request | TR: Request of Quotation BAUMER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 14:40 | 2025-10-01 18:22 | 1 | 0 |  | je suis désolé mais je ne peux pas m'occuper de définir un MIT3D50H23 (or 25) (ou MAX indien et au pire MEX), ils ont des vendeurs capables de définir un produit et même chez nous ça ne devrais pas être un soucis quand même. même BAFR n'a pas osé me demander des trucs aussi simple et je les ai renvoyé a leurs titres de technico-commerciaux plus d'une fois pour des définitions de produits qu'ils sont sensés faire.
TCK70203 | User | Service Request | TR: Référence : Solyro Notre commande HA2507189 AFS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 14:07 | 2025-10-01 18:03 | 1 | 0 |  | oui lacétylène et loxygène sont les 2 gaz a avoir un traitement spécial dans l'EN. mais pour lacétylène la norme dit juste 2 choses, le mano doit être de sécurité (S3) et il doit avoir le marquage acétylène sur le cadran. il n'y a pas d'autres actions spécifiques (comme le dégraissage pour l'O2) les options a mettre sont: 9936/EF04 (on en est pas content avec Nico F, ce serait a revoir).
TCK70147 | User | Service Request | MEP5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 09:50 | 2025-10-01 17:39 | 1 | 0 |  | it's spare parts for satorp, we already have it directly, don't try to quote it.
TCK70122 | User | Service Request | Y91 For LPG | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-10-01 07:24 | 2025-10-01 09:37 | 1 | 0 |  | Every time you have this type of compatibility requests, please remember that only the user has the full picture of his process (components and proportions, temperature, pressure, moisture, contact with air, etc...) and it can drastically change the behaviour of mateirals. this means he's the only one entitled to make material compatibility choices. But with the experience we have with our instruments and the informations we can research (on websites from gasket specialists for example), we can act as advisors. here Viton felt like a good options and several gasket specialists confirmed it (below an example): [Image]
TCK70029 | User | Service Request | MG5-D211/LRS1 / Chemicoat | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-30 09:49 | 2025-09-30 10:44 | 0 | 0 |  | you're right, the code 4710 doesn't appear in the P2F, in fact it's the standard feature (connection on the left). otherwise the code you defined was pretty good, it's just NPT on both sides: MG5-4D60.B20/4710 + DT2.2N20.0.2N0.01
TCK70025 | User | Service Request | BAIT / customer VANTA / request Manometer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-30 09:30 | 2025-09-30 10:39 | 0 | 0 |  | I don't know this customer and it's special products (MFE, railway), the numbering system comes from the french railway authority (SNCF). we cannot disclose the drawings to an unknown railway actor and there is 2 digits missing to define clearly the products. for example for the first line, I have those 2 matching references: 6346 14 006-00-c 6346 14 006-01-D can you clarify the context of this request and get more details ? then we can surely see what to do next.
TCK70003 | User | Service Request | TR: Pression épreuve ARPX - CMP Arles | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-29 16:44 | 2025-09-30 10:30 | 1 | 0 |  | oui, c'est typiquement ce qu'on entend par pression dépreuve, une pression ponctuelle et pas trop longue (quelques secondes ça ne pose aucun problème).
TCK69974 | User | Service Request | WG: Baumer Manometer EN-837-1 -1 bis 15 Bar R 134a Kältemittel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-29 13:11 | 2025-09-30 10:26 | 1 | 0 |  | as we don't have direct replacement for the DRO80 and it's still available, I suggest to replace by the same reference. -1...17 doesn't exist in the EN norm, it'll be either -1...15 or -1...24bar, otherwise a request for special dial has to be made (it'll have an impact on costs and leadtime). the references needed are below: DRO 80/311.111/-1...15BAR R134A/0004 (10320116) DRO 80/311.111/-1...24BAR R134A/0004 (to be created).
TCK69929 | User | Service Request | Manual agfre | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-26 19:10 | 2025-09-30 10:14 | 1 | 0 |  | you're right, I don't thinks we have a manual on the AGF, but I don't know what we could write on it either as it's just a solid sleeve of metal. the only thing relevant would be the tightening of the connection but the customer orders usually a connection he knows. sorry I don't have anything to give you on this.
TCK69863 | User | Service Request | TR: RFQ 301166710/200 MN 107.235.164.001 PROTECTING TUBE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-25 15:39 | 2025-09-26 09:36 | 1 | 0 |  | on ne fabrique pas ces produits (et je ne pense pas non plus qu'haenni les fabriquait, cétait surement déjà sous traité). je pense qu'il faut le traiter comme une gaine spéciale dont on a le plan et simplement faire une consultation de nos fournisseurs sur plan.
TCK69785 | User | Service Request | vdav -BIO-INOX- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 17:21 | 2025-10-02 10:45 | 1 | 1 |  | salut j'ai eu la réponse on est toujours limité a 6bar mini désolé. il y a des tests en cour et on va tenter le 4bar mais pour l'instant on ne le fait pas.
TCK69779 | User | Service Request | TR: price request for BHFR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 16:11 | 2025-09-25 09:27 | 1 | 0 |  | Salut Yannick, j'ai pas vraiment de doc sur les C0xx, juste les config, qui seront surement ouverts chez vous bientôt ;) le C015 c'est de l'UNF femelle, le plus standard: [Image] Best, Nicolas
TCK69760 | User | Service Request | TR: Dépressiomètre | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 14:59 | 2025-09-24 15:07 | 0 | 0 |  | Salut Yannick, si on le fait toujours mais les gens doivent l'acheter directement a Mercedes, on ne peut pas le vendre directement. Best, Nicolas
TCK69757 | User | Service Request | creation // sames // 20103096 // 3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 14:10 | 2025-09-24 15:27 | 0 | 0 |  | Salut Yannick, je pense que c'est plutôt une demande de code pour la base de donnée.
TCK69754 | User | Service Request | FW: Pressure gauge - MO 10461377 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 14:07 | 2025-09-24 15:41 | 0 | 0 |  | their mounting is not nice and not safe (no ejection can be done as it's blocked by the panel. the size of the hole is probably more meant for a diameter 100 I would then offer the MEP5 in panel mounting fitted with a DT2. we don't offer the ptfe coated bottom anymore on 1/2NPT (because it gets damaged by the threads tightening but an option is exotic material is possible: Order code: MEP5-261.B20 MEP old Code / SAP code: + Order code: DT2-2N22.0.660.01 DT old Code / SAP code: and if necessary, the manifold can also be replaced by an AMFD02N
TCK69744 | User | Service Request | RE: BAKR_PSC_Inquiry regarding Pressure Gauges and Temperature Gauges(250923) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Park, Seong Cheol [pseo] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 12:42 | 2025-09-24 15:16 | 1 | 0 |  | for low pressures, please check our product the MCX: https://www.bourdon-instruments.com/fr/en/product-overview/pressure-gauges/capsule-pressure-gauges/mcx5/p/37172 the TBI wiill be possible with special length, you choose special length in the P2F (if you don't know this software ask your PI referent or contactme) and then the option 9003/xxxx (xxxx being the length in mm). for the AGF, you just replace the 3 digits of the length by your length in mm (be careful, a sliding connection on the stem of the TBI will compensate slight errors in length calculation but if you chose a fix connection, everything needs to be very specific to your exact length).
TCK69737 | User | Service Request | Difference 9704 ET 9709 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-24 11:36 | 2025-09-24 12:02 | 0 | 0 |  | MDR pareil pour moi... ces options ne sont pas très claires et du coup ça va dépendre un peu de qui crée ton code. on en a discuté avec Jorg mais c'est extrêmement dur de créer des options de cadran précises et sans équivoque et de toute façon on aura besoin d'options plus larges pour prendre le hors standard mais comme plus large, elles sont aussi moins précises et soumises à interprétation.
TCK69661 | User | Service Request | FW: Customer request: product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-23 12:10 | 2025-09-23 16:42 | 1 | 0 |  | yes it's some TG bimetal thermometers from India, I suggest to offer TBI080 in replacement, the design will be different but it'll be equivalent in functionnality.
TCK69635 | User | Service Request | EQUIVALENCE WIKA TOTAL YOKO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-23 08:52 | 2025-10-02 17:26 | 1 | 1 |  | pour les transmetteurs yoko avec séparateurs Baumer je te laisse définir selon la REQ, il y a juste quelques D920 simples. pour les D912 avec anneaux spéciaux par contre on a pas mal bossé avec Jérôme pour retrouver des affaire ou on avait proposé des choses similaires et te proposer un plan pour arriver a une offre avec autant déléments prédéfinis que possible. les modèles de séparateurs pour remplacer les plans wika 14153751_01 et 14153597_01 vont ressembler à ça: 14153751_01: D9126D0FQ9602623G0A0XX *2 (avec XX selon nouveau plan similaire à 55916-01 mais avec 4 trous (2 fermés par bouchons et 2 avec les tubes + brides 3/4 300 ASME RF HC276) (longueur capillaire et LRS pas sur plan ou descriptifs donc 4.5m et LRS9 en standard a confirmer). 14153597_01: D9126D0FQ9602627G0A0XX *2 (avec XX selon nouveau plan similaire à 55916-01 mais avec 4 trous (2 fermés par bouchons et 2 avec les tubes + brides 2" 300 ASME RF HC276) (longueur capillaire et LRS pas sur plan ou descriptifs donc 4.5m et LRS9 en standard a confirmer). ci-dessous le plan qui pourra servir de base, il faut demander le dessin et le prix à partir des éléments que je te fournis + des 2 plans Wika (14153751_01 et 14153597_01).
TCK69632 | User | Service Request | BTIN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-23 08:43 | 2025-09-23 10:31 | 0 | 0 |  | ça doit être pour Anil en Inde, je lui ai déjà répondu il m'a fait la demande en direct, tu peux classer.
TCK69620 | User | Service Request | Enquiry For Pressure Senosr for Acetylene Application / Herbert | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-23 04:32 | 2025-09-23 10:33 | 1 | 0 |  | this is the code for O2 application: 11011218 Y91-3.3.3.B22.R/0765
TCK69517 | User | Service Request | TR: 견적의뢰 - PRESSURE GAUGE 1품목 / (엔엔에스) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-19 10:38 | 2025-09-22 11:03 | 1 | 0 |  | oui les silicones (BH3 ici) sont beaucoup plus exigeants en termes de fuites donc ils ne seront pas possibles dans les MMD5. mais part sur un MEX5 c'est pas vraiment plus cher et tu auras une meilleure résistance et qualité générale et laccès à toutes les options, une traçabilité etc...
TCK69492 | User | Service Request | set configuration ¦ Y913 0-10bar DN 40 ¦ odp ¦ Zapytanie ofertowe | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-19 08:49 | 2025-09-22 10:58 | 1 | 0 |  | we only provided the Y913 to BAPL on this order, it means the diaphragm seal was probably mounted by a distributor. we do not usually advise it but we can't really prevent it. I would advise to offer the complete replacement this time but I don't have enough details about the diaphram seals. it most probably is the one below, but, the customer will need to confirm: Y91333B22R/2166 (10301459) + DADL042L2A.01/0592/9010. a similar code for reference would be the 11195415
TCK69467 | User | Service Request | Baumer India Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-18 17:25 | 2025-09-23 16:57 | 1 | 0 |  | I had the same request with pictures from Andrew. this is TG bimetal thermometers from BTIN (63mm) we don't have direct replacement but a TBI080 with bottom or everyangle type should replace it.
TCK69451 | User | Service Request | WEP5 Approvals/Certificates | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-18 14:36 | 2025-09-18 15:17 | 0 | 0 |  | if we need to disclose the complete report, I'll need to ask the PM and head of R&D. please send this request to my personnal mailbox and I'll make the request from there (the ticketing system is not great to handle conversation with multiple actors).
TCK69214 | User | Service Request | FW: **SPAM** FW: Pressure Gauge with contact | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-18 09:02 | 2025-09-18 09:59 | 0 | 0 |  | it's hard to be sure about the connection but I'll offer this replacement: 11133651 MR57-D30.B74/4710 (the -1...0.6 is too low for 2 contacts) also I choose the contacts 7 (NO/NC) because I think I see the pointer dragging a contact but it's really not sure and must be confirmed. and the standard electrical connection is on the other side now but it should not be an issue as far as I see on the picture.
TCK69155 | User | Service Request | Need technical confirmation for model DPX100 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-17 13:10 | 2025-09-17 14:18 | 0 | 0 |  | I don't think it makes any sense to stick with DPx150, please offer MIX7 + D821. the bar Bourdon scale (A76) -1...3 bar bourdon scale and dual scale in negative with -760mmHg here is an example of a A79 (-1...9bar bourdon) sorry that's the only I had in negative scale Bourdon but it show you the general idea: [Image]
TCK69084 | User | Service Request | 3712102 / Anfrage Quinta Radisson / 927448 / AG 20652137 / MEP5 with Seal / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-16 16:20 | 2025-09-17 17:30 | 1 | 0 |  | Please set me a call for this, I'll need more information I think. Best, Nicolas
TCK69083 | User | Service Request | FW: BAUMER type Y913 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-16 16:17 | 2025-09-16 16:42 | 0 | 0 |  | sure, here is a screen print of all the details of your order: [Image]
TCK68980 | User | Service Request | SETP check | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 16:07 | 2025-09-15 17:39 | 0 | 0 |  | the official minimum is 145 kPa falling and I have none set below. the 140 would require a derogation from R&D (probably via an SPR).
TCK68976 | User | Service Request | TBI100-254.140.13T/9003/0680/0110/0751/0727/0581. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 15:38 | 2025-09-15 16:00 | 0 | 0 |  | as long as you can select it it's fine, it's allowed, even partial evaporation is fine as it's not a filling liquid but just a grease for the sensing element. please find enclosed the MSDS if your customer needs more details.
TCK68971 | User | Service Request | VDAV -- 30305446 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 14:47 | 2025-09-22 16:00 | 0 | 1 |  | cch vient de passer, attention aux changements de température mais pas de soucis sinon c'est OK: MEX5F31K19 + ARIDF5F53 + Y91333B19R/0592/9010 + D295-2L20.000 +
TCK68943 | User | Service Request | TR: O ring of Y92 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 12:33 | 2025-09-15 15:07 | 1 | 0 |  | voici la référence si c'est du NBR (3L dans le code) mais ça peut être du viton ou de l'EPDM (9L ou 5L) 10303360 JT TOR 22X2 NBR
TCK68938 | User | Service Request | WG: ask | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Berberich, Nina [bnin] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 11:59 | 2025-09-15 15:45 | 1 | 0 |  | this is the RP2-ER.354 (SAP: 10296812), order 1942327/10 in June 2022 (it was handled by 2680 so I have no access to the order to check who was the customer) 3/4NPT is the thread for the cable gland in the case.
TCK68918 | User | Service Request | Temperature Gauge Pockets | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 10:25 | 2025-09-15 10:44 | 0 | 0 |  | this is very short (40mm) what you found is pretty much the best we can do in standard. we don't go lower because on a 6mm stem, the sensing element is longer than 40mm and needs to be immersed. so without our standard TBxxx you can't go smaller than this. I have another solution, you can always do an SPR for special length AGFR, but fot this type of market I think it'll probably be expensive.
TCK68911 | User | Service Request | MG5 ATEX Certificate | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-15 09:55 | 2025-09-15 10:46 | 0 | 0 |  | please find enclosed the right certificate:
TCK68851 | User | Service Request | Repair Cost | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-12 11:48 | 2025-09-12 15:53 | 1 | 0 |  | this is the connection and the sensing element of a kind of RP2. but we won't just replace it as it's not only the connection, it makes everything inside work, they need to replace the full RP2 Sorry.
TCK68848 | User | Service Request | BAES --- TALGO - 4500441958 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-12 11:25 | 2025-09-12 15:51 | 1 | 0 |  | handled via teams
TCK68845 | User | Service Request | MEX5 with range -76cmHg to 2 Kg/cm2 / Azzam | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-12 11:09 | 2025-09-12 15:50 | 1 | 0 |  | it doesn't exist and it shouldn't exist as nothing here is conform to anything (the units, the range, nothing). I would prefer for us and for the customer to offer something more standard (so easier to replace also after) A76 is the bourdon standard bar and will be -760mmHg/bar - 3bar or A74 -760mmHg/bar - 1.5bar here is what an A79 looks like (same design but range -1...9bar) [Image]
TCK68829 | User | Service Request | Certificat KHK Japon - PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-12 09:32 | 2025-09-12 10:49 | 0 | 0 |  | je ne suis jamais contre ces demandes mais une nouvelle certification coûte en general un an de tests et de dossier et pour faire une moyenne en général 10k pour la certif et un truc comme 5k par an pour les renouvellements. en plus je préférerais pour la pérennité le faire sur un autre produit que le PBSN. si l'affaire que tu as peux couvrir ces coûts il faudra demander a Mario. je peux ajouter que nous sommes conformes à l'équivalent Européen de cette norme la PED (DESP en fr)
TCK68792 | User | Service Request | E913 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-11 14:05 | 2025-09-11 15:54 | 0 | 0 |  | on a stoppé le E913 et gardé seulement sa version ATEX pour couvrir les 2 application, standardiser rendais les 2 versions moins cher, sinon le E913 etait plus cher et moins bien. ils on juste l'ATEX en plus et pour moins cher avec le Y913.
TCK68744 | User | Service Request | MX7 with Nitrogen application | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-11 06:17 | 2025-09-11 15:47 | 1 | 0 |  | our gauges are used on a wide variety of fluid and there is no specific option per fluids. differential gauge on Nitrogen could be an application on closed and pressurized Nitrogen storage tank. in this example, the 316L of the parts in contacts is compatible and widely used on this fluid, there is normally no particles or viscuous fluids (or filters are used) so the differential can be used. there might be cold effect as it can be stored compressed or cold (or both) but if the Nitrogen when in contact with the gauges is within the authorized temperatures, I don't see any problems. this is an example of a very common use case to illustrate what you need to pay attention to but your could be different. but the Nitrogen is not really the question.
TCK68736 | User | Service Request | TR: Inquiry on ZDPHN-158CHM/SETP Options | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-10 17:17 | 2025-09-10 17:37 | 0 | 0 |  | oui, 470mbar à la baisse avec 9bar statique c'est ok.
TCK68698 | User | Service Request | TR: Inquiry on ZDPHN-158CHM/SETP Options | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-10 10:04 | 2025-09-10 11:42 | 0 | 0 |  | le SHM est assez standard et le reste du code est ok, pas de points bloquants pour moi, en plus j'ai ouverts tous les codes ressemblants et ils sont en RL (ou RB car sans mouvement), je ne trouve aucun point bloquant tu peux proposer. bien sur il faudra que le SETP soit defini pour creer le code.
TCK68662 | User | Service Request | Baumer India Gauge? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-09 17:40 | 2025-09-09 18:09 | 0 | 0 |  | could be but the connection looks more like a Clamp one, so I guessed the AY, but it's not available with this case type so I asked india directly. Anyway it'll be probably a DP100 or DPC100 on our side, you can start checking this out for the right connection with the customer.
TCK68654 | User | Service Request | Contact Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-09 15:57 | 2025-09-09 17:01 | 0 | 0 |  | I would go with an MR as I "think I see" the magnet on the contact inside (very difficult to be sure) MR54Dx0B24/4710/xxxx the reason why it's class 2.5 is unknown to me, I don't find any reason. and the connection is not visible so I left an X for confirmation by the customer.
TCK68646 | User | Service Request | RPPY-7.M3-206/SETP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-09 14:50 | 2025-09-09 17:07 | 1 | 0 |  | sure, don't offer it ! it's not adapted. you can send me the request on my email so I can take a look and maybe help you define something else.
TCK68603 | User | Service Request | PBMN Low Pressure sur séparateur D820 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-09 10:07 | 2025-09-09 10:42 | 0 | 0 |  | le D820 est le bon choix (D920 est pour les field transmitters genre YOKO, honeywell etc...). le PN10 est ok pour moi mais c'est moins lié au transmetteur qu'au PN du piping client il faudra confirmer avec lui que c'est OK. et pour le LRS4, si tu prends un fluide dédié aux applications O2, il faut penser au dégraissage (0765) 1 x PBMN-24B81RA11403201010_D820-10L.42CH.B20_LRS4_9010_0765 (ou PBMN-24B81RA11403201010_D820-10L.42DH.B20_LRS4_9010_0765 si on doit passer en PN16) 1 x PBMN-24B81RA15403201010_D820-10L.42CH.B20_LRS4_9010_0765 (ou PBMN-24B81RA15403201010_D820-10L.42DH.B20_LRS4_9010_0765 si on doit passer en PN16)
TCK68589 | User | Service Request | TIX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-09 07:58 | 2025-09-09 09:55 | 0 | 0 |  | le code de TIX que tu me donne ne correspond pas a la datasheet. par contre la datasheet nous dit que le TIX est limité a -10C en température ambiante (ne pas en tirer de conclusion hâtive, les composants et les matières sont proches de celles d'un manomètre (tube inox, mouvement inox, boitier inox, cadran alu, aiguille alu etc...) et même si le mouvement est un peu plus sensible et que vous allez perdre en précision hors de la plage ambiante autorisée, il ne devrait pas y avoir de danger pour le thermo (ce que prouve ses longues années de service). je joint l'ancienne datasheet du TIX (maintenant TSS/TSF).
TCK68567 | User | Service Request | TR: Inquiry 72532364 + Extra attachment | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-08 16:34 | 2025-09-08 16:50 | 0 | 0 |  | voici la bonne ref: 1020.0101 S/E SDI 439-AXFLOW 0.7...2M 10323207
TCK68543 | User | Service Request | Gauge and Seal assembly Naphthalene | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-08 13:02 | 2025-09-08 15:03 | 0 | 0 |  | I don't know much about Naphtalene production but I feel like if it cristallize in a gauge it'll cristallize in the seal as the biggest issue will be the 1/4" connection, it feel like a design flaw. it's the one that'll have the biggest entry to prevent clogging. 11125844 D044-2H02 (6mm hole in the connection) [Image] I think it's the best cost benefit solution.
TCK68541 | User | Service Request | TR: Opportunité 01286152 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-08 12:30 | 2025-09-08 14:27 | 0 | 0 |  | a cette pression le 600 / RTJ me parait aussi inapproprié si il n'y a pas de limiteur. mais c'est juste a confirmer, si c'est bon pour eux c'est ok pour nous (ils peuvent avoir standardisé les connections en RTJ 600 mais n'avoir qu'un rating reel plus bas a cet endroit et le mano en 6bar est justifié).
TCK68540 | User | Service Request | ARIN 2L0N | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-08 12:25 | 2025-09-08 14:30 | 1 | 0 |  | it would be this one but it's rare to sell F1/2NPT G1/2F adaptor (no sales since 10y+) so you'll have to request the sales to reopen the code (they also update the price). 10265111 ARIN INOX F1/2NPT G1/2F
TCK68524 | User | Service Request | vdav- etude EVOLLUTECH --AIR LIQUIE DECARBONATION DUNKERQUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-08 10:11 | 2025-12-18 10:07 | 1 | 1 |  | error
TCK68399 | User | Service Request | CTX-3.B.5.B29.A configuré en -1 ... 59Bar | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 17:07 | 2025-09-05 12:04 | 1 | 0 |  | je pense que c'est possible techniquement car on fait bien du -1...15bar avec la cellule du 16bar, par contre il faut quand même faire une demande d'ouverture de code spécial et Mario doit confirmer que c'est ok pour lui (ça nécessite la création dans les config d'une étendue de mesure spéciale et peu orthodoxe).
TCK68397 | User | Service Request | TR: Urgent - CTL-3.2.3.B24.1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 16:50 | 2025-09-05 11:51 | 1 | 0 |  | Je suggérerais de remplacer par le même mais en Inox:
TCK68365 | User | Service Request | WG: MEX5-B60.B33/0684_0771 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 11:29 | 2025-09-05 11:47 | 1 | 0 |  | I think the option 0684 is wrong, as defined, you'll get a single red pointer. I think what you want is more the index pointer as below (0052):
TCK68352 | User | Service Request | TR: MEX5-B60.B33/0684_0771 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 09:36 | 2025-09-05 11:48 | 1 | 0 |  | Répondu sur le ticket d'Andrea, option 0052.
TCK68351 | User | Service Request | Vertification UKCA TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 09:22 | 2025-09-05 11:00 | 1 | 0 |  | Non on a fait la démarche sur aucun produits chez nous. mais je croyais avoir entendue que cette demande de lAngleterre était un peu tombée a l'eau vu que beaucoup de fournisseur ne pouvaient pas refaire certifier tous leurs produits seulement pour lAngleterre. Oui je viens de regarder d'un peu plus prés et pour moi ça n'a jamais été vraiment applicable et la demande est en train dêtre retirée.
TCK68345 | User | Service Request | D912 20645660 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-04 09:00 | 2025-09-08 10:48 | 1 | 0 |  | pour moi tout est en ordre, le D912 peut descendre plus bas que 25mbar et est compatible pour le EJX110A avec les capsules type M et les flasques sont en low volume (W). pas de point bloquant pour moi mais on passe comme tu le suggerais de 6 en D pour le moins de 25mbar. EJX110A-JMSWG-919EB/KU22/D3/C2/CA + D912LD0EF9606L22G0M111 / échelle 588 mmCE EJX110A-JMSWG-919EB/KU22/D3/C2/CA + D912LD0EF9D06L24G0M111/ échelle 246 mmCE
TCK68286 | User | Service Request | replacement ¦ ALH-2-4NM-B (0/10 BAR).PY_PW_GH ¦ Zapytania ofertowe KM/2841/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-03 11:31 | 2025-09-03 14:18 | 1 | 0 |  | the ALH is a standard gauge from our sister company in india (BTIN) I don't know very well if sales unit can order to BTIN a direct replacement but here is a replacement possible from Vendome: MIX7-D61.B22/Q008/Q003
TCK68281 | User | Service Request | FW: MPG6DX Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-03 10:56 | 2025-09-03 14:24 | 1 | 0 |  | yes it's spare part request for SATORP, but it'll be handled locally, they'll not get the business on this. it's a tube connection MPG6 mounted on a parker flange.
TCK68268 | User | Service Request | MCD7 Alteration | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-03 10:10 | 2025-09-03 10:21 | 0 | 0 |  | I think technically it would be possible (with the need off a new calibration because of the welding impact). and very important we need to make sure it's traced on label, markings etc... so that when they order a replacement we know not to send another without the wall mounting. but this is not my call, it'll be quality dept handling this and so their decision. you need to start a return procedure for this.
TCK68239 | User | Service Request | FW: Problems! P/O32982 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-02 20:30 | 2025-09-03 09:28 | 0 | 0 |  | the norm for thermometers is a bit different than for gauges. for gauges it's straightforward, class 1 means +/-1% of the full range. but for TBI according to the norm, here is what the classes means: [Image] so on a 0...120C, we have +/- 2C and both are within the tolerance.
TCK68231 | User | Service Request | TR: SAP code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-02 16:49 | 2025-09-02 17:35 | 0 | 0 |  | c'est un 11012497 MEX2-F20.B22, il vient de l'OF: 103075353
TCK68229 | User | Service Request | TED50 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-02 16:33 | 2025-09-02 17:37 | 0 | 0 |  | bien sur, la voici:
TCK68224 | User | Service Request | duquene -- point sur echelle B 59 en bar bourdon en bar en est en code X. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-02 16:08 | 2025-09-02 16:56 | 0 | 0 |  | Vu par teams
TCK68117 | User | Service Request | AREMECA - CP 1H-1250B Baumer / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-02 06:04 | 2025-09-02 10:34 | 1 | 0 |  | yes, for the manufacturing of those pressure generator, we work with another company from our city, I asked them and they have a seals kit for the model you have so that you have all necessary seals for maintenance. the ref is OP0089 seals kit and I'll send the price and details to sales internally so they can make you an offer for it.
TCK68110 | User | Service Request | TR: TRES URGENT RFQ AVI08049 -- CL174685.1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-01 16:00 | 2025-09-02 16:29 | 1 | 0 |  | voici le cadran que tu auras pour ce code que je te suggère: 11156131 MEX8-F30.B35/0840_0670_0764 [Image]
TCK68018 | User | Service Request | Enquiry For MPE6-D61D22+D204/LRS1 / Pearltech | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-09-01 05:10 | 2025-09-01 09:53 | 1 | 0 |  | it's very old so there is not much I can add, the MPG6D61D22 is available as standard on the website, so I don't think there is any issue. for the D204 I can add that it had a cleaning ring with 1/8NPT hole and was based on a DT2. as we don't have other information, if we go standard on the rest, I would offer a: DT2-2N20.2.260.01
TCK67904 | User | Service Request | TR: MTTFd- Werte / 30303024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-29 08:47 | 2025-08-29 09:04 | 0 | 0 |  | as the Y92 is not in the catalogue since long, I don't have it for it. but it's really similar to a Y913 (same xcept with a diaphragm on the front) and we have the MTTf of this one (see enclosed).
TCK67879 | User | Service Request | 50134922 / Anfrage Druckschalter / Andritz Hydro / AG 20649007 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 15:18 | 2025-08-29 15:33 | 1 | 0 |  | Solved over teams.
TCK67874 | User | Service Request | TR: Inquiry item : Bourdon Sedeme PT100-TR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 14:11 | 2025-08-28 16:51 | 1 | 0 |  | the PT100 TR is a temperature transmitter from the 90s, I don't have a manual anymore but I have the paper version of the datasheet (in french only). it'll be replaced by products like the TE2 from Baumer switzerland.
TCK67859 | User | Service Request | TR: MFX für Firma TL Lausanne | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 12:24 | 2025-08-28 14:07 | 0 | 0 |  | je touche pas du tout aux prix, il faut voir avec le responsable produit.
TCK67851 | User | Service Request | D912 20645660 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 11:16 | 2025-08-28 11:48 | 0 | 0 |  | je ne sais pas avec quel transmetteur tu veux le monter mais on est déjà descendu a des étendues comme ça. pour une vrai confirmation il me faudrait les montages complets. 60mbar 25mbar
TCK67850 | User | Service Request | Satherm Anfrage Nr. 2310/78065/ MEX5 with OD 12x1,5mm / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 11:12 | 2025-08-28 11:58 | 0 | 0 |  | here the product is defined, I feel like your questions are more sales oriented, they can probably tell you which certificates were sold with the product on the last order and they handle the pricing. 11142976 MEX5-DX0.B24/0816_0751_0078
TCK67846 | User | Service Request | TR: [바우머코리아] Bourdon Haenni Pressure Gauge 견적 의뢰드립니다. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 10:58 | 2025-08-28 17:10 | 1 | 0 |  | I checked and we have no traceability of the serial numbers from KHW. I don't think we can completely identify visually but it's for sure a DR on a small fud, probably clamp as far as I can imagin considering how it's tightenned. I would offer something like below and ask the to confirm if it fits their need. MEX5D30B20 + 1621 + LRS1 (could be a 1627 but then you need a MX35 or it could be a 1622, impossible to define just with the picture. the electropolishing is to offer)
TCK67827 | User | Service Request | TR: Anfrage 25-34-1643 //30301526 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 09:41 | 2025-08-28 10:25 | 0 | 0 |  | le E7 pourra être remplacé par un Y9 ou par un PBMN par exemple selon la précision nécessaire. par contre cétait une connexion spéciale que je ne reconnais pas (type flasque supérieur de DT2) je doute qu'on (ou n'importe quel fournisseur, puisse remplacer a l'identique. on peut proposer d'autres séparateurs selon le besoin mais ca veut dire que le client change sa connexion process. sinon si ils veulent remplacer a l'identique il faut le démonter et nous envoyer des photos et toutes les cotes, en espérant que ça corresponde toujours a un de nos flasques de DT actuels...
TCK67816 | User | Service Request | TR: Customer request: product WEP 5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-28 08:43 | 2025-08-29 16:33 | 0 | 0 |  | tu pourras organiser un call avec Mamour et moi, on pourra demander des precisions, peut etre le depanner et sinon on a des moyens de test pour checker le fonctionnement des WEP5.
TCK67800 | User | Service Request | WG: 21107875 / Anfrage Torricelli / Mano + Drumi // AG 20649074 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 17:11 | 2025-08-28 09:31 | 1 | 0 |  | here is the SAP code for D4252L60214G20, including the details for each digits: 10266658 D425-2L60.214.G20: D425 SEAL MOUNTED ON FLANGE BRP WITH WELDED MEMBRANE 2 TOP HOUSING INOX 316L (SS) L G 1/2 FEMALE 6 HASTELLOY C 276 (HC) DIAPHRAGM 0 DIAPHRAGM WITHOUT COATING 2 FLANGE TO STANDARD ANSI B 16-5 1 CLASS 150 4 DN 1" G RAISED FACE (RF) (FACTORY STD) 2 STAINLES STEEL BOTTOM HOUSING 316L (SS) 0 NO COATING
TCK67799 | User | Service Request | vdav--SNCF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 16:50 | 2025-08-28 09:40 | 1 | 0 |  | j'ai trouvé la référence 10516774, c'est en fait un numéro de plan 10 516 774 des années 60. c'est pour un MFE42 avec aiguille avant rouge et arrière blanche. vu le marquage du plan je dirais que cétait un 10 bar (pas complètement sur). il faut repartir sur du standard a partir de ces informations.
TCK67781 | User | Service Request | inquiry Fagerberg / AGW+ Thermometer / AWG / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 14:42 | 2025-08-27 14:57 | 0 | 0 |  | I think you don't have the information for the B because you're looking at the datasheet of the AGWR (threaded) and not the AGW (flanged) I add the datasheet in the ticket for reference. I'm not allowed to give price, this will be a question to ask sales then.
TCK67777 | User | Service Request | 50134922 / Anfrage Druckschalter / Andritz Hydro / AG 20649007 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 14:33 | 2025-08-27 15:15 | 0 | 0 |  | here are the current replacements but most of the codes were correct: RPPAAA3209 2 pcs RPPN7A3209 (10297520) RPPNGA3203 2 pcs RPPN7G3203 (setp option to confirm as limits on G microswitch are complex). RPPAAA3204 2 pcs RPPN7A3204 (11164003) RT-NAB.415.226E3 (11231671) 2 pcs RTAA64152A.6E3(capillary temperature switch) 2 pcs mistake in the code I think they meant this: RT-AAG.415.2A6E3 (10298279) RPPN-7B3.206 (10297606) 2 pcs
TCK67760 | User | Service Request | minimum scale division | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 11:59 | 2025-08-27 14:00 | 0 | 0 |  | this is the scale you'll get for the requested gauge as defined by the EN standard: MEX5-A60.B24/0840_0670 11278567 [Image]
TCK67759 | User | Service Request | 21107875 / Anfrage Torricelli / Mano + Drumi // AG 20649074 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 11:48 | 2025-08-27 12:03 | 0 | 0 |  | My Bad, I took the EN norm with ANSI characteristics, MEX5D30B24/0191 + D4252L60214G2009 (09 is for the filling , I choose a standard but LRS1 could also be fine (01))
TCK67752 | User | Service Request | Coperion / Order 2326883 / A8250243 + Order 2304950 / A8241388 / overpressure by MEP8 / MEP5 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-27 11:08 | 2025-08-27 11:40 | 0 | 0 |  | the RAS2 means changing the tube, the connection etc and all is welded so it' would be really much more expensive than just ordering a new one. considering the RAS1, I think it would be possible in principle but with the cost of sending it back and force + the handling process, adding the option ar redoing calibration, tests, packaging etc... I don't think it really worth it.
TCK67725 | User | Service Request | Coperion / Order 2326883 / A8250243 + Order 2304950 / A8241388 / overpressure by MEP8 / MEP5 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 15:29 | 2025-08-26 17:57 | 1 | 0 |  | Solved on other ticket number
TCK67724 | User | Service Request | Satherm / Eilt Anfrage Nr. 202509132 /Bitte prüfen / Daten unklar / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 15:24 | 2025-08-26 17:55 | 1 | 0 |  | MG5-1D30.B20/XRIN.AKPL.1621/LRS1
TCK67722 | User | Service Request | Confirm model code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 14:55 | 2025-08-26 16:40 | 0 | 0 |  | I had the request from sales and answered last week already. here is the answer I gave: It wasn't easy to find because some parts are missing from the photo (the window and the front bezel). But it seems to be this code. The G1/2 is still to be confirmed; that's more what I see in the photos, but their descriptions say 3/8. 11201833 MR5-4B30.B31/4710_9695_9696_D000
TCK67712 | User | Service Request | WG: MEX5 plus D820 inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 13:32 | 2025-08-26 15:08 | 0 | 0 |  | here is the height of MEX5-D30.B22.D820.10L.42GE.B3001/0078_0751_9941 [Image]
TCK67708 | User | Service Request | SPR BABR 2580 AGWR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 12:43 | 2025-08-26 15:06 | 1 | 0 |  | Je vois, il était en dessous de la longueur mini et on a bloqué, je pense que pour 1200 pièces la bonne façon de faire - confirmer avec le client la longueur (elle contient le filetage donc c'est effectivement très court (1/2NPT = 18mm environ -2mm de fond = 30mm d'insertion réelle) AGWR-0260N2D.050 [Image] et lui dire qu'en cas de d'impossibilité en AGWR, on proposera une alternative en AGFR avec une épaisseur de mur de 8mm int et 14mm ext: 11225142 AGFR-0260N2H.050P [Image] si les dessins sont validés (ou seulement un des 2) on pourra lancer les demandes de prix (le AGWR devrait être un peu moins cher que le AGFR mais le deuxième est plus solide).
TCK67707 | User | Service Request | TOTAL-WEP5-SUPPORT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 12:24 | 2025-08-26 15:41 | 1 | 0 |  | le plus simple si il n'est pas en zone atex est d'enlever les 6 vis a larrière et d'enlever et de remettre la batterie. (les paramètre sont sauvegardés et seront tous repris au redémarrage avec l'aimant). si ce n'est pas possible pour diverses raison (zone Atex, pas daccès à larrière ou autre), voici la procédure. bien mettre l'aimant sur la rainure et contre le coté inox pendant des longues secondes (on a toujours tendance a compter des secondes courtes, et il ne prendra pas en compte si l'appui est de 4 secondes mais 6 secondes ça marchera). si ça ne marche toujours pas, il nous faudra le numéro de série du mano et on verra plus en détails comment vous aider.
TCK67695 | User | Service Request | WG: 55269266 / Anfrage Breuer / Nr. 9990-0213-444 / Offer 20627370 / Articel 10340301 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:22 | 2025-08-26 14:17 | 1 | 0 |  | the 10340301 FUD186/10/1/3 has an Ra < 0.8microns as standard. the version with a lower Ra has not been created yet but is planned in the drawing, it'll need to be created in SAP and a price need to be defined. you can request this code to be created at the material data adress (please make sure in your request to tell them I checked it for feasability). the code to request is FUD 186/10/1/3/1359 (1359 = Ra < 0.5microns)
TCK67694 | User | Service Request | inquiry Fagerberg / AGW+ Thermometer / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:20 | 2025-08-26 17:44 | 1 | 0 |  | I think what you want is the price of this thermowell (76 x AGW-B02GGL2D190) and a thermometer to go with it. if I'm right, the code of the AGW is fine, you can get a price for it through sales. for the thermometer I can suggest a: TBI100-254.160.55T/9003_0200 (price available in the P2F). I took the closest range but other are possible.
TCK67693 | User | Service Request | Anfrage ZNS/ 520512 / Mano + Drumi / Offer / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:19 | 2025-08-26 12:09 | 0 | 0 |  | we don't have the tracking of KHW marking on those seals but I can help you and the customer identify it. in the future as standard there is no marking on our tubular seals but it can be requested as an option. Considering the pictures, I guess it's a 1622 (1") but you can confirm by checking the diameter A (should be 50.5mm) and the shape of the internal diaphragm (should be 4 sided). [Image]
TCK67692 | User | Service Request | Order 2326883 / A8250243 + Order 2304950 / A8241388 / overpressure by MEP8 / MEP5 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:19 | 2025-08-26 17:34 | 1 | 0 |  | we have option for improved overpressure resistance, I'll forward you the slides explaining it. but here I would take another perspective, the overpressure according to norms and international standard of 30% (that was rightfully requested before annotation) covers the design pressure and the case of the need for an external overpressure protector is planned. I'm more worried about the person requesting a special 50% overpressure, because it's possible with RAS1 or 2 but then it's not conform to the norms and standard (EN / ASME) and I doubt it'll comply with the general requirements for his project. I suggest to offer standard and explain him that 50% (or more depending on range and RAS) is possible but will cause him to loose norm compliance. below the table of the resistance depending on range, material of the tube and RAS type. [Image]
TCK67691 | User | Service Request | 55220183 / Anfrage Idoneus / Mano mit Drumi / Offer 20646924 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:18 | 2025-08-27 10:13 | 1 | 0 |  | please find in the ticket the 3D drawing, it'll be the same for both as they are dimensionnally the same (only the diaphragm material changes).
TCK67690 | User | Service Request | 50107653 / Anfrage Centec / F250569 / AG 20648103 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:18 | 2025-08-26 17:13 | 1 | 0 |  | for now on the TBI and TBHI we only have an option for clamp, but on the TSF, they have some choice of sanitary connections. but this needs to be checked from stockach as you handle stiko products.
TCK67689 | User | Service Request | Anfrage Bayer / Webanfrage / Pressure Gauge with Seal / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:17 | 2025-08-26 17:10 | 1 | 0 |  | we don't have DIN11864 diaphragm seals in our portfolio for now. I created a few for an SPR this year but it was the bolted flange form. To me, especially in your sales unit, it's a topic to discuss with our product manager, I see quite often requests for ASME BPE and DIN11864 diaphragm seals and right now except for a few sizes and form, we cannot answer (and even for those sizes, our sales are not informed).
TCK67688 | User | Service Request | 55242456 / Anfrage Steinecker / Ersatzteilanfrage Manometer / Offer / 20647492 / Nr. Gauge / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:16 | 2025-08-26 16:55 | 1 | 0 |  | this is the label on the back of an MX35 but as they are only sold with diaphragm seals, I prefer to give you the code of the complete mounting that it's a part of: 11225660 MX35-F30.B22.DANC08.2L2A.01/0751_9750_++
TCK67686 | User | Service Request | 21107875 / Anfrage Torricelli / Mano + Drumi // | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-26 10:11 | 2025-08-26 16:48 | 1 | 0 |  | I don't have attachment with your request, but I saw you wanted to offer a D425 in inconel if I understand correctly (probably inconel 625). once you found the compatibility on the seals Navigator, you can build in details from the configurators, P2F or datasheets. Inconel is not available as standard for now (request was done and the update of the exotic materials we offer is ongoing but not yet finished). I would suggest to offer another alloys with similar properties but from another trademark, the Hastelloy C276 you'll be able to select it in datasheets and other configuration tools and you should get a price from sales quite easily.
TCK67647 | User | Service Request | HSAN-LYONDELL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 16:24 | 2025-08-25 16:38 | 0 | 0 |  | je ne pense pas que ça vienne de chez nous, la façon dont c'est designé ne me parle pas, la vis est une purge mais il y a pas mal de types de raccords qui peuvent avoir des purges. le type de numéro de plan n'appelle rien non plus chez nous, désolé.
TCK67644 | User | Service Request | TR: MFT accessory | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 15:41 | 2025-08-25 16:33 | 0 | 0 |  | c'est les accessoires de montage sur tube 2", ils sont automatiquement inclus si tu choisis ce montage dans le code du MFT.
TCK67642 | User | Service Request | TR: POMA Systems GmbH - Einkaufsanfrage AF25219 / 30302539 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 15:33 | 2025-08-25 16:32 | 0 | 0 |  | les Y92 avaient été stoppés (pas complètement) mais ils doivent être réintégrés a la gamme, je suggère donc juste de remplacer a l'identique (avec l'accord de Mario pour ouvrir le code). par contre je ne vois pas assez bien sur les photos pour avoir le code sap ou le numéro de commande.
TCK67610 | User | Service Request | MPG6(R) / REX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 10:35 | 2025-09-01 09:03 | 1 | 0 |  | those dials exist on the MPG6 new design (PBT), here is the example of the 16kg/cm [Image]
TCK67608 | User | Service Request | TR: Subsea pressure gauge enquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 10:21 | 2025-08-25 14:58 | 1 | 0 |  | non on ne fait plus le MAH3, et je ne vois pas de remplacement pour des manos submersibles similaires.
TCK67588 | User | Service Request | MEX5-D60.H27/0751_0760_XXXX For Oxygen Service / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-25 08:16 | 2025-08-25 09:20 | 0 | 0 |  | 0765 and 0764 are separated because the European norm states only safety gauges should be used on oxygen so we need to have an option conform to the EN (0765, compatible with safety gauges and keeps the EN markings) and a non conform to EN (0764 to use on standard gauges like MEX and which removes the EN markings). in the case of a psi gauge, there is no EN marking anyway so it would not matter so much but I still suggest to use the same method (0764 for standard gauges and 0765 for safety gauges).
TCK67565 | User | Service Request | Enquiry for TBI100 (Quantity: 5 Nos.) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Manpurkar, Manish [mmani] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-22 15:16 | 2025-08-22 16:45 | 0 | 0 |  | our TBI is IP67 as standard (so better than 66), we have a close available range of -30...170 C/F (code 55V) and painting is potentially something we can do. I would suggest to do it this way: see if they can accept with the 316L case without painting and standard range. if not and if there is more than a couple units, open an SPR, a specific range closer to their need can be defined (there will be a cost for the creation in our system , printer and design) considering the painting, we will need in the SPR the details of the ADNOC specific painting and will be able to tell you if we can do this special painting. you'll receive after the SPR approval process an offer for the compliant product or the closest improvement we can offer).
TCK67534 | User | Service Request | TR: contact -- mano différentiel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-22 08:29 | 2025-08-22 09:21 | 0 | 0 |  | sous 230V AC, le contact 4001 supporte jusqu'a 45mA: [Image] et pour le schema elec, voici les pin de raccordement et la connexion est un presse étoupe de M20x1.5 pour câble de 7 a 13 mm avec à lintérieur un bloc de connexion: [Image] pin5 = terre
TCK67499 | User | Service Request | Enquiry For Pressure Switches ZPN204SH, ZPN206SH, ZPN204SHH / Kah Fei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-21 11:43 | 2025-08-21 15:08 | 1 | 0 |  | those are very old and a lot of things have changed in 30 years but I'm able to define the closest possible current reference. the customer will still need to review the current caracteristics to confirm it's still conform to it's application. also in those old codes, specific instructions, settings etc... could have been added in text in the order, I'll be completely blind to it. I suggest to at least offer the setting for each as it's the most common on those products. ZPN204SH becomes RPPN-7C3.205/SETP ZPN206SH becomes RPPN-7C3.206/SETP ZPN204SHH becomes RPPN-7W3.205/SETP ZTN408SIIINT becomes RT-AYW.408.120E0/SETP
TCK67458 | User | Service Request | SERAC--CADON--AEKT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-20 16:31 | 2025-08-21 10:25 | 1 | 0 |  | Ci-dessous les diamètres des tubulures (j'ai mis les diamètres internes en tolérance mini, c'est le diamètre de passage mini assuré). Product SAP Code Tube used (mm) AEKT-2H
TCK67429 | User | Service Request | Tickets Non relevant to close | Service Desk > General | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Mosberger, Marco [mmos] | 2025-08-20 10:59 | 2025-08-20 11:22 | 0 | 0 |  | Solved with IT.
TCK67397 | User | Service Request | Need confirmation for setpoint in rising pressure with adjustable differential type micro switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-19 15:12 | 2025-08-19 15:35 | 0 | 0 |  | your device can be set at 4.5bar rising, if not specified the deadband will be set at the minimum which means the following characteristics: RPPE-7B6.204 + D425-2N2-021-6G20-01 setp: 4.5bar rising pressure Resetpoint: 4.15bar falling pressure so a deadband of 350mbar (+/-20mbar, I did the calculation with approximation).
TCK67349 | User | Service Request | RP2-EL.N46 with unit Kg/cm2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-18 16:56 | 2025-08-19 15:30 | 1 | 0 |  | there is no scale or display on the RP2 so nobody will read something in any unit. it's only noted on the device plate the range and the setp and I think you could have the unit changed to Kg/cm2 but it'll be a special request a maybe a bit complicated to codify (and not necessary as bar and kg/cm2 are the same value and also you don't read this device). picture of an RT2 (the case is the same, it'd just be bar instead of temperature).
TCK67319 | User | Service Request | Enquiry for MEX3, 0-700 bar /IPSH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-18 11:32 | 2025-08-19 15:41 | 1 | 0 |  | it's possible on the bayonet version but need to be created in SAP. the code below is an example, just without the filling: 11213900 MEX3-D50.B40 on the crimped version (0133) we are limited to 400bar (the higher pressures are still under validation)
TCK67315 | User | Service Request | LLA--- NQ ---RFQ AVI08049 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-18 10:49 | 2025-08-19 17:10 | 1 | 0 |  | Vu via teams
TCK67012 | User | Service Request | EC Type Certificate LCIE 03 ATEX 6300X | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-11 14:57 | 2025-08-19 15:47 | 1 | 1 |  | I attached the ATEX declaration for the YTED but please note it's oblolete and we stopped the renewal of this certification but it was valid for the product they ordered at the time.
TCK66880 | User | Service Request | Enquiry For M12E 4D63 Contact Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-08 05:42 | 2025-08-19 16:00 | 1 | 1 |  | this is very old and so we won't find more details in the system and will be missing information. we'll have to re-define a solution with the customer but I can help you with some information already: Left reference: M12E: snap action contact gauge diameter 150 explosion proof case (we don't have explosion proof cases on contact gauges anymore, I think an RP2E solution will be a good replacement or an ATEX contact gauge but with a different protection like intrisic safety). 4: min/max contacts (5021 or NC/NO) D: bottom connection 6: 1/2NPT 3: BH3 + D101 20 LRS1: DT1 (316L I guess) filled with LRS1 Right reference: M12E: snap action contact gauge diameter 150 explosion proof case (we don't have explosion proof cases on contact gauges anymore, I think an RP2E solution will be a good replacement or an ATEX contact gauge but with a different protection like intrisic safety). 4: min/max contacts (5021 or NC/NO) D: bottom connection 6: 1/2NPT 3: BH3 + D101 20 LRS1: DT2 (316L I guess) filled with LRS1 (to confirm after redefining the max pressure and device to mount with).
TCK66695 | User | Service Request | WG: 50164555 / Anfrage Wagner Group/Offer / 3D Step / File | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-05 09:06 | 2025-08-19 16:30 | 1 | 1 |  | We don't have the 3D available for this one. but you can order it if necessary (there is a SP code for that, it's called something like technical drawing but the sales would know better than me). on most of the KHW/german references, it's a work that was never really started, sorry.
TCK66636 | User | Service Request | DANC Compatibility | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-04 09:46 | 2025-08-19 15:10 | 1 | 1 |  | we have the request sometimes and the issue is the external diameter is the same but we don't know about the gasket, the compression, the design of the grooves etc... I asked to open a project for a new DAxx according to ASME BPE or even to a new seals range to make the clamp and the flanges (ASME BPE also hae flange types) but for now it's not possible to be sure it fits.
TCK66615 | User | Service Request | Need replacement of model DRO 100/413.111/065/0794/9711/1020 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-04 07:58 | 2025-08-19 15:05 | 1 | 1 |  | I can't find any trace of a 413, only 213 but it's about the material and here not in contact with the fluid and the MEX5 is 316L so anyway the best solution. the code for replacement would be: MEX5F31B20/0794/9711 (9711 to be defined if still necessary, otherwise remove) + 1020 + LRS7
TCK66597 | User | Service Request | TR: Need details & price | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-01 16:09 | 2025-08-19 14:55 | 1 | 1 |  | ça n'a pas été facile à retrouver vu qu'il manque des pièces sur la photo (voyant et collerette avant). mais cétait a priori ce code, le G1/2 est à confirmé, c'est plutôt ce que je vois sur les photos mais dans leurs descriptions ils disent 3/8). 11201833 MR5-4B30.B31/4710_9695_9696_D000
TCK66593 | User | Service Request | TR: Pricing Update | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-01 14:58 | 2025-08-01 15:31 | 0 | 0 |  | cette demande contient une incertitude, ils mentionnent RDPN (non ATEX) et ATEX plus bas, et le switch tropicalisé n'existe plus il faudra en choisir un autre c'est donc: Soit un RDN4x6121/SETP soit RDY4x6121/SETP
TCK66576 | User | Service Request | Responds time of the ZTN402CHM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-01 10:45 | 2025-08-25 16:27 | 0 | 1 |  | as there is a special temperature limit, our standard response time cannot really match, I'm sorry but we don't have accurate information for your application.
TCK66572 | User | Service Request | EQUIVALENCES BADOTHERM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-01 09:33 | 2025-08-01 10:23 | 0 | 0 |  | j'ai eu la même question de Gregory en début de semaine, c'est vraiment Yannick qui fait vos définitions de produits même quand vous avez tous les éléments dans un tableau ? je ne le ferais pas pour vous, par contre si tu rencontre un point bloquant nhésite pas a me demander, mais en gros je vois un TBI100, un MEP5 + D820 et un MEP5. je te joint la fiche technique des BDT20 vu que je l'ai a dispo
TCK66564 | User | Service Request | replacements ¦ MPE6 ¦ MPX ¦ Zapytanie ofertowe Baumer - UP/1479/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-08-01 08:17 | 2025-08-01 10:04 | 0 | 0 |  | that's some old gauges and seals (1990s like me). the MPX100 is completely stopped but the MEP5 would be the best replacement and the SE6 was an old special flush seal forpulp and paper applications. we would suggest to replace with a more up to date standard like the D295 or simply D820 (Standard ISO/EN or ANSI flange, flush...). [Image] the MPE is replaced with the MPG new design that is in PBT which is as strong mechanicaly as PP and as fire retardant as Phenolic (UL94 V0) even better.
TCK66534 | User | Service Request | MIX5 PETR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-31 16:08 | 2025-07-31 16:42 | 0 | 0 |  | this french norm isn't valid anymore (since 2006). you can offer MEX5 (direct replacement of what they had) or MEP5 which is more used nowadays with the more restrictive safety norms. then in terms of norms we are covered with the most common and usually accepted in the O&G market (EN837, Nace for the parts in contact and availability of all kinds of specific certificate on request (calibration, material, NACE etc according to EN10204) and conformity possible to ATEX (with option 0078) and S3 if they chose an MEP.
TCK66527 | User | Service Request | Demande de traçabilité / WEIL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-31 15:29 | 2025-07-31 16:34 | 0 | 0 |  | c'est le mano ci-dessous: NumRefArticle : 11151949 DesignationSAP : MEX3-D50.K35/0841_0753_0769 Num commande: 2317669
TCK66472 | User | Service Request | capillaires pour la mesure de pression | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-30 15:50 | 2025-07-30 16:51 | 0 | 0 |  | ça a été remonté et ça devrait être corrigé bientôt. en attendant pour ta demande je proposerais en standard un: AKPL-2F5305 10264792
TCK66460 | User | Service Request | Option 0693 ¦ 11240866 ¦ MJX7-663.B24.D820-10N.2227.G20.09/0693++ ¦ 20645494 ¦Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-30 12:16 | 2025-07-30 16:47 | 1 | 0 |  | G means Gauge pressure which means the standard way of measuring pressure against the atmospheric pressure as a reference. A is used for absolute when absolute vaccum is the reference. most people consider that is there is nothing specified the standard way is relative / gauge pressure measurment and do not specify it. some customer require it to be specified.
TCK66437 | User | Service Request | TR: **Price Required** BHDE 2680 MFT7813B18N32000/V20+AMFD05N+ARIN26F5/MFT7813B15N32000/V20+AMFD05N+ARIN26F5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-30 09:34 | 2025-07-30 16:44 | 1 | 0 |  | Offre Fantinelli clarifiée, demandée et reçu et transférée a Priscilla.
TCK66430 | User | Service Request | TR: Création BAFR // Client : EDF UTO // Cde Clt : 5100-4310874550 // MEX5-F33.B19/0771_9923 // Position 30 | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-30 08:49 | 2025-07-30 09:04 | 0 | 0 |  | je pense que c'est une erreur d'aiguillage, tu es au support technique.
TCK66405 | User | Service Request | D020 (SE20) Diaphragm Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-29 14:59 | 2025-07-29 15:17 | 0 | 0 |  | we did higher pressure in the past (over 1600bar) but I'm not aware of a seal that goes this high. as you can see on the picture below (catalogue from 1981) the SE20 was a tongue seal (we still do this type with the 15xx) but it was not going to 3000bar, it was limited to 1000bar. [Image] if it was not an SE20 but an SE2, then it'll be the DT5 that replaces it (up to 1000bar). and it can be mounted on MEX or MEP series. Concerning the MFP, Jorg added it as a trade product (originally from stiko) and we can offer it but it's handled by BHDE and I'm not sure if diaphragm seals mounting will be possible but you can ask BHDE.
TCK66375 | User | Service Request | SDI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-29 11:36 | 2025-08-22 16:37 | 1 | 1 |  | il savère que non, on en a jamais fait, donc pas possible pour l'instant. ce n'est pas impossible a faire mais il faudrait créer un poste de test, un Template etc... vu lhésitation sur qui fera les SDI à l'avenir, tu pourrais faire une SPR mais je doute que ce soit envisagé.
TCK66371 | User | Service Request | AD model 3d | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-29 11:24 | 2025-07-29 14:34 | 1 | 0 |  | this is a product from our sister company in India, we don't handle it here. he need to contact them (probably via the person who sold him the parts. otherwise here is a direct contact, he can say he comes on my reccomendation (Vijayvargia Anil <avijayvargia@baumer.com>)
TCK66370 | User | Service Request | ZDPWN material. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-29 11:17 | 2025-07-29 15:24 | 1 | 0 |  | you can find most information in the nuclear databook (and also in the validation file EDF and the SEPTEN did that they have access to). please find enclosed the first part of the nuclear databook, the information you're looking for and more are in it. here is the extract of page 19 with the materials of this model: [Image] Best,
TCK66369 | User | Service Request | TSF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-29 11:14 | 2025-07-29 14:27 | 1 | 0 |  | according to what you're telling me you're right, but I only find one code that is close to what you have and it's not really matching, the requested length was 150mm. we would need to compare the code ordered with what was delivered and also you need to ask this to the responsible sales unit (stockach BHDE) please find the product order and give this information and the pictures to BHDE, they should be able to help.
TCK66333 | User | Service Request | Confirmation références Yoko | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 15:29 | 2025-07-29 15:00 | 1 | 0 |  | le D912 est probablement le bon choix et tes codifs sont cohérentes avec la demande. il faudra un peu plus d'info sur le transmetteur pour valider le digit F et peut être réviser les longueurs de capillaires mais pour une première approche ça me parait bon.
TCK66329 | User | Service Request | RE: Demande d'équivalence / GRCH / GNT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 14:52 | 2025-07-28 15:30 | 0 | 0 |  | je ne peux pas faire le boulot a ta place mais je peux t'aiguiller et te donner les bons éléments et si tu as une question tu pourras bien sur me la poser. les FP6 sont des pressostats standards de Georgin et je pense que notre meilleure chance est de proposer des RP2. ci-dessous l'exemple du premier et sachant que leur connexion standard est du G1/2, la définition d'un RP2 se fera toute seule il te reste juste à définir léchelle et le setp pour chaque et beaucoup sont les mêmes. donc ici par exemple pour le repere PSLL22115: RP2NR355 (seuil bas a 2.5bar au lieu de 2bar et attention a bien relire nos datasheet, note pouvoir de coupure est un peu plus bas que celui de Georgin). Je joint aussi la datasheet de Georgin.
TCK66317 | User | Service Request | Document technique MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 12:12 | 2025-07-29 16:28 | 1 | 0 |  | ce sont les SP2138 et 2139 qui définissent les limites, le nombre de cycles de résistance et le temps. elles sont validées en interne suite à des essais labo mais ne sont pas diffusables sans accord express de la direction. ci-dessous un extrait pour t'aider avec le client, sinon je devrais demander l'autorisation de diffusion officiellement.
TCK66315 | User | Service Request | Pin-out ED510 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 11:45 | 2025-07-28 16:56 | 1 | 0 |  | C'est assez vieux, coup de chance j'ai gardé le catalogue complet, ci joint la datasheet avec les details des cablages
TCK66307 | User | Service Request | VCR Female Connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 09:56 | 2025-07-28 17:07 | 1 | 0 |  | as far as I know VCR connectors are used only on UHP gauges which we stopped offering long ago. also the VCR is a registered connection which means I don't even know if we could design and offer it. the best this is to first check if it's Ultra High Purity, then we just cannot offer. if not I would suggest to use more standard connections.
TCK66300 | User | Service Request | TR: SDI enquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 09:14 | 2025-07-28 15:35 | 1 | 0 |  | Je n'ai aucun SDI avec une échelle basée sur 57000L existant donc ce serait un nouveau produit, pour ça il faut faire une SPR directement a Jorg, il n'y a que lui qui maîtrise vraiment ces produits au point de savoir ce qu'il faut faire pour en créer un nouveau.
TCK66296 | User | Service Request | DV-387598 - LILLY | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-28 08:48 | 2025-07-28 14:59 | 1 | 0 |  | le remplacement pour le code 10279966 serait pour moi le 11244998: MEX5-A33.X81/9712_D000 (-1...15bar black / NH3 blue/red)
TCK66206 | User | Service Request | Demande d'équivalence / GRCH / GNT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-25 15:35 | 2025-07-28 15:00 | 1 | 0 |  | Ticket will be answered on TCK66329
TCK66197 | User | Service Request | Barflex 4 ¦ no: 4500394603710 ¦ Zapytanie ofertowe Baumer - UP/1442/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-25 13:51 | 2025-07-25 16:20 | 1 | 0 |  | It's too old, I don't think I can find out with this number, but below there is a date code, with a customer name and pictures of the full product I could narrow the search. please reopen the ticket if you can get those informations. Best, Nicolas
TCK66184 | User | Service Request | TR: RE: Demande d'information sur un matériel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-25 10:52 | 2025-07-25 11:49 | 0 | 0 |  | je ne trouve rien même dans les vieux catalogues SEDEME, mais je dirais que ça ressemble a un codeur, je dirais qu'il faut voir chez Baumer mais je n'ai vraiment pas de détails sur ce produit. Désolé.
TCK66169 | User | Service Request | TR: BAUMER E106NV8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-25 08:59 | 2025-07-25 11:30 | 1 | 0 |  | je peux te dire ce que c'est et te donner la datasheet (ci-jointe la seule difference c'est que c'est une valve a 90) mais c'est un produit indien et je ne sais pas si ils peuvent le commander à l'inde. on a rien au catalogue Bourdon pour remplacer, sinon il faut qu'on fasse une demande équivalent à nos fournisseurs de manifolds (aptek ou schneider) mais je trouve que ça va compliquer les choses et nous faire tourner en rond (mais c'est possible).
TCK66156 | User | Service Request | NIH --- TMH --- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-24 17:31 | 2025-07-24 18:05 | 0 | 0 |  | ce ne sera plus a l'identique et si les 2 aiguilles sont liées a 2 tubes on ne fait plus du tout mais on peut remplacer par 2 manos, plus simple et moins couteux et compliqué a long terme (et meme faire un marquage (voie 1 / voie 2 par exemple) si c'est une aiguille et un repère (mais j'en doute) il faudra faire une proposition plus standard, le cadran a 90 n'a pas d'interet, la peinture n'est pas necessaire etc...
TCK66149 | User | Service Request | Bourdon Gauge Issue FW: [EXT] Bourdon Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-24 16:14 | 2025-07-24 18:01 | 0 | 0 |  | it's from an order of 2024 and there is clearly something wrong (and their explaination could be correct, a leak where they point it could lead to this kind of behaviour) I suggest they contact the seller and start the claim process. in my opinion they'll be replaced with new ones (at least the diaphragm seals).
TCK66096 | User | Service Request | Confirm if we can still supply model DRO63+1095 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-24 08:45 | 2025-07-29 10:47 | 0 | 0 |  | this mounting is validated. MEX3-F21.B77/0751 + 1095/1359 + LRS1
TCK66078 | User | Service Request | MEX2 on Inert Nitrogen Gas Line | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-23 17:19 | 2025-07-24 09:27 | 0 | 0 |  | Nitrogen pure and gazeous (what you have here) is inert, it doesn't react with anything so no problem. Nitrogen is also used in it's liquid form as a cryogenic liquid and so even if corrosions or reactions are still not an issue the temperature is. on material resistance it's really complex because you need to know very good basics in chemistry, the big use case in industries and the different material behaviours. I drafted a presentation and we could go through it when it's ready.
TCK66015 | User | Service Request | Baumer Mexico ¦ Bourdon Requeriments | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Alfonso, Ivan [ial] | Birrien, Nicolas [bin] | MXQUE - El Marques | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-23 02:31 | 2025-07-23 09:29 | 0 | 0 |  | I think you're looking for a sales contact and I can only solve the technical / product related ones. in terms of sales contact, here is who you should direct your request to: BHFR business unit export sales: contact: De Lestrange Priscilla <PDeLestrange@baumer.com
TCK65978 | User | Service Request | DEMANDE DEFINITION PRODUIT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-22 10:25 | 2025-07-22 15:59 | 1 | 0 |  | on a trouvé aucune traça avec les TAGs, Numéros de série ou autre pourtant on trace depuis 2012. je n'ai aucune certitude vu les photos que les séparateurs et le montage viennent de chez nous. pour le montage d'un STD720 tel que sur la photo sur un séparateur, on proposerais un: D912L100A9202L19G0A000 (montage direct selon datasheet figure 3) à voir si ça peut leurs convenir.
TCK65974 | User | Service Request | MEP5-D61.F24.DT2-2N20.0260.09 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-22 10:12 | 2025-07-22 16:06 | 1 | 0 |  | DT2 will be limited in temperature but DT8 is OK, you can say I Validated it.
TCK65937 | User | Service Request | Example for application for the paper industry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-21 17:38 | 2025-07-21 18:08 | 0 | 0 |  | I'd like to help here but the best way to search is in the CRM with the category of customer "paper industry" but I don't have logins for CRM anymore, you need to ask sales. What I can say is the D295 was developped for the needs of this industry. but I checked in trainings and success stories and I have no examples (our D295 are most probably all on applications with pulp and paper but we supply most of it through distribution network. if sales and CRM can't help, come back to me, I could still ask our french sales unit, they sell most of the D295.
TCK65920 | User | Service Request | ETE/FLOEX - grch - 2502860 - | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-21 15:01 | 2025-07-21 18:01 | 1 | 0 |  | tes 2 solutions me paraissent bien selon le DN/PN, peut etre que le mieux pour un anneau en simple taraudage 1/8 est un anneau sous un D820 ou D821.
TCK65861 | User | Service Request | TR: TVE-FRAMATOME- SDI439 / EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-21 08:57 | 2025-07-23 12:05 | 0 | 0 |  | j'ai une seule correspondance dans le système, je te mets le détail ci-dessous pour confirmation: Points essentiels: cadran 100 000L Cadran de dessous 1...3M 11075864 [Image] [Image]
TCK65859 | User | Service Request | TR: SDI enquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-21 08:47 | 2025-07-21 09:53 | 0 | 0 |  | les 2 premières photos sont le même code: 11120206 la 3ème photo (second cadran) est un: 11121937
TCK65807 | User | Service Request | replacement ¦ 11005670 ¦ DRO 100/211.111/0..10 BAR/0730 ¦ zapytanie PILNE! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-18 11:46 | 2025-07-18 16:52 | 1 | 0 |  | it is still available. but the edgewise tip pointer in the filling of the gauge is something we would not reccommend. it acts like a paddle and put stress on the axis that could lead if any vibrations or pulsation (what we usually use the dampening liquid for) lead to the pointer going off the axis. my reccomendation would be this gauge which is cheaper, full stainless, 316L parts in contact and fine graduation, just not with the knife edge pointer: Order code: MEX5-D31.A22 MEX old Code / SAP code:
TCK65765 | User | Service Request | 55269266 / Anfrage Breuer / BW1 TMP0000025 / gleiche Anfrage SAR 03/04 + Breuer Anfrage 07/04/ Offer 20638443 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-17 16:24 | 2025-07-23 14:34 | 1 | 0 |  | the mounting is possible except it'll be a Y91333B26R/2164 + MEXD3xB26/0751/0078/9711/0770 + ARIDxFF3 + 1266/25/9/3 + LRS (to be completed and confirmed, I based it on the drawing, I might miss something). upon reception of confirmation we can do the 1266 this afternoon, the mounting would be possible.
TCK65755 | User | Service Request | D055 Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-17 14:41 | 2025-07-18 16:42 | 1 | 0 |  | we reviewed the limit on the D055 this year because of the differences in limits with the same product from the german factory and our new capabilities in terms of thin diaphragm welding. it was integrated in the Seal Navigator and updated for the next round of updated datasheet but it might take a while. the Seal Navigator is the most updated tool please follow it.
TCK65733 | User | Service Request | Datasheet or Drawing MIT2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-17 10:24 | 2025-07-18 16:39 | 1 | 0 |  | MIT2 is a very specifi product (only like 5 codes exists) I'm not aware of any datasheet. but as it's a product handled by Stockach (BHDE) if there is anything available, Jorg will be aware and responsible for this. please get in touch with him for this.
TCK65686 | User | Service Request | D820.10L.42GE.B20 in 904L Stainless Steel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-16 15:23 | 2025-07-16 16:41 | 0 | 0 |  | the 904L is also called Uranus B6 and you'll find it in our datasheets. depending on if you need only the diaphragm or all parts in contact in 904L, the DN/PN and the type of face, you can find the different solutions in this table.
TCK65670 | User | Service Request | TR: AMF ¦ RFQ AMF25000283 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-16 11:33 | 2025-07-16 14:40 | 0 | 0 |  | le voila: [Image]
TCK65667 | User | Service Request | Stop pin 0722 & 0718 For MEP5 / Azzam | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-16 10:59 | 2025-07-16 16:56 | 1 | 0 |  | the stop pin at 0 cannot be done because here we have to go below 0. but the stop pin at 6 is possible and available on the MEP5 including on the new design.
TCK65657 | User | Service Request | Sealing for ARIN26F8 / APS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-16 10:08 | 2025-07-16 16:36 | 1 | 0 |  | you're right and I had to check this out, when it's G connections, the design is like this and we produce internally: [Image] for NPT, we buy it to a third party and the design is different to make it work with NPT: [Image]
TCK65599 | User | Service Request | TR: libelle separateur projet ELYSE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 16:21 | 2025-07-15 18:33 | 0 | 0 |  | oui c'est le flasque sup des DT2, D425 et D621 (en 316L) si c'est pour l'origine, je pense que tous les codes ne sont pas encore convertis.
TCK65591 | User | Service Request | A. Beranek Bestellung Order B250388 / Offer 20641389-2 + Order 250392 / Offer 20642727 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 15:53 | 2025-07-15 18:30 | 0 | 0 |  | please find enclosed the declaration for the filling fluid and for the diaphragm seals in contact (potentially) with the process.
TCK65573 | User | Service Request | 55269266 / Anfrage Breuer / 9990-0085-579 / Offer 20643240 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 13:20 | 2025-07-23 17:03 | 0 | 1 |  | this is possible but Christophe told me the DN08 doesn't exist, it should be a DN10 which would make the code: Type: DRC 100/813.133/065/0677 (or MEX5) + 1627/1310/1359
TCK65570 | User | Service Request | MEX3+DANC with options | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 12:47 | 2025-07-15 14:20 | 0 | 0 |  | it's simply because those options doesn't exist on MEX3 (the dashpot movement only exist on the big size movement (diam100 and over and the adjustable micro pointer would be too big for the small gauges). but here is a mounting that will give you the same effects and protections (and possible to mount on a DANC): code: MEX3-D20.B19/0679_0770_0751 MEX old Code / SAP code:
TCK65568 | User | Service Request | Certificat Feu fumé EF6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 12:04 | 2025-07-15 14:08 | 0 | 0 |  | voici les certificats feu fumée des EF6 avec les différentes sorties électriques.
TCK65564 | User | Service Request | ASME BPE Certification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 11:27 | 2025-07-15 11:54 | 0 | 0 |  | the ASME BPE is not so much a certification as it is a design convention, our clamp design and sizes is done according to the european ISO 2852 or DIN32676. sorry but we don't have yet ASME clamps, they can check if the size they want are cross compatible with ISO.
TCK65552 | User | Service Request | TR: [BAKR]_MEP3, MEP5 Ambient temperature inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 09:27 | 2025-07-16 16:59 | 0 | 0 |  | on sait que ça ne pause pas de problème mais c'est rentré dans des considérations de test labo etc... on aura pas de réponse dans limmédiat, idéalement il faut faire une demande de dérogation officielle a David qui dit que c'est OK pour une utilisation a -25 au lieu de -20C et dire au client que ça pourra se faire mais que ça prendra un peu de temps.
TCK65550 | User | Service Request | TR: Material inquiry / 오스 문의 건 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-15 09:23 | 2025-07-15 09:41 | 0 | 0 |  | it's a bit tricky here because I don't find any match but we know the application, we'll need to check with the customer if the replacement is OK. you can tell them the one they have is quite old and since then, on this application we mostly provide the one below. all mechanical parts are the same (0.7...2m sensing element and NG8mm cable gland connection) but the subsuppliers in charge of it might have changed. he can check and if ok replace with this one: 10262332 SDI 439/131.211 0,7...2M
TCK65506 | User | Service Request | WG: 80116082 / Anfrage Ludwig Maschinenbau / Block&Bleed valve GV8 Type FF + MF / Offer / Offer 20642674 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-14 11:48 | 2025-07-15 09:31 | 1 | 0 |  | Simone, Thanks for your feedback. Best, Nicolas
TCK65454 | User | Service Request | MPG6DX1B19+FC_M12A8T600-XXXXSP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-11 14:42 | 2025-07-11 16:43 | 0 | 0 |  | it should be one of those codes, but to determine which one you might have to check if you sold them one of the reference in the past or give them the detail of each to make their choice: MPG6-DX0.122.XBR:FCBM12A8T600/0816_9711+ 11168457 MPG6-DX0.176.XBR:FCBM12A8T600/0191++ 11168458 MPG6-DX0.B19.XBR:FCBM12A8T600/0191++ 11184675
TCK65450 | User | Service Request | EIL - Anfrage-Nr.: 04/20230509 / Angebot 20520563 = neu 55269266 / Anfrage Breuer / 9990-0234-507 / Offer 20642871 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-11 14:18 | 2025-07-11 14:29 | 0 | 0 |  | yes it's fine: [Image]
TCK65433 | User | Service Request | Matériaux combustibles PP10R-1.3B22R.A144.062200.050 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-11 11:07 | 2025-07-11 14:25 | 1 | 0 |  | ce n'est pas vraiment possible de répondre comme ça je ne sais pas quelle définition ils font d'un "matériel combustible". par contre si c'est du PP10R je suppose que c'est pour le ferro et qu'ils en ont besoin pour confirmer la conformité feu fumée selon la 45545-2:2020 et dans ce cas la j'ai le document qu'il te faut.
TCK65418 | User | Service Request | TR: manual for assembling of 11125168 from EF6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-11 08:46 | 2025-07-11 09:51 | 0 | 0 |  | the instruction for the wiring of the DIN43650 on an EF6 is in the box with the product. I enclosed a copie.
TCK65409 | User | Service Request | product recognition ¦ Zapytanie czujnik | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-11 07:19 | 2025-07-11 09:36 | 0 | 0 |  | this is a swiss product, you can check with them.
TCK65394 | User | Service Request | 80116082 / Anfrage Ludwig Maschinenbau / Block&Bleed valve GV8 Type FF + MF / Offer / Offer 20642674 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-10 15:32 | 2025-07-10 16:22 | 0 | 0 |  | the GV8 valves are products from BTIN (in india) I don't know if we can offer it. but if Jorg approves, then you can just send the request for code creation (with the approval) to the material data ticketing system. they'll create the codes and consult india for the prices. if it's just for an offer for now, you can ask sales (also if jorg is OK).
TCK65377 | User | Service Request | FW: MEX5-D30.F29.D824-10L.2227.G70.02 (11261092) / IPS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-10 12:22 | 2025-07-10 15:01 | 1 | 0 |  | it's a good question but good practice says that you should use gauges inaround the middle of the range, not in the extreme, so a 60bar gauge is probably selected to measure 30bar which is in the resistance of the flange. but I'm a bit curious of how you get the number of 51.7bar at 50C, because I don't have the same.
TCK65343 | User | Service Request | 55269266 / Anfrage Breuer / BW2 0184.9037K / Offer 20638547 / Anfrage auch 27/05 /+ SAR Hr. Kruk 27/05 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 17:01 | 2025-07-23 14:47 | 0 | 1 |  | this code is technically feasible: MEX5-F30.B26/0751_0760_0794_FÜD1266_LRS1 Industrial pressure gauge (MEX) with stainless steel Bourdon tube Housing can be filled with damping fluid Nominal size 100 (5) Housing 1.4301 (304) Connection at the rear (F) Process connection G1/2 (3) without damping fluid (0) Measuring range 0...25 bar (EN) (B26) Multilayer safety glass (0751) Blow-out device on the rear of the housing (add. ) (0760) silicone-free incl. dial imprint (0794) 1951 Inline diaphragm seal longitudinal axis Vertical attached to inline diaphragm seal (FÜD1266) Connection DN 25 Sterile screw connection ECO Tube Pressure transmission fluid -15C...150) LRS1
TCK65342 | User | Service Request | 55269266 / Anfrage Breuer / 9990-094-377 / DRCE + 5056 / mit BBS Drumi / Offer 20628674 / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 16:59 | 2025-07-23 14:36 | 0 | 1 |  | the code is possible to produce. 0186.9403 DRCE100/811.133/6 B/FÜD187/20/1/3/5056 10343702
TCK65337 | User | Service Request | WG: Customer request: product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 16:29 | 2025-07-09 17:39 | 0 | 0 |  | what he defined is a good choice is he need a differential pressure gauge, you take the closest range available. but he says "pressostat" which means pressure switch and if it's the case you need to check the RDN (maybe it's just choosing the wrong word). and considering compatibility, in both cases, you'll have mainly 316L stainless which is perfect and gaskets chosen mainly for the oil and gas business so no issues with this.
TCK65313 | User | Service Request | Transmetteur ED701 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 14:12 | 2025-07-09 17:32 | 1 | 0 |  | oui c'est une connexion spécifique client (FEDEGARI type) mais on ne l'a pas faite depuis des années, vu le reste de la codif (j'ai trouvé une approchante) je proposerais un PBMH avec les connexions hygiéniques plus actuelles. ED 701 PIEZORESISTIVE INDUSTRIAL PRESSURE TRANSMITTER R20 CONNECTION FLANGE TYP FEDEGARI OUTPUT SIGNAL 4...20 MA CE CONFORM 4R4 ELECTRICAL CONNECTION DIN 43650 ACCURACY: +/- 0,1 %FS ABSOLUTE PRESSURE A25 0...5 BAR 11 WETTED PARTS STAINLESS STEEL OIL FILLING WHITE OIL A Specific label and test certificate Fedegari Autoklaven AG * * * * * * 9460 FEDEGARI PROCESS CONNECTION ACCORDING DRAWING JZ D 13808.5000 (TO39) 2219 CALIBRATION: MEDIUM TEMPERATURE 50...150C WITH AMBIENT TEMPERATURE 9007/0120 TEMPERATUR OF MEDIUM 120 DEGREES C
TCK65281 | User | Service Request | Anfrage / IGB Fraunhofer / Diff. Druckmessung / Offer 20642349 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 10:11 | 2025-07-09 17:10 | 1 | 0 |  | it's Female G1/2 to male G3/4 in 1.4404. there is no article text, just this description: "RI F G1/2 M G3/4 1.4404" "RI" means "Raccord Intermediaire" in french which translates to "intermediate connector".
TCK65265 | User | Service Request | TR: Recommend item : Level gauge. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-09 08:38 | 2025-07-09 09:17 | 0 | 0 |  | c'est un 10262332
TCK65219 | User | Service Request | AKPL F F | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-08 12:09 | 2025-07-08 14:18 | 1 | 0 |  | c'est possible mais spécial, comme d'habitude, question prix si tu as peu de pièce un raccord adaptateur sera plus intéressant sinon tu dois faire une demande de prix pour un AKPL-BF6H02. mais ça passera par une demande de prix, un plan, les fournisseurs etc... et le nombre de pièces jouera bcp sur le prix.
TCK65205 | User | Service Request | EN837 Dials MEX8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-08 10:35 | 2025-07-09 09:11 | 0 | 0 |  | please find enclosed the example of a diameter 160mm in EN837 psi/bar. I want to emphasize that those dials are non standard and will need efforts to be created on top of not being explicitely compliant to EN (it's not completely clear on the interpretation of the norm). I want to emphasize, that for replacement, compatibility, simplicity, price, mistakes avoidance etc..., I think standard bar/psi EN dials would be much more effective for everyone. but here is an example dial and the EN marking would be over the center. [Image]
TCK65197 | User | Service Request | RFQ for PGs in 10P176 RAS TANURA (need offer for Non-metallic seal model D030) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-08 09:40 | 2025-07-08 12:14 | 1 | 0 |  | here is the file filled with the product and deviations. for 7bar max, I offered some 10bars but 6bar range could also be an option make sure you account for the certificates required.
TCK65153 | User | Service Request | TR: RFQ 445165208 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 17:12 | 2025-07-08 09:56 | 0 | 0 |  | je pense qu'ils parlent de ce SDI: 10262332 [Image]
TCK65145 | User | Service Request | TR: LTSD Needed 4101269034/4101386733/4101419208/4101439549/4101579613 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 15:25 | 2025-07-07 15:34 | 0 | 0 |  | j'ai passé les numéros de DTR en recherche mais je ne trouve rien et vu les noms des composant qui semblent être des diodes, je ne pense pas que ce soit des choses de chez nous, on a qu'une diode qui va chez Alstom et elle est dans les MFE et on ne la livre pas comme ça donc elle n'est pas sensé avoir un numéro de DTR. la demande a du arriver chez nous par erreur.
TCK65119 | User | Service Request | Pressure Transmitter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 12:26 | 2025-07-07 15:22 | 1 | 0 |  | there is definitely some mistakes in the definition, the code seems to be Y91333B22R (the 0 added was a common mistake, I don't know where it comes from but we see it often). the drawing number is not registered on our side and doesn't match our standard. for the connection, the M20; 6mm and DIN could discribe a D43650 standard electrical connection (the diameter of cable entry fits 6mm and the thread of the cable gland is M20) but the mismatch is clear in the range, B22 is 10bar and they mention 24bar. I suggest to try to clarify with the customer to be sure to offer the right replacement.
TCK65117 | User | Service Request | DEMANDE HONEYWELL SEPARATEUR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 12:07 | 2025-07-09 09:25 | 1 | 0 |  | Salut Vero, tu as raison, les photos des montages ne correspondent pas et je dirais même plus que sur les photos tu as 2 montages différents (un standard et un différentiel avec 2 capteurs de marques différentes, ou au moins dépoques différentes). je peux t'orienter vers un D920 mais avec juste des photos je ne ferais pas beaucoup plus. je pense que pour de l'historique le mieux est que tu contacte Jérôme, il a une mémoire géniale et si c'est des pièces de rechange je suis sur qu'il pourra t'aider. si ce n'en est pas, c'est les datasheets et les réquisitions du projet qu'il nous faudra pour définir les produits en partant de la demande du client. Best, Nicolas
TCK65095 | User | Service Request | TR: DDP URGENT D910-L100A960 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 10:45 | 2025-07-07 11:40 | 0 | 0 |  | on peut faire des montages directs, mais ce n'est pas un D910 sur la photo, ça ressemble plus a un D920. ton choix de transmetteur honeywell est un différentiel alors que sur la photo ce n'en est pas un, je pense que Nathalie a raison, il y a des choses a revoir ensemble pour cette demande, appel moi et on regardera la demande et on définira le produit ensemble.
TCK65090 | User | Service Request | RE: DDP URGENT D910-L100A960 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-07 10:18 | 2025-07-07 11:41 | 0 | 0 |  | Cloturé sur TCK65095
TCK65025 | User | Service Request | MX7 MCD7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-04 14:45 | 2025-07-04 17:21 | 1 | 0 |  | j'en trouve très peu mais le composant pour le raccord 1/4 NPT sur un MX7 existe, il faudra que tu fasse une demande de prix avec ta quantité et un exemple de code approchant (exemple: 11271838)
TCK65004 | User | Service Request | BAIT / Customer Blackston / alternative Temperatura Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-04 11:28 | 2025-07-04 11:37 | 0 | 0 |  | the first picture (order n676740) is this: 11048665 TSF 100/231.1B2 0...200C KR=2M/738C the second (order n211092/20): 11048665 TSF 100/231.1B2 0...200C KR=2M/738C
TCK64925 | User | Service Request | dossier sur les EF6 71507-02 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-03 16:48 | 2025-07-04 10:13 | 1 | 0 |  | je n'ai pas les specs à respecter de chez Alstom et le numéro de plan est générique, ce n'est pas un produit spécifique. ça ne me donne pas grand chose pour comparer. après, de toute façon c'est soit le PP10R soit le PP20R qui peuvent remplacer mais ils seront de toute façon différents de l'EF6. Pour moi il est compliqué de remplacer l'EF6 sur des projets en cours car les produits sont passés par une qualification selon les specs du projet, il faudra tout refaire si vous voulez changer le produit. si c'est pour de nouveaux projets, il faut commencer le process de qualification des transmetteurs avec Belfort, ça avait pris plus d'un an pour les EF6
TCK64912 | User | Service Request | TR: Aanvraag - Eurotrade offre pour Dennis Groosjohan | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-03 15:24 | 2025-07-04 10:34 | 1 | 0 |  | a part que ça ne se monte pas, je pense que la solution MEX + AKPL de rik est très bien, moins de raccords intermédiaire, tout 316L et robuste. Order code: MEX5-D61.K33 MEX old Code / SAP code: + Order code: AKPL30N601 AKPL old Code / SAP code:
TCK64894 | User | Service Request | TR: demande pour thermowell | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-03 13:43 | 2025-07-04 09:35 | 1 | 0 |  | la c'est une décision de chez vous, nous techniquement on peut surement faire, on bosse avec des spécialistes que ce soit Traimeca ou Metron. mais si c'est pour 1 pièce le prix va être indécent surtout sur une pièce spécifique comme ça et il faudra passer par une validation de plan, une création de code SAP etc... si le client n'est pas prêt à payer 5000 euros pour sa pièce (et je dirais dans les 12 semaines de délai) ça ne sert a rien d'aller plus loin. si c'est pour 20 pièces ça vaudra déjà beaucoup plus le coup.
TCK64853 | User | Service Request | TR: Y92-2.3G.B79.R/0832 process connection problem | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-03 09:03 | 2025-07-04 09:28 | 1 | 0 |  | j'ai envoyé sans faire exprès, j'ai regardé les plans de connexion et on a changé la membrane mais pas de cote extérieures en cause pour le montage. ça doit se monter chez le client. su'il fasse bien attention que sa connexion soit propre, surtout sur de la peinture, une petite couche pourrait empêcher le montage. Si il est sur que ça ne vient pas de la, il peut vérifier les cotes de montage (plan sur la datasheet) et si elles ne collent pas a la réalité il peux faire une claim.
TCK64750 | User | Service Request | replacement ¦ E923 ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-02 07:04 | 2025-07-02 09:21 | 0 | 0 |  | the number was not complete on the picture, the order number was 302300/20 and the product was: 10269981 _E923 G1/2 1JT NBR 0+10BARR I suggest to replace with: 10301879 Y9233GB22R this is the same but we only kept the ATEX version to cover for both ATEX and non-ATEX applications.
TCK64735 | User | Service Request | TR: Quote request mano-meters Baumer 20640468 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-01 14:55 | 2025-07-01 16:15 | 0 | 0 |  | Répondu sur thread mail
TCK64734 | User | Service Request | Old Gauge Identification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-01 14:49 | 2025-07-04 10:37 | 1 | 0 |  | I had to ask for help on old SAP research for this but here is what product it was: EA4-AP3.B38R 10270530 code is blocked but we can still do it, you can request the reopening of the code for a replacement order.
TCK64684 | User | Service Request | DP100-hyg | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-01 08:58 | 2025-07-01 10:03 | 0 | 0 |  | yes those products are available.
TCK64682 | User | Service Request | TR: unidentified gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-07-01 08:54 | 2025-07-01 10:01 | 0 | 0 |  | c'est un mano aéro pour messier dowty, mais ils sont confidentiels et sous copyright, ils doivent sadresser a messier pour le remplacement.
TCK64365 | User | Service Request | TR: Consultation SNCF : n° SOURC-0000312460 Phase 1 Version 1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-27 10:01 | 2025-06-27 11:27 | 0 | 0 |  | Jai retrouve ces informations Ref sap 10263784 [Basic data text, data migration/2009-05-28] THERMOSTAT ELECTRONIQUE 1/2NPT -40+125C SEUIL 0+105C A LA HAUSSE PLAN 71364-01 NDO le 17/09/2021 : statut IB suite mail AGABR : "Malheureusement nous navons plus les équipements pour réaliser ces pièces."
TCK64348 | User | Service Request | tve-naval-larrieu | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-27 09:02 | 2025-06-27 09:33 | 0 | 0 |  | Voici: [Image]
TCK64301 | User | Service Request | Barflex à calibrer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-26 10:04 | 2025-06-26 10:13 | 0 | 0 |  | si c'est pour les renvoyer pour calibration, le récent on peut le faire par contre sur l'ancien design on aura plus les composants si quelque chose est a remplacer. sinon la procédure pour les renvois pour calibration est comme une claim mais tu précise que c'est un retour pour calibration.
TCK64278 | User | Service Request | DRO 63/411.111/145 / Christine | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-26 08:34 | 2025-06-26 09:42 | 0 | 0 |  | I would prefer to offer a MEX3 (less expensive, more standard and full stainless): MEX3D20B35 or MEX3D20B35/0133
TCK64240 | User | Service Request | TR: JP. Alcaraz Electrónica ¦ Manometro Full Inox DN100 Class 0,6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-25 14:59 | 2025-06-25 16:54 | 0 | 0 |  | je n'ai rien trouvé de ressemblant les M20 possibles sont ci-dessous (celui de gauche est celui que tu as avec le digit "9") a mon avis si ils veulent cette connection sans adaptateur, il faudra qu'ils nous donnent la norme ou le dessin et il faudra un processus de creation et de validation. [Image]
TCK64219 | User | Service Request | Restrictor Screw for MPG6(R) / GTS | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-25 12:02 | 2025-06-26 17:20 | 1 | 0 |  | yes, the restrictor screw is in the same material as the connection (1.4404 stainless steel).
TCK64202 | User | Service Request | MFT7 Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-25 11:08 | 2025-06-25 11:46 | 0 | 0 |  | V20 is an option for external 0 adjustment but it comes from the subsupplier's codification system, I'm not sure how we will reflect it in the final codification internally.
TCK64145 | User | Service Request | TR: Thermomètre bimétallique | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-24 16:05 | 2025-06-24 18:03 | 0 | 0 |  | nos TBI ne sont pas sensés être mis sous pression, les bonnes pratiques veulent qu'ils soit installé dans un thermowell. si ce n'est pas le cas, le client prendra la responsabilité. nous estimons que sans flux, et sur de l'eau, les raccords coulissant bien installés / serrés tiennent un maximum de 20bar.
TCK64139 | User | Service Request | TR: item inquiry : see attached picture. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-24 15:29 | 2025-06-25 09:35 | 1 | 0 |  | Nathalie, c'est toujours compliqué dêtre sur sans avoir la photo du cadran du dessous mais j'ai eu la demande par la Belgique avec exactement la même photo d'application et ils on envoyé ensuite des photos des cadrans et on avait déterminé que cétait ce code. c'est aussi un des seuls codes vendu à Singapour: 10262332 si ils veulent une confirmation formelle, il faudra une photo des 2 cadrans (celui du dessus et du dessous et la mesure du diamètre de passage du câble de pression (en bas a gauche de l'appareil).
TCK64137 | User | Service Request | Return of PCBA (10165792) to BECH without information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Peter, Jérôme [pjer] |  | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Peter, Jérôme [pjer] | 2025-06-24 15:14 | 2025-06-24 15:44 | 0 | 0 |  | Emailadresse war falsch
TCK64073 | User | Service Request | Y922 ¦ product identification ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-24 08:06 | 2025-06-24 09:45 | 0 | 0 |  | it's this code and it's released: Y92-2.3G.B28.R/0832 10301865 enclosed is the old datasheet (a new one is in the making). and there is no special connection.
TCK64066 | User | Service Request | PBCN / APS | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-24 06:44 | 2025-06-24 09:38 | 1 | 0 |  | we stopped the PBCn and didn't replace the capacitive cell sensor. I would suggest to offer this as a replacement: Y91-2.B.3.B22.R/2164
TCK64064 | User | Service Request | Order for MPG6-D60.D31/0771 / GTS | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-24 05:57 | 2025-06-24 10:16 | 0 | 0 |  | the MPG6 PP is not available anymore and is replaced by the new design in PBT (especially for new code creation). there is a debate on the digit to be used for kPa dial on the MPG6 (as the design of MPG is ANSI, and not EN) but with X or D this is the dial you'll get: [Image]
TCK64042 | User | Service Request | Durée de vie batterie Barflex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-23 14:43 | 2025-06-23 18:06 | 1 | 0 |  | je pourrais vérifier mais pour moi le système du Barflex est plutôt simple tu as la conso de lécran qui ne bouge pas vraiment sur le long terme et la consommation de la carte qui dépend de la pression mesurée (et pour le calcul on a surement pris une mesure soit moyenne soit haute (12mA ou 20mA de signal constant). on pourrait refaire une estimation mais ça ne changera pas les 150h en 1an, tu gagnera 10 ou 20 % au max et on ne peut pas changer la pile sans perdre la qualification ATEX. pour 1 an en fonctionnement continu sur batterie je ne sais pas si il trouvera une solution.
TCK64033 | User | Service Request | Sample dial | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-23 13:37 | 2025-06-24 09:05 | 0 | 0 |  | thee you go: [Image]
TCK63950 | User | Service Request | séparateur D294 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-20 15:21 | 2025-06-20 16:02 | 0 | 0 |  | je proposerais ce code si tu dois remplacer aussi la partie a souder et que c'est du 3" (11202177) Sinon: TED-5.3.3.B19.R.D295-2L20.000.01/0592/0607/2037/9010 (si tu le monte sur la bottom part en place).
TCK63915 | User | Service Request | MIX7-D50-B22/Q1070 /Alex | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-20 10:47 | 2025-06-20 11:25 | 0 | 0 |  | Q1070 wa the former code for the EN837 calibration certificate. here is the current code: [Image]
TCK63907 | User | Service Request | 25-05-ATALIAN FACTURE 250507983.pdf | Service Desk > Non IT | Low |  |  | Closed | Lambert, Thierry [tla] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-20 10:03 | 2025-06-20 10:45 | 0 | 0 |  | Salut Thierry, je pense qu'il y a une erreur d'aiguillage, j'ai reçu ta facture pour Atalian de l'IT en suisse. Best, Nicolas
TCK63901 | User | Service Request | Enquiry for TXA7 / Alex | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-20 09:00 | 2025-06-20 09:21 | 0 | 0 |  | the replacement for the TXA7 is the TSF or TSS 160 (the question of the capillary is not clear as it is mentionned but I don't see a length anywhere).
TCK63890 | User | Service Request | IBAIT / Customer Precision Fluid : controllo/riparazione pressostato BAUMER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-20 07:07 | 2025-06-20 09:26 | 0 | 0 |  | the product you have on the picture is a PFSX (basically a TED5 that is SIL2 and ATEX). the replacement will be the TED5 but without those approvals: TED525B16R
TCK63870 | User | Service Request | TR: FIRM BUYING REQUIREMENT 1045 GAUGE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-19 14:45 | 2025-06-19 14:56 | 0 | 0 |  | il veut ce code mais en C24 au lieu de C26: 11188937 MPG6-DX0.C26.XBR:FCBM12A8T600/0693++ il faut faire une demande a parker pour la bride pour 1piece , le mano ne changera pas de prix. et ne pas oublier les certificats demandés pour le mano et pour la bride: "CERTIIFCATION REQUIRED: FOR PRESSURE GAUGES: Q008/Q002/Q1201/Q003/Q017 FOR FLANGE: Q002/Q001/Q1201/Q003/Q017"
TCK63857 | User | Service Request | Gauge Dial Issue | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-19 12:33 | 2025-06-19 14:28 | 0 | 0 |  | I would agree with you but I can't act on claims. the best way to go is to open te claim, and say you agree to just received new dials and change it locally (I think they would accept as there is no technical issue and the dials are easy to replace (on dry gauges)
TCK63844 | User | Service Request | TR: item inquiry : ZDPHN-158-CHD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-19 10:28 | 2025-06-19 11:59 | 0 | 0 |  | CHD n'existe pas mais on voit souvent des erreurs de recopie sur les codes de ZP (beaucoup de numérisation on étaient faites avec des logiciels de reconnaissance de caractères pas très fiables). pour moi ils veulent ce code, c'est le seul existant en CH: ZDPHN-158CHM/SETP 11039198
TCK63807 | User | Service Request | vdav- etude EVOLLUTECH --AIR LIQUIE DECARBONATION DUNKERQUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-18 15:30 | 2025-06-27 14:32 | 0 | 1 |  | voici le step qui te manquait, les questions sur la qualif des joints sont toujours en attente, je clos le ticket et je te mets dans la boucles de la question au fournisseur.
TCK63769 | User | Service Request | Imprécision PBMN-23.B22R.DT2-2L22.0.000.04/0765 (11119093). | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-18 10:11 | 2025-06-18 11:41 | 0 | 0 |  | la précision du montage 11119093 est de 0.5%
TCK63768 | User | Service Request | TR: GB2S POC 10001 - Consultation manomètres KF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-18 09:52 | 2025-06-30 16:56 | 1 | 1 |  | on a rien de déjà fait, j'ai cherché et j'ai fait chercher les experts séparateurs et on a rien trouvé donc ce sera de toute façon une SPR pour création.
TCK63764 | User | Service Request | FW: Material inquiry / 피앤아이코리아 | Service Desk > Non IT | Low |  |  | Closed | Lee, Na Young [lena] | Mao, Yihua [maoy] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-18 09:26 | 2025-06-18 10:07 | 0 | 0 |  | you can request the reopening of this code to this adress (they handle code creation), you can precise I validated the code is ok for reopening. FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org>
TCK63737 | User | Service Request | TR: Demande de devis | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-17 16:35 | 2025-06-17 16:47 | 0 | 0 |  | je vois juste que le BAT est a valider a la commande, il semble que pour l'instant on ait juste fait une offre et pas encore de commande. tu peux demander le BAT, valider et il sera déblocable.
TCK63711 | User | Service Request | Etalonnage Barflex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-17 12:06 | 2025-06-17 14:49 | 1 | 0 |  | oui c'est possible et ça ne s'appelle pas claim mais c'est un processus similaire, mais demande a Nathalie, elle saura mieux et surtout elle pourra te donner les détails de la procédure et le tarif.
TCK63683 | User | Service Request | BAIT /customer CE.VI.V. / richiesta PBSN - 1.3.B.R.A1.14.99.2.1.CO01 4 pezzi | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-17 07:55 | 2025-06-17 10:13 | 1 | 0 |  | this is a specific product dedicated to the customer Bucher Vaslin, and the connection is a specific flush diaphragm on a G1/2 + cone connection.
TCK63682 | User | Service Request | BAIT / customer Valvitalia / UL approved pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-17 07:28 | 2025-06-17 10:05 | 1 | 0 |  | I'm sorry the answer will be easy on this, we didn't do any UL qualification on pressure switches (RP2 and RPP). but Baumer has options in electronic pressure sensors with switches and UL (PFMN for example).
TCK63662 | User | Service Request | TR: Phenolic cases bezel ring | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 16:01 | 2025-06-17 17:21 | 1 | 0 |  | Handled by FENI
TCK63629 | User | Service Request | Need details - DRO63 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 12:06 | 2025-06-16 15:17 | 1 | 0 |  | I'm pretty sure it's this one but the connection size is too confirm, I can't be sure from a picture: 10342669 DRO 63/811.133.545/1095_1359_9712_0720+
TCK63627 | User | Service Request | BAIT / Customer MIRECO S.r.l. / I: Richiesta | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 12:00 | 2025-06-16 15:24 | 1 | 0 |  | the product on the pictures is a 11150031 PBSN-1.2.B29.R.A1.53.02.2.3 it could be replaced by a Y9 (analogic option): Y91-3.2.9.B29.R/2160 or maybe by a PP something but I'm not an expert for those products.
TCK63623 | User | Service Request | ARA06 ARA10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 11:47 | 2025-06-16 15:48 | 1 | 0 |  | I cannot share detailled information about the internal construction the only thing I have in terms of information is the older datasheet that has a bit of complementary information.
TCK63610 | User | Service Request | PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 10:31 | 2025-06-16 15:01 | 1 | 0 |  | Le voici ci-joint.
TCK63602 | User | Service Request | TED avec label oxygène | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-16 10:00 | 2025-06-16 15:00 | 1 | 0 |  | On peut le fabriquer mais les TED sont bloqués et on évite d'en créer de nouveau normalement. mais demande a Mario, il peut accepter et après de notre coté ce sera bon. attention par contre, l'objet de ta demande parle doxygène mais je ne vois pas d'option de dégraissage dans ton code, il faut peut être la rajouter.
TCK63510 | User | Service Request | replacement ¦ EDX 520/324.221/756 ¦ AKP_INQ_20250613_0833 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-13 08:51 | 2025-06-13 12:18 | 1 | 0 |  | I would offer this as a replacement, the range is to check because 756 doesn't exist, and the connection is to be confirmed but should be the right one: PBMN-24B11RA14444201010 [Image]
TCK63483 | User | Service Request | Technical assistance - switches | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-12 17:49 | 2025-06-13 09:51 | 1 | 0 |  | Discussed, will be handled via teams calibration session with Pinnacle
TCK63406 | User | Service Request | Special Dial | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-11 16:54 | 2025-06-11 17:24 | 0 | 0 |  | I think it's really close to the Bar Bourdon (A) design. I sent you an example.
TCK63398 | User | Service Request | WG: EA5-AF3.B2DR/0587 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Marolf, Mario [marom] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-11 15:48 | 2025-06-11 16:15 | 0 | 0 |  | it's a different cell in the 0.1% (EA5A) and there is no gascket between the cell and the body. so what is stated on the datasheet is right
TCK63318 | User | Service Request | D820 316SS NACE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-10 18:17 | 2025-06-11 09:12 | 0 | 0 |  | our supply of 316L for diaphragms and flanges is NACE (marqued on the side of the flanges) and you can request a certificate (that will be for the wetted parts by adding a Q017 to your order).
TCK63311 | User | Service Request | TR: 10304986 PLAQUE D'ECHELLE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-10 16:36 | 2025-06-10 18:02 | 0 | 0 |  | la plaque déchelle est la plaque métallique avec les réglages de setp et décart sur un pressostat. elles ont toute le même code sap et sont repérées par le code de léchelle (ici ) pour les différentier: 10304986 PLAQUE D'ECHELLE "https://servicedesk.baumernet.org/wm/" plaque 234
TCK63284 | User | Service Request | Assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-10 15:05 | 2025-06-11 09:54 | 1 | 0 |  | H = 94mm on the small one. H = 168mm on the long one.
TCK63278 | User | Service Request | PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-10 14:25 | 2025-06-10 18:11 | 1 | 0 |  | Sure, there you go:
TCK63251 | User | Service Request | TR: 10269759 - E913 GEC PL 71059-01 REP 101 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-10 11:31 | 2025-06-10 12:09 | 0 | 0 |  | ci-joint le plan de référence du E913 GEC, on ne le fait plus mais on peut proposer un standard avec des connexion et une fonctionnalité similaire: Y91339B20R il faudra lui proposer avec la datasheet pour qu'il puisse comparer.
TCK63188 | User | Service Request | RDN-8A3.234 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-09 16:23 | 2025-06-10 10:34 | 1 | 0 |  | here is the SAP code for the plate but you'll need to precise the code of the full RDN with it as the different plates don't have an individual code and they need to identify the one you need (the 234): 10304986 PLAQUE D'ECHELLE I let you check with sales for the rest.
TCK63185 | User | Service Request | FW: 20615322 ¦ 20540330 ¦ zapytanie 44229/25/ZD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-09 14:32 | 2025-06-10 10:18 | 1 | 0 |  | I was about to tell you it was OK from the top of my head but luckily I checked, your customer needs MR5-1D30.B22: [Image]
TCK63180 | User | Service Request | weld-in sleeve for TBI ¦ zapytanie czujnik temp | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-09 11:59 | 2025-06-10 10:10 | 1 | 0 |  | Bellow are 2 options that would fit your request, directly a thermowell to weld on piping for the models without sliding connection: [Image] or a welded accessory and a threaded thermowell to go in: [Image] With below if thermometer without sliding connection: [Image] with below if sliding G1/2 connection on themometer: AGWR0230L2D100
TCK63102 | User | Service Request | TR: taminco FW: [EXTERNAL] RE: Manometer onderdruk | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-06 08:45 | 2025-06-06 09:36 | 0 | 0 |  | Oui c'est sur que ça ne passera pas au P2F mais il faut surtout vérifier dans SAP si il est créé sur un MEX8 et c'est le cas: et en cas de doute nhésitez pas a demander le cadran a la sérigraphie. [Image]
TCK63081 | User | Service Request | TED sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-05 14:16 | 2025-06-05 14:39 | 0 | 0 |  | le code du client est tronqué, il manque des infos, mais je t'ai fait un tableau des codes anciens correspondant (en D294) et des nouveaux codes existants pour remplacement. j'ai ajouté tous les codes options et mis en commentaires les particularités de chaque montage pour que tu te repère facilement. ci-joint aussi la datasheet du D295 et tous les montages que j'ai trouvés étaient remplis en LRS1. Old Designation Old SAP Comment _TED-5.3.3.B19.R.D294012L202/LRS1_0591/0592/2037
TCK63034 | User | Service Request | TR: the regulation conformity of ship recycling and HAZARDOUS MATERIALS for MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-05 08:36 | 2025-06-05 10:07 | 0 | 0 |  | Bonjour Nathalie, voici la lettre de conformité matière globale Baumer qui mentionne tous les standards qu'on respecte globalement. en plus le RoHS est mentionné sur le certificat de conformité aux normes CE ci-joint. J'ajoute ci dessous un tableau des matières contenues dans le produit composant par composant (les noms des composants sont cachés pour des raisons de confidentialité). PN Name Qty Unit Weight (kg) Material 10308592 Best, Nicolas
TCK63016 | User | Service Request | TR: MFT Information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-04 16:55 | 2025-06-04 18:04 | 0 | 0 |  | the MAR5 no longer exists but we decided to offer a possible replacement with a MEX including the main options of a MAR. hereafter a slide of how to convert a MAR5 into a MEX. as an example, the first of your list would be a MEX5A30A20/0751/0151/0412/0670/A009
TCK63012 | User | Service Request | TR: Baumer - AB: 2363788 / Testmaterial | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-04 16:16 | 2025-06-05 11:04 | 1 | 0 |  | le marquage des voies n'avait pas été prévu par Philippe, mais il est possible. sur un MFE42 ça ressemble à ça et on peut l'ajouter a la demande via Philippe (et ajout d'une sérigraphie peut être). [Image] Par contre comme tu peux le voir la place est limité donc le marquage avec les mots allemands long ne sera pas possible, il faut nous donner quelque chose de court. - Labelling: BZ (white) BSII (red) - Labelling: Hauptleitung (white) Hauptluftbehälter (red)
TCK63010 | User | Service Request | TR: ) Quotation request / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-04 16:13 | 2025-06-04 18:02 | 0 | 0 |  | the MAR5 no longer exists but we decided to offer a possible replacement with a MEX including the main options of a MAR. hereafter a slide of how to convert a MAR5 into a MEX. as an example, the first of your list would be a MEX5A30A20/0751/0151/0412/0670/A009
TCK62960 | User | Service Request | Info | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-04 12:12 | 2025-06-04 14:27 | 1 | 0 |  | we still do both the 1005 and the DRCG101, but the issue is the 1005 only exists in M12*100 female connection and we never did it on the DRCG. if there is only one unit it makes no sense, it'll be too expensive but for some units, we could do something. also it'll not be 3A as we would need to make the complete certification for it. if you think it's ok without 3A and there is some quantity, please send the request for the complete code to both me ans sales and they'll make the price for the special DRCG based on the quantity of the first order.
TCK62956 | User | Service Request | TR: CTX Pressure Transmitter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Nathalie, Billault [nabi] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-04 11:32 | 2025-06-06 16:35 | 1 | 0 |  | Nathalie, on a pas d'adaptateur compatible en stock ou en standard. il faudrait faire les plans, le créer et le commander. ce sera plus rapide et moins cher de commander les CTX avec la bonne connexion. Best, Nicolas
TCK62917 | User | Service Request | Problems with NeoAPI and GigE cams | Service Desk > Non IT | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-03 21:44 | 2025-06-04 09:24 | 0 | 0 |  | seems not to be a request for baumer and certainly not for france or Bourdon (cameras). answered via email with clarifications and a contact in brazil.
TCK62911 | User | Service Request | 80115096 --- GRCH--- ID:10103840 urgent !! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-03 16:14 | 2025-06-04 10:46 | 1 | 0 |  | Demande comme ça, et n'oublie pas le logo en .jpeg de bonne qualité: MEX5-D30.X77/0751_0228_9750_9714 (Bar/noir; psi/vert; C/bleu; F/rouge (linéaires))
TCK62905 | User | Service Request | TR: Demande information AD00003377287 / 6343 13 010 – 05 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-03 14:57 | 2025-06-03 17:28 | 1 | 0 |  | Je peux déjà clarifier un peu: A3 est le format du plan, pas un indice. le numéro du plan est 6343 13 010-05 et il n'a jamais été changé et n'a donc pas d'indice. a titre d'exemple ci-joint, l'autre plan mentionné par le client est le 6343 13 010-00-d qui est un autre numéro de plan et qui est a l'indice d. si il y a d'autres questions après ça tu peux me les envoyer sur le support.
TCK62898 | User | Service Request | TR: ddp 2503ISNA105 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-03 13:51 | 2025-06-03 17:50 | 1 | 0 |  | je ne pense pas mais on pourrait proposer un robinet schneider qui l'a. je pense que c'est le plus simple.
TCK62885 | User | Service Request | GRCH - Projet 25-0430 / demande info manifold | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-03 11:58 | 2025-06-03 16:29 | 1 | 0 |  | vu que lélément limitant est le PTFE et en prenant en compte ce qu'on dit sur d'autres produits avec cette matière comme joint et en prenant un coef de securité, -30C c'est OK.
TCK62809 | User | Service Request | BAES --- Guardian Industries Navarra --- SDI 439 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-02 15:02 | 2025-06-03 16:31 | 1 | 0 |  | A adresser par email stp je dois consulter Jorg pour ça.
TCK62782 | User | Service Request | TR: coordonnées Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-02 12:07 | 2025-06-03 17:55 | 1 | 0 |  | A gérer hors tickets, renvois moi la demande sur la boite mail. on ne pourra pas remplacer mais on est peut être équipé pour restaurer. A voir si c'est accepté en interne et il me faut plus d'infos. demande à ton collègue si je peux leurs passer un coup de fil pour avoir plus d'infos stp.
TCK62769 | User | Service Request | replacement ¦ E913-53-B22A ¦ Zapytanie ofertowe MH/2219/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-02 11:20 | 2025-06-02 14:23 | 1 | 0 |  | sure, the E91 was not really phased out, it's just that we only kept the ATEX version (Y91) to cover all applications as our standard. so you can offer a: Y91353B22A
TCK62748 | User | Service Request | TR: DDP 175961 BAUMER URGENT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-02 09:22 | 2025-06-02 10:39 | 0 | 0 |  | ce code a juste été bloqué car il manque la longueur et le setp mais sinon tu peux le garder tel quel: RTNAA.415.23XE6/1990/SETP (X = P = 450mm) setp à confirmer. après c'est un raccord coulissant donc je conseille quand même de rester si possible sur 400mm ou 600mm et de prendre un code plus simple et moins cher pour le client.
TCK62742 | User | Service Request | MEP5-460.F72.AORPB502/SETP.ARIN2606.D631-2L20.217.G20.09/0678_0693_0763_9030_9941 / GTS | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-06-02 08:55 | 2025-06-02 10:31 | 0 | 0 |  | I just made 2 changes but it's fine like below: MEP5-460.F72.AORPB502/SETP(0.6bar).ARIN2606.D631-2N20.217.G20.09/0678_0693_0763_9030_9941
TCK62690 | User | Service Request | FW: PCC Rokita ¦ Pressure gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-30 13:44 | 2025-06-02 10:23 | 1 | 0 |  | I do not handle anything related to prices, but our sales dept will be happy to help you here. please send the request as is, they might require you to fill an SPR (special price/product request) and we well send the requirement to our supplier for the different connections. you can add the extract from the datasheet for a clearer request and also confirm you want Male/Female M20x1,5:
TCK62664 | User | Service Request | Enquiry on Pressure Gauge, Pressure Switch and Temperature Switch /Nyein Winsys | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-29 03:58 | 2025-06-02 17:45 | 1 | 0 |  | there is no real reference from us and it's very old, it'll have to be defined again. I suggest you direct them to our website and you can give them already the datasheet of the TSFE/TSSE (it's most probably similar to the TXE that might have been an electric contact version of the TXA) for the pressure switches, considering they called switches a TSS/FE, I think we might offer a MR5/MG5 and we can say we're unclear about the switches vs contact definition in their request and sy they can check both on our website. a switch main function is cuttin power, it doesn't deliver measurment but switch at a defined setp and can cut from high to very low current/voltage. a contact gauge main function is measuring and has a dial to read the actual pressure or temperature and contacts that can cut mid/low power. I hope it'll help you guide them to the right products.
TCK62593 | User | Service Request | DANC compliance information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-28 09:52 | 2025-05-28 14:26 | 1 | 0 |  | enclosed is the certificate for the BSE/TSE for our filling liquid LRS1 (and the MSDS of the fluid). I also attached the global Baumer statement on the most common material requirements. the rest of the parts are just pure stainless 316L. if you need anything else don't hesitate.
TCK62589 | User | Service Request | MEX3 For Refrigerant R449a / GUE | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-28 09:28 | 2025-05-28 13:51 | 1 | 0 |  | yes we can do this, the connection code is "I" so for example: MEX3BI0X2D/97xx (dual scale R449a in C) I don't have enough information to define the dial precisely, this is an example. [Image]
TCK62588 | User | Service Request | GRCH --- SUEZ -- cmp-036-rfq | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-28 09:21 | 2025-06-05 14:19 | 1 | 1 |  | ci-joint le fichier STEP du AMFD03N.
TCK62489 | User | Service Request | Regarding Accuracy Class 0.6 Option for MEX5 Series | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lee, Bryan [lb] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-27 10:31 | 2025-05-27 11:28 | 0 | 0 |  | the MEX5 was never available with both filling and special accuracy (the knife edge pointer for accurate reading paddles in the liquid and it put to much stress on the axis). We have both option 0.5% and 0.6% available on dry MEX5 but we advise the 0.6% one as 0.6% is an accuracy recognized by the EN norm and your gauge will remain conform and have the EN markings. and there is one code already created that might be what they ordered previously but it was with no specific accuracy (so 1%) the code already existing is: old code: MEX5A63B18-0201 new code: MEX5-A63.B18/0838 10280374 and if now they want the special accuracy it'll be without filling and I suggest 0.6%: MEX5-A60.B18/0838/0840
TCK62481 | User | Service Request | MPG6(R)-A62.X20.ASIP20N604/0718/9941 / Exxon | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-27 09:46 | 2025-05-27 10:52 | 0 | 0 |  | that's a hard question. all MPG6 new design have a pointer stopper at 6 o'clock (which is a good news for you I think because it's requested on your market). but as you request a specific dial, we will remove the existing and replace it by a blank one we will print and those don't have a pointer stopper. so if you don't add the option, I'm sure it'll not be added. so it's possible but make sure you keep the option 0718 in your request for code creation. in the future we'll try to make it standard on all.
TCK62477 | User | Service Request | manufacture standard of BOURDON‘s diff switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-27 09:24 | 2025-05-27 09:56 | 0 | 0 |  | do you want to know the norms we comply to ? please find enclosed our most requested sdandards regarding those products. (RoHS, PED, ATEX and other CE norms)
TCK62471 | User | Service Request | GRCH - RP2-CHAUDRON- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-27 09:01 | 2025-05-27 09:49 | 0 | 0 |  | question pas évidente, ce n'est pas mentionné sur le plan mais j'ai regardé sur un plan similaire qu'on fait chez le même usineur et le Ra demandé est 6.3 et 3.2 micromètres. [Image] par contre le fluide sera aussi en contact avec la membrane en PTFE et pour ça je n'aurais vraiment pas de Ra.
TCK62447 | User | Service Request | TR: BAIT / Customer: Goriziane / Marine certification Y91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-26 17:18 | 2025-05-26 17:33 | 0 | 0 |  | we had it before on the Y913 which could fit your request, but we stopped renewing it because there was too few request for it. though the product technically didn't change, I have an official paper stating it (enclosed, you can fill it with date and customer information). some of our marine customer still use it under internal derogation as it was certified in the past.
TCK62386 | User | Service Request | SERIC-- BITUME - BFI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-26 10:33 | 2025-05-26 11:24 | 0 | 0 |  | je pense que tu es sur la bonne piste, MEX5 + D418 avec purge ou évent et purge, propose la vanne qui va avec les normes (Métrique si le reste est en métrique et impérial sinon). et le mano avec montage mural si ils veulent des trous de fixation.
TCK62349 | User | Service Request | Thermowell ATEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-23 22:26 | 2025-05-26 17:21 | 1 | 0 |  | the thermowell is purely a static, mechanical component made of stainless steel with no inherent energy source or moving parts that could generate heat, sparks, or other ignition sources, it would fall outside the scope of the ATEX Directive and do not require any ATEX approval to be installed even in ATEX zones.
TCK62344 | User | Service Request | ATEX Certificate | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-23 15:23 | 2025-05-23 16:44 | 0 | 0 |  | Handled via teams
TCK62322 | User | Service Request | Battery information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-23 11:35 | 2025-05-23 11:54 | 0 | 0 |  | here is an extract of the manual with the instruction for changing the battery and the model. and the picture with the reference on the RS website. 1 year with good network and hybride mode is the battery lasting time with data interval of 10-15min approx. [Image] [Image]
TCK62312 | User | Service Request | calogero -- projet RIVARD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-23 10:21 | 2025-05-23 12:19 | 0 | 0 |  | Vu par teams, solution ARID Y91 MEP5 attente info client.
TCK62299 | User | Service Request | product identification ¦ S401.59-03033 ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-23 09:32 | 2025-05-26 14:47 | 0 | 0 |  | Anna, after answer from BTIN there was no hidden option, so what we suggested is a good replacement. so no further action here.
TCK62266 | User | Service Request | UV protection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-22 16:57 | 2025-05-23 09:30 | 1 | 0 |  | we test for UV resistance dials and paintings, but also the window gasket (to make sure it doesn't color the filling fluid in the gauge) and the EPDM plug gasket: here are the 2 components of the plug (the first one ensures the tightness). 11079536 EVENT EPDM GRADE 5 (SUIVANT SPECIF2005) 11077356 BOUCHON EVENT ROUGE GPN 350 I 5 (solid polyéthylène low density (PE-LD)) and here an extract of the SP specifying the tests done on it to validate it is suitable for our application (including UV tests). [Image]
TCK62187 | User | Service Request | MSDS Sheet LRS7 and BH10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-22 10:49 | 2025-05-22 12:01 | 0 | 0 |  | here is the LRS7 and BH10 MSDS. the BH10 is only in french, if you need it in english I have to make a request to the supplier (i didn't found it on their website).
TCK62115 | User | Service Request | TR: Lead 30276376 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-21 16:08 | 2025-05-21 16:39 | 0 | 0 |  | Je te conseille de répondre comme ceci: pour remplacer le manomètre en photo: MG5.4D33B22 + DANC012L2A + 01 (double contact NF/NO a confirmer et détails invisibles a confirmer). pour le manomètre différentiel -1...5bar avec clamp 38 dont il parle, ça fait peu d'info mais on peut faire une propal: M33 4A3B20H + AKPLBF5303 + DANC022L2A + 01 (DANC pas possible en DN38, mini DN51/50) -1...5 est plus une étendue relative (compound) que différentiel, il y a peut être un quiproco de ce coté.
TCK62107 | User | Service Request | question sur un ED522 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-21 14:17 | 2025-05-21 15:38 | 0 | 0 |  | je te joint la datasheet du produit, mais il a été phase out il y a plus de 20ans, même dans les archives papier je ne trouverais rien et cétait le système informatique Macpac. on a gardé des sauvegardes de l'ancien système jusquà l'année dernière on y a plus accès, désolé. de toute façon pour le certif de calibration, il ne vaut plus rien après 20ans. par contre pour la matière je suis désolé. selon l'étendue de mesure, il pourra au besoin être remplacé par le nouveau Y92 des que la datasheet et le config seront prêts (ou par un Y91 + DANC en attendant) ou un PBMH/PP20H si la précision est très importante.
TCK61985 | User | Service Request | RDN-8A3.234 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-20 10:13 | 2025-05-20 15:11 | 1 | 0 |  | I have a similar one here and for what I can see, they cut the connector to change it. The fact it doesnt work could come from here, please find below the wiring for your product. at a minimum, you should be connected on white and either blue or red (and grounded). [Image]
TCK61984 | User | Service Request | TR: Product Info | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-20 09:53 | 2025-05-20 11:56 | 0 | 0 |  | this is a mounting done by ESP with our Y91 on one of their seals. I would offer this as a replacement: Y913-33-H26R + DANC012L2A.01 but Mexico is under the responsability of ESP.
TCK61946 | User | Service Request | BAIT / Customer: Fluid processing / SEP535-CC1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-19 16:56 | 2025-05-19 17:48 | 0 | 0 |  | Perfect, thanks for the update, I close the ticket then. Best, Nicolas
TCK61920 | User | Service Request | replacement ¦ ED 517/314.231 / 155 ¦ RECTUS, zapytanie ofertowe | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-19 14:25 | 2025-05-19 14:59 | 0 | 0 |  | please find enclosed the old datasheet (sorry we just have old paper versions in french). so it was a transmitter, Relative 400bar, DIN 43650, 4...20mA, G1/4. you could replace by a Y913.23.B38R
TCK61902 | User | Service Request | 转发: 转发: 回复: 回复: 关于电接点压力表的选型咨询 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-19 11:53 | 2025-05-19 14:54 | 1 | 0 |  | I'll not have some in production right now (we do very few of those). but he can already get most information from the complete datasheets (dimensions, pictures and all caracteristics). please find the datasheet enclosed in the ticket. [Image]
TCK61893 | User | Service Request | Contact Gauge and In lIne Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-19 10:31 | 2025-05-20 09:09 | 1 | 0 |  | Please go with one of those, they are both OK with ATEX and silicon free for Paint application. MG5-1F2K.B26/4710/0831 + 1545 + LRS7 MG5-1F3K.B26/4710/0831 + 1542 + LRS7
TCK61864 | User | Service Request | thread of DIN connector of Y91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-19 07:50 | 2025-05-19 09:47 | 0 | 0 |  | The definition of PG9 cable gland includes the thread, please see below the table of the PG9 thread definition and the complete documentation. [Image]
TCK61831 | User | Service Request | TR: Sensor Support | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-16 14:50 | 2025-05-16 16:34 | 0 | 0 |  | pour l'instant tu peux lui envoyer l'ancienne mais uniquement pour référence. Mario vient de me confirmer qu'il essayait de finir la nouvelle ce mois ci.
TCK61779 | User | Service Request | AW: Solicitud de cotización ¦ pressure gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-16 08:41 | 2025-05-16 09:17 | 0 | 0 |  | sure, the first one can be replaced by a simple MEX3 and the second by a MEX3 + DANC, they are both available and selectable in the P2F and website. I need to warn you that what they had (ADD model and AYD) are references from our factory in India.
TCK61744 | User | Service Request | sensor can measure pressure and temperature simutaneously | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 15:19 | 2025-05-15 17:19 | 0 | 0 |  | sorry, I don't see which product it could be, it's possible but I don't find it. and maybe we still have an option to offer but you need to contact sensor division for this, I don't know their products enough.
TCK61739 | User | Service Request | GOAY-MDJ-DEXIS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 14:36 | 2025-05-15 17:13 | 1 | 0 |  | en PN100 ca me parait bien, mais en PN10/40, ca veut dire que si la pression va en haut déchelle, le séparateur va lâcher (ou fuir). je conseillerais de prendre un Mano 40 bar si de toute façon leurs process ne résiste pas au dessus. ou un 60bar et un PN63 ou 100 si ils peuvent vraiment atteindre cette pression, l'important est de garder la cohérence entre les instruments et le piping.
TCK61735 | User | Service Request | CPX electrical connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 13:40 | 2025-05-15 17:08 | 1 | 0 |  | the code you gave me is in M12*5 (11189875). but you're right, it's unclear from the datasheet. for me the standard was M12 but better to check with me if you have a doubt.
TCK61720 | User | Service Request | Hydrogen Application | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 11:38 | 2025-05-15 11:55 | 0 | 0 |  | Handled over teams. Cone and thread connection will be advised with the link to Bourdon H2 page. otherwise 1/4NPT solution will be provided with warning that it's not the most suitable connection especially at 400bar.
TCK61715 | User | Service Request | [BHFR Material Data] Compléter opération de gamme article 10321286 | Service Desk > General | Low |  |  | Closed | Riou, Marc [rima] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 10:59 | 2025-05-15 12:02 | 0 | 0 |  | Resent to FUNC material data
TCK61607 | User | Service Request | MCX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-15 08:07 | 2025-05-15 17:50 | 1 | 0 |  | je vais t'envoyer une photo mais sur les MCX/F, au dessus de la Pmax de la capsule (400mbar dans ton cas) une vis clapet vient fermer l'entrée de la capsule (et tient jusqu'a 20x la pression). je t'envoie une photo du systeme par teams.
TCK61565 | User | Service Request | validation d'un montage sur DAEF14 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-14 11:50 | 2025-05-14 16:24 | 1 | 0 |  | comme discuté via teams, c'est validé avec un ASIPD au lieu du ZPX
TCK61495 | User | Service Request | RDN-8A3.234 Gaskets. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-13 15:54 | 2025-05-14 09:58 | 0 | 0 |  | Below the SAP code of the gasket they broke: 10302462 Sensing element gasket BUT, this means that their took appart the sensing element, which is not necessary in the calibration process and will act on the placement of all elements. after that, we cannot keep any warranty on this device and we would strongly reccomend that they stop the process. the product will not be covered by any safety norms from us and this needs to be clearly stated to the customer. if they only opened one, I would suggest to start a return process for the rest of the devices for calibration, please check with sales how to handle the next steps.
TCK61492 | User | Service Request | Y92-2.3G.B29R/0201 - RENAULT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-13 15:30 | 2025-05-13 16:58 | 0 | 0 |  | je ne peux pas retrouver ce qu'il y avait dans l'option 0201. il y a de bonne chances que ce soit l'option 0832 "Label peinture" mais je ne peux pas être sur. autrement, pas de soucis, il faut juste créer le code: Y92-2.3G.B29.R/0832 (exemple pour création: Y9223GB28R-0201 Y92-2.3G.B28.R/0832 10301865)
TCK61489 | User | Service Request | VDAV -- AZ-industrie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-13 15:15 | 2025-05-13 16:52 | 0 | 0 |  | Tu as raison, je rajouterais un point, quand on est sur des appareils sensible, attention à bien les vérifier droits. et bien tapoter le boitier entre chaque point pour assurer la stabilité du point.
TCK61458 | User | Service Request | Y91 sur séparateur avec capillaire | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-13 11:00 | 2025-05-13 14:11 | 1 | 0 |  | on pourra faire soit max 400bar et 400C avec le montage ci-dessous, soit 600bar et 200C. pour un réglage de léchelle a 420bar, tu peux remplacer le Y91 par un PBMN mais les limites de montages restent les mêmes, elles viennent du séparateur.
TCK61411 | User | Service Request | TR: Urgent - Certificats | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-12 16:38 | 2025-05-12 17:58 | 0 | 0 |  | si on a bien un certificat ci-joint, et tous les certificats LCIE sont sur le site LCIE si besoin https://www.iecex-certs.com/#/deliverables/CERT/12803/view
TCK60995 | User | Service Request | _DR 63/452.111 0..6BAR 0201 (10319886) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Gitau, Kenjie [gik] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-07 09:21 | 2025-05-07 14:04 | 1 | 0 |  | I would sugget this to replace: MEX3G20B20/0769 and a 10bar range or an accessory if overpressure protection is necessary. the main difference is this one is full stanless steel with 316L parts in contact with the fluid.
TCK60888 | User | Service Request | PDHAEN ---BABE -- gsk PO 0002217467 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-06 09:43 | 2025-05-06 15:19 | 1 | 0 |  | on ne fera plus les DPDx, mais on peut les remplacer par un MFT5 à contacts. par contre le contact 5033 je ne vois pas comment il peut exister, un seul prend déjà 4 fils a mon avis c'est plus une erreur ou ils ont considéré 2 contacts car il y a 2 points de contact.[Image] il faudra faire la demande aux a fantinelli via l'ADV mais tu peux le demander comme ça: Order code: MFT5-A00.B15N.32/5003 MFT old Code / SAP code: il peut même y avoir un model un peu moins cher avec une P statique max moins haute. Attention l'écartement des connections de l'ancien DPDE n'était pas standard (70mm) le nouvel écartement sera standard (54mm).
TCK60840 | User | Service Request | Urgent | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Aernouts, Senne [asen] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-05 16:17 | 2025-05-05 17:55 | 0 | 0 |  | You need to send this request to a commercial contact, I can only deal with technical questions.
TCK60770 | User | Service Request | GRCH-SCHNEIDER-GASPAR-- 31032025AG00 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-05 09:59 | 2025-05-05 10:16 | 0 | 0 |  | les montages avec des produits extérieurs sont possible mais il y a plein de trucs a vérifier et un process a mettre en place, je pense que le meilleur moyen de bien valider tout est de passer par une SPR.
TCK60734 | User | Service Request | Y913 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Soskic, Orhan [soo] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-02 14:44 | 2025-05-05 10:11 | 1 | 0 |  | Sorry, we stopped renewing the marine approval as there was not enough request. Please find enclosed the official letter stating this and that the product conception didn't change. some of our marine customer used it to keep working with the Y91 during some time before finding a replacement.
TCK60730 | User | Service Request | D043 3.1 traceability | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-02 11:23 | 2025-05-05 11:36 | 1 | 0 |  | I like what they are doing here, it's pretty neat. we handle traceability a bit differently, our seals are not sold alone so we make traceability of the full system with the gauge barcode. the barcode is linked to the manufacturing order and at eah production step, the material certificates of the components are scanned and automatically added to the documents linked to the product. but I'll still tell quality that the PN marking is a good idea, so is a direct link from the seal to the material certs.
TCK60726 | User | Service Request | MPG6-A60.X18(R)/9709 / Kah Fei | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-02 10:04 | 2025-05-06 10:33 | 1 | 0 |  | the drawing on the MPG6 R for 0...250kPa ANSI will have to be created.
TCK60715 | User | Service Request | MEP3-420.B25/0764 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-01 10:59 | 2025-05-05 09:17 | 1 | 0 |  | go with 0765, and it'll be EN. sorry for this, I noticed a couple month ago that the different O2 options were not understood the same way by codification and P2F, we'll be working on this with Jorg.
TCK60714 | User | Service Request | RDN8 Pressure Switches Issue | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-05-01 10:19 | 2025-05-05 09:19 | 1 | 0 |  | send further communication via email (there might be large files and some follow up, it'll be easier without the ticketing on this issue).
TCK60668 | User | Service Request | TBI080-212.141.54T possible ? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-30 17:14 | 2025-05-05 10:25 | 0 | 0 |  | it's OK, Marc confirmed it's possible: TBI080-212.141.54T [Image]
TCK60617 | User | Service Request | TR: inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-30 10:19 | 2025-04-30 12:12 | 0 | 0 |  | Philippe a raison pour le SETP, ce qu'ils demandent n'est pas possible, ci-dessous le minimum de reglage possible et le tableau ainsi que le calcul pour le deadband mini: ZPN201CHM6/SETP (SETP: 0.1bar falling; DB: 0.165bar mini) Calcul: Range = -1 to 2.5bar = 3.5bar so 10% = 0.35bar Usable range (10 to 90%) = -0.65 to 2.15bar (so the SETP of 0.1 is within) deadband mini is 0.15bar at -0.65 (10%) and 0.2bar at 2.15 (90%) at your setp (0.1 which is at 32% of the scale), the min deadband is ((0.2-0.15)*0.32)+0.15 = 0.166bar your minimum deadband at this setpoint is 0.166bar [Image]
TCK60615 | User | Service Request | MEX5 With Range -760 mmHg…1.5 Kg/Cm2. / Theerasak | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-30 09:57 | 2025-04-30 11:54 | 0 | 0 |  | yes it exists and there is several codes open with this dial, here is the most recent code example: 11204243
TCK60587 | User | Service Request | small screw on the back of TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 17:54 | 2025-04-30 09:07 | 0 | 0 |  | this is a breathing hole so that the cell is in contact with the atmospheric pressure. it's basically a filter letting the air in but keeping the particles out.
TCK60579 | User | Service Request | MEX3 class 1.0 accuracy 转发: 转发：Fw:Re: 回复: 压力表替代咨询 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 16:41 | 2025-04-29 16:55 | 0 | 0 |  | yes, with the option 0841, a MEX3 can be class 1 but I don't see why you would want this, your customer only require class 2.5... otherwise, here is an example of code with the option: MEX3-D20.E26/0841
TCK60569 | User | Service Request | LAA-LARGILLIERE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 15:22 | 2025-04-29 16:51 | 0 | 0 |  | pour la compatibilité je laisse le soin au client de le dire mais 10% dhydrogène ça ne devrait pas avoir d'impact, ça me parait simplement être du gaz naturel pré-traitement (ou entre la désulfurisation et létape de séparation des éléments lourds). je ne vois pas de contre-indication majeure de mon coté avec les composants du gaz.
TCK60568 | User | Service Request | GOAVEC-GRCH-250012 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 15:14 | 2025-04-29 15:17 | 0 | 0 |  | tu as raison, pas de norme ASME, on propose du clamp norme EN.
TCK60557 | User | Service Request | Kindly confirm | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 14:19 | 2025-04-29 15:15 | 0 | 0 |  | this is a request for spare parts for a KMMC order from 2012. please check this request with Jerome and Shafi, I think it might be handled locally with them.
TCK60535 | User | Service Request | MEX5-D60.X73/9709_0751_0764 | Service Desk > Non IT | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 10:33 | 2025-04-29 11:21 | 0 | 0 |  | sorry, only 2 codes exist in MEX5 X73 and it's not the scale you need. it'll need to be created.
TCK60519 | User | Service Request | General information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-29 08:58 | 2025-04-29 09:12 | 0 | 0 |  | I would need to know on which product you need it to give you the right certificate but here is were to find it: RohS is on the CE certificate: [Image] And REACH is notified in Baumer's global material compliance declaration at the group level (enclosed). along with a lot of other material regulations.
TCK60487 | User | Service Request | TR: Rfq transmitter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-28 16:15 | 2025-04-28 16:33 | 0 | 0 |  | ci-dessous le teams que j'ai envoyé a Angel, on a une solution en PFMN + D82x mais vu la demande c'est plus pour du transmetteur type Yoko avec un D9xx et si c'est ca il faudra passer par un SPR. I got your request for a transmitter on diapgragm seal, if the Hart communication is not mandatory we could offer a PFMN mounted on a D82x (to be confirmed depending on the parts in contacts and diaphragm material) but I feel like this application is for a field transmitter type yokogawa and then we don't offer full mounting, for long term business we can validate a mounting but it needs to go through an SPR.
TCK60480 | User | Service Request | MFT + D821 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-28 15:25 | 2025-04-29 14:50 | 0 | 0 |  | issue with customer request (static 250bar / class 150) to be reviewed with the customer
TCK60434 | User | Service Request | FW: 11135834 ¦ MEP5-461.B18.D621-2N60.213.G6002/0191++ ¦ Zapytanie ofertowe // BAUMER Gauge MEP [10.04] | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-28 10:12 | 2025-04-28 11:19 | 0 | 0 |  | I guess it is this one: MEP5-460.B18.D621-2N60.213.G6001/0191_++ 11206151
TCK60429 | User | Service Request | [BHFR Material Data] Modification code article 10320544 | Service Desk > General | Low |  |  | Closed | Riou, Marc [rima] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2025-04-28 09:33 | 2025-08-20 11:20 | 0 | 1 |  | Not relevant to technical support.
TCK60419 | User | Service Request | CONFIRMATION REFERENCE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-28 08:55 | 2025-04-28 09:36 | 0 | 0 |  | ça match pour moi avec les numéros de plans, ci-dessous les résultats en recherchant par ancien n d'articles: 53077-01-00 JOINT GRAPHITE EN 10307171 53077-02-00 JOINT GRAPHITE PUR EN 10307172
TCK60393 | User | Service Request | plastic diaphragm seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-27 12:30 | 2025-04-28 09:33 | 0 | 0 |  | it is still available, but the limits for mounting are slightly different: [Image]
TCK60383 | User | Service Request | MPG6DX1B19+FC_M12A8T600-X | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-25 17:57 | 2025-04-28 09:28 | 0 | 0 |  | I found a list of codes using this flange but no existing combination of B19 and filled. you can have it created by refering to a recent exemple (11225551) MPG6-DX0.122.XBR:FCBM12A8T600/0816_9711+ 11168457 MPG6-DX0.176.XBR:FCBM12A8T600/0191++ 11168458 MPG6-DX0.B19.XBR:FCBM12A8T600/0191++ 11184675 MPG6-DX0.C26.XBR:FCBM12A8T600/0693++ 11188937 MPG6-DX0.C76.XBR:FCBM12A8T600/0816_9711+ 11193012 MPG6-DX1.B22.XBR:FCBM12A8T600/C014++ 11147248 MPG6-DX1.B26.XBR:FCBM12A8T600/C014++ 11225551
TCK60378 | User | Service Request | WG: TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-25 16:00 | 2025-04-25 16:16 | 0 | 0 |  | here are our standard CE certificate and the 3A certificate. the oil filling is FDA, I also included the FDA certificate for it.
TCK60375 | User | Service Request | AW: TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Ammann, Christof [amc] |  | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Ammann, Christof [amc] | 2025-04-25 15:42 | 2025-04-25 15:56 | 0 | 0 |  | Hallo Miteinander Ich wollte das Mail eigentlich an Bourdon weiterleiten. Komischweise wurde daraus ein Ticket erstellt. Es kann gelsöcht werden. Gruss Christof
TCK60330 | User | Service Request | 3D Step Drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-25 10:11 | 2025-04-25 14:56 | 1 | 0 |  | There is the 3D drawing.
TCK60328 | User | Service Request | E913 avec connecteur spécifique | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-25 09:58 | 2025-04-25 16:09 | 1 | 0 |  | il faudrait vous coordonner, la demande a été faite à Céline ce matin, vous pouvez toujours proposer ce produit, la connexion est une fourniture client donc a la commande il faut qu'ils pensent a nous envoyer les raccords.
TCK60296 | User | Service Request | ZDPN121CGM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-24 17:04 | 2025-04-24 17:21 | 0 | 0 |  | you're perfectly right in your approach. you defined your safe zone for settings (6.8 to 45.2mbar). you just need now to be sure both set and reset point are inside this zone. if you set at 10mbar rising, you'll close 0.4mbar lower (deadband) which is 9.6mbar and still in your setting zone so everything is fine.
TCK60240 | User | Service Request | TR: RFQ MEX5/AKPL/1005 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-24 08:58 | 2025-04-24 10:07 | 0 | 0 |  | the requested mounting is not possible at the moment. the 1005 only exists with an M12*1 instrument connection which is not available on the AKPL. Also with the very small diaphragm, 9m of capillary might absorb the movement of the diaphragm too much. For 1 unit this should be refused. For big volumes, a study can be done but I would suggest still a direct mounting of a Y91 and a cable and display.
TCK60208 | User | Service Request | TED sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-23 15:53 | 2025-04-24 10:38 | 1 | 0 |  | le code que tu as est un flasque sup de DT2, il est utilisé dans de nombreux montages avec des brides d'adaptation mais en gros j'ai trouvé 2 codes qui pourraient potentiellement matcher et il y a des chances que ce soit le premier je pense. ci-joint le plan du flasque de DT2 et de la bride a confirmer avec le client quand même, notamment avec le code ou plus d'infos sur le TED (marquage).
TCK60084 | User | Service Request | D043 Questions | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-22 12:47 | 2025-04-22 17:08 | 1 | 0 |  | the D043 are indian products we optimised with BTIN production and R1D before testing it and validating. but it is low cost stock products so this means that making specific options or markings is not the goal here. they are all ordered with material certificates but that's the only available option. as it's an indian product made in BTIN, it is possible to make changes (they have the production tools to make marking or machine a different connection) but it needs to be negociated through an SPR and will only make sense for big qty and if we keep stock somewhere. otherwise, the DT will be the right option, there is more connection available (and we could add G3/8 F for lower qty than with D04) and labels or tag plate with the heat number is possible via options 9941 or 9911 (engraving would be slightly more complicated). [Image]
TCK59953 | User | Service Request | Confirm the model code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-18 10:57 | 2025-04-22 10:45 | 1 | 0 |  | I don't find the exact one you ask for in the system, so either it's very old and was lost in the transfer between old and new SAP or there is a wrong digit in the customer's code. anyway, here is the list of what I can find with RP, range 205 and option 0201: RPPAAA3205-0201 _RPPA SI STD G1/2 3,5+25 BAR CONSIGNE EN 10297263 RPPAGA3205-0201 _RPPA SAM STD G1/2 3,5+25 BAR CONSIGNE EN 10297330 RPPE7B6205-0201 RPPE-7B6.205/SETP EN 10297418 RPPE7F6205-0201 RPPE-7F6.205/SETP EN 10297441 RPPE7G6205-0201 RPPE-7G6.205/SETP EN 10297450 RPPE8B6205-0201 RPPE-8B6.205/SETP EN 10297467 RPPNAA3205-0201 RPPN-7A3.205/0765 EN 10297508 RPPNBA3205-0201 RPPN-7B3.205/SETP/2281 EN 10297605 RPPNBA6205-0201 RPPN-7B6.205/SETP EN 10297615 RPPNCA3205-0201 RPPN-7C3.205/SETPS EN 10297638 RPPNGA3205-0201 RPPN-7G3.205/SETP EN 10297747 RPPNGA6205-0201 RPPN-7G6.205/SETP EN 10297762 RPPNNA6205-0201 _RPPNNA6205/SETP EN 10297793 RPPXAA3205-0201 RPPN-8A3.205/SETP EN 10297831 RPPXGA6205-0201 RPPN-8G6.205/SETP EN 10297847 RPPY7C3205-0201 RPPY-7C3.205/SETP EN 10297901 RPPY7C6205-0201 RPPY-7C6.205/SETP/2261 EN 11013704 RPPY7M6205-0201 RPPY-7M6.205/SETP EN 10297933 RPPY7S3205-0201 RPPY-7S3.205/SETP EN 10297944 RPPY7S6205-0201 _RPPY-7S6.205/SETP EN 10297948
TCK59878 | User | Service Request | need confirmation for model code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-16 18:10 | 2025-04-17 11:44 | 1 | 0 |  | here is the first code you requested: 11206151 MEP5-460.B18.D621-2N60.213.G6001/0191_++ the second is not existing (B26 with US5A is not existing, only in psi) but here is the closest example of an existing one with the details of the US5A options. [Image] the US label I don't think we do it anymore otherwise you can request the same options.
TCK59876 | User | Service Request | Technical approval | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-16 16:49 | 2025-04-16 17:40 | 0 | 0 |  | in terms of compatibility of mounting it is fine but there is a couple comments: 0751 is not necessary, it's standard on MQ7. B11J is possible but you'll have 5% accuracy and a reduced scale (probably 180) 4m is not standard, but it's possible (F 4.5m is standard and will be available in configuration, 4m will be sent to R&D for confirmation). Thanks for the request, I detected that on differential mounting I still use the scale to define the Pmax in the seals navigator, I need to change this. MQ7-831.B11J.D820.G4L.2229.G2001/0679_0078 [Image]
TCK59733 | User | Service Request | TotalEnergies-ZEM3-EF-4551538106-F13-YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-15 09:58 | 2025-04-15 10:47 | 0 | 0 |  | Salut Véronique, c'est écrit dans la datasheet: [Image] Par contre je proposerais du LRS8, le EJX310A est un absolu. Best, Nicolas
TCK59725 | User | Service Request | COMPATIBILITE SEPARATEURS YOKO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-15 09:10 | 2025-04-15 10:44 | 0 | 0 |  | pour moi pas besoin de Low volume sur le premier ce sont des flasques standards et le premier est un relatif donc je mettrais du LRS9 si rien de spécial n'est précisé. Pour EJX430A-JASWG-919EB/KU22/D3/C2/CA : D912-LD0.3Q960.6L33Q0M111 (11273101 par exemple) Pour EJX310A-JAS5G-919EB/KU22/D3/C2/CA : D912-LD0.9Q860.6L33Q0M111
TCK59655 | User | Service Request | EA4AP9B2DR (10270588) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Gitau, Kenjie [gik] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-14 10:04 | 2025-04-14 17:30 | 1 | 0 |  | I think you wanted to send an email to FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> (for code creation) the code you need is closed but can be reopen as is as it's for replacement, you can just send the request for opening the code in your sales unit to the above adress.
TCK59646 | User | Service Request | AROE-HENDSOLDT-- MEP3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-14 09:21 | 2025-04-14 17:26 | 1 | 0 |  | comme discuté au téléphone, pour l'instant la meilleure approche globale est de tout proposer en MEP5 C&T. on bosse sur d'autres versions et sur le MEP3 600 et 1000bar en parallèle mais a plus long terme. si ils ne veulent pas de diamètre 100, revenir vers moi et on étudiera leur demande plus spécifiquement.
TCK59588 | User | Service Request | à propos des RDE8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-11 12:16 | 2025-04-11 15:43 | 1 | 0 |  | sous la vitre (à gauche) la plaque de réglage du setp et de l'hysteresys. et a droite la plaque signalétique.
TCK59555 | User | Service Request | DEFINITION D912 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-11 09:11 | 2025-04-11 11:38 | 0 | 0 |  | je suis avec Christophe, a 50mbar abs en usage seul le LRS8 marchera et il ne tient pas 400C, on a pas de solution, désolé.
TCK59551 | User | Service Request | TR: Création BAFR // Client : Naval // Cde Clt : 7027054 // TSF100-231.1B0.65T.20/738C_9003_1000 // Position 10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-11 08:39 | 2025-04-11 09:02 | 0 | 0 |  | je pense que c'est une erreur d'envoi, tu est au support technique a Vendôme.
TCK59533 | User | Service Request | WG: Customer request: productInfo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Spricigo, Roberto [spro] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-10 15:33 | 2025-04-10 17:00 | 0 | 0 |  | no specific precaution for the transport it's a 400bar tube so the fluctuation of atmospheric pressure cannot damage it. You need to respect the temperature limits of -20...70C. Before strating it once arrive at destination, please make sure you open the top plug according to the sticker to balance the atmospheric pressure inside and outside the case. it's needed to avoid chocs as much as possible because the outside pressure will lower and th oil will push on the triplex window. it is safe but if something bump on the glass it could break more easily than usual due to the lower external pressure.
TCK59512 | User | Service Request | Confirm for DPG with diaphragm seal along with capillary is possible | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-10 11:10 | 2025-04-10 15:03 | 1 | 0 |  | here is the code with the full details of the one I think you need: 11227202 MX7-831.X16J.AKPL.DT3-2L90.0.960.01/++++ MX7 DIFFERENTIAL PRESSURE GAUGE BELLOWS ALL STAINL. STEEL DIFFERENTIAL PRESSURE, DN=150MM FILLABLE CASE 8 BOTTOM CONNECTION, PIPE CLAMP 2" CASE AND BEZEL IN 316L 3 CONNECTION G 1/2" 1 LIQUID FILLED BH1 X16 SENSING ELEMENT 0+1,6 BAR J STATIC PRESSURE 0+16 BAR 9712 DOUBLE SCALE: QUADRATIC SCALE 0M3/H = 0% = 0 MBAR 2000M3/H = 100% = 1216 MBAR 9711 TECHNICAL INSCRIPTION MADE IN FRANCE ON DIAL 0678 MICRO ADJUSTABLE POINTER AKPL CAPILLARY B STAINLESS STEEL DIA 2.5 + STAINLESS STEEL PROTECTION F5 ORIENTABLE SCREW G 1/2 3 CONNECTION G 1/2 01 LENGTH 1M DT THREADED DIAPHRAGM SEAL 3 LOW PRESSURE RANGE (D3XX) 2 TOP HOUSING INOX 316L (SS) L G 1/2 FEMALE 9 MONEL 400 (MO) DIAPHRAGM 0 DIAPHRAGM WITHOUT COATING 0 WITHOUT CLEANING RING 9 BOTTOM HOUSING MONEL 400 (MO) 6 THREAD 1/2 NPT 0 NO COATING 01 LRS1 (FOR TEMP.-15..150C)VASELINE
TCK59510 | User | Service Request | BYO-GRCH - RAPT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-10 10:42 | 2025-04-10 17:00 | 1 | 0 |  | Solved via teams meeting.
TCK59457 | User | Service Request | EA4AP9B2DR (10270588) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Gitau, Kenjie [gik] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-09 16:48 | 2025-04-09 17:28 | 0 | 0 |  | I'm sorry, I guess you wanted to send this email to FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> they handle code creation fr BHFR products.
TCK59397 | User | Service Request | DEFINITION D912 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-09 09:12 | 2025-04-10 16:50 | 0 | 0 |  | peux tu demander si le vide (Tour sous vide) descend sous les -500mbar R ? si ce n'est pas le cas, le montage sera OK en LRS9, sinon reviens vers nous.
TCK59375 | User | Service Request | Barflex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-08 15:27 | 2025-04-08 15:49 | 0 | 0 |  | voici le certifoicat CE des barflex.
TCK59345 | User | Service Request | Ex II2GDc T6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-08 12:43 | 2025-04-08 14:56 | 0 | 0 |  | this is the ATEX option (0078) that call for a marking. and the product on the picture is a 11137244 MEX3-D20.B15/0751_0078. but the marking on the back of the case was not made by us, I think it's a copy of the marking on the dial we do (maybe to be sure it was visible from all sides).
TCK59331 | User | Service Request | GRCH --- NOVONORDISK | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-08 11:07 | 2025-04-08 12:19 | 0 | 0 |  | bizarrement, sur le LRS11 on a jamais mis de plage de température car je pense que cétait une gestion et une imposition client) mais j'ai la FDS (ci-jointe).
TCK59251 | User | Service Request | I: BAIT / Customer: Loccioni / I: HPC Project - ATEX Certificate 21004596/OP) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-07 11:53 | 2025-04-07 15:17 | 1 | 0 |  | the LCIE othe other notified body do not issue certificates for ATEX, the manufacturer makes all the tests and verification, issue a complete file with the result and send it to LCIE for storage, they just give a certificate of reception of the file and an authorization to make the markings (they don't read or check the file). so the real ATEX certificate legally is the CoC we issue that is the one you sent to the customer. but if it can help, here is the LCIE 18 AR 015 (certificate of reception of our file).
TCK59229 | User | Service Request | EMP--- BOILLOT -- MDJ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-07 09:14 | 2025-04-07 09:36 | 0 | 0 |  | ce ne sont pas des MFE, (ARMANO) mais je pense qu'on peut proposer ca: MFE41_20.04.JB3/0420 MFE42_20.20.JB2/0420 MFE42_20.20.JB2/0420 Attention, il faudra surement ajouter une ou 2 options de cadran (pour la marque rouge, les noms des circuits et le n client), mais ca te donne une base sur quoi partir.
TCK59126 | User | Service Request | DEPRESSION MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-04 08:24 | 2025-04-04 09:14 | 0 | 0 |  | j'ai envie de dire que tant que le client appelle ça dépression je suis content et on tient mais je vais essayer dêtre plus précis. notre tube le plus faible est utilisé pour faire le -1...0bar et on assure une précision de 1% ce qui veut dire que notre tube résiste avec le coefficient de sécurité 10-3bar absolue. c'est la seule valeur exacte avec une résistance sure a 100% mais nous savons que nos manos sont utilisés dans de nombreuses applications industrielles qui peuvent supporter jusquà 10-6bar abs (d'habitude c'est ce que je donne comme règle générale pour notre tenue au vide, ça correspond au vide grossier).
TCK59045 | User | Service Request | VDAV -- examesure | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-03 11:08 | 2025-04-03 16:25 | 1 | 0 |  | c'est surement un MFE52, ça doit être remplaçable, et c'est probablement du 3/8 (qui est standard sur ces produits). il faudrait une meilleure photo et si possible avec étiquette si il y a, ça nous donnera peut être un code. et ce ne sera pas à l'identique car des choses sur la carte et les LED ont été changées (validé ferro toujours et approuvé par Alstom et la SNCF). sinon on peut redéfinir le produit a partir de la datasheet du MFE.
TCK58996 | User | Service Request | sensing element of ZDPWN 158 CHM/1930 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-02 16:19 | 2025-04-03 16:28 | 1 | 0 |  | as you can see on what you sent me, the materials (316L + Viton) are applicable to both 15x and 16x in the table so there is no point of changing for this reason. But the 16x have a different more robust designso it could help a bit with the tighness and resistance, but all materials are the same.
TCK58969 | User | Service Request | ATEX declaration for MEX3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-02 13:10 | 2025-04-02 14:28 | 0 | 0 |  | completely agree with you, all our conformity document are legally correct but are cruelly missing details to makes things clear. I have not much to offer on this as the certificates of conformity are handled at the group level but I'll relay the information to Jörg. in the meantime, let me send you the certificate stating that we are authorized to do the marking under our responsability following the reception of our technical homologation test report and the picture of the marking on a gauge (that's what's making our marking legal and the details). I'll push at the quality dept for more detailled CoC. [Image]
TCK58966 | User | Service Request | Vacuum protection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-02 12:35 | 2025-04-02 13:53 | 0 | 0 |  | no gauge in the worl would resist -1bar (absolute vaccum). But all our tubes are design to ensure the resistance and accuracy in vaccum which match rough vaccum conditions. so we can ensure the weakest tube will resist 10-3 bar (0.001bar absolute) and a 10bar tube will resist better. this will ensure safety for all accidental vaccum (you don't create high quality vaccum by accident). the issue that we could havr if they use a 0...10bar MEX3 for example (that has a pin at 0) is that the pointer will force on the stopper and might be damaged. in this case, a product without stopper (MEX5) or a -1...9bar scale will be suitable solutions.
TCK58954 | User | Service Request | Liability insurance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-02 11:02 | 2025-04-02 13:45 | 1 | 0 |  | please find enclosed the group liability insurance applicable for all our products.
TCK58867 | User | Service Request | BAIT ¦ Customer: WATER SYSTEMS ¦ Membrane pressure gauges ¦ MEX5 broken | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Melini, Alessandro [alme] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 17:17 | 2025-04-02 09:48 | 0 | 0 |  | they must be carefull because if it happened on both gauges quickly it means the solution is not adapted to the situation. the issue might come from various reasons, I would need a deeper understanding of their system to confirm but there is a few things I can say first. the usual suspects here would be: overpressure, choc, vibration or pulsation (or several together). and I see a pump right next to the gauge, which usually means pulsation, vibrations and sometimes chocs from water-hammers at start (the placement of the diaphragm relatively to potential water hammers is important to check, it might require the addition of an elbow befor the connection). I would suggest to share this with the customer and adapt our offer for the solution.
TCK58861 | User | Service Request | ADV-AROE--vacuomètre PC064193 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 16:36 | 2025-04-02 09:13 | 1 | 0 |  | je ne pense pas qu'on fera pour 44 pièces, en plus avec du non linéaire et non symétrique, je ne sais même pas quel tube et mouvements pourraient êtres utilisés. j'ai bien peur que ce soit mieux de décliner.
TCK58852 | User | Service Request | 3D drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Stojanovic, Zoran [stoz] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 14:50 | 2025-04-02 10:02 | 1 | 0 |  | I'm sorry this 3D drawing doesn't exist yet, but you can order it from the sales (technical document). the 2D dimensions are also available on the datasheet and if you need to check something specific, I can get the specific parts from production and check physically.
TCK58837 | User | Service Request | FW: REAR FLANGE FOR 150MM GAUGE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 13:15 | 2025-04-02 17:18 | 0 | 0 |  | After checking with Jorg, he aproved we can send those part to Elcon and they'll take the responsability from here. here are the part you requested, the front flange is to be welded behinf the bezel on the case. Bezel 10308749 LUNETTE150INOX POLIE316L Window
TCK58815 | User | Service Request | CONFIRMATION REFERENCE Huile Chaude /// QUO-93673 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 11:21 | 2025-04-01 14:13 | 1 | 0 |  | pas de soucis pour moi, le montage est compatible, le capillaire seul fera baisser la température à environs 100C donc sans problème dans les tolérances de l'EJX , et le reste me semble respecter la RC INS 402.
TCK58784 | User | Service Request | [BHFR Material Data] Modification quantité dans un article | Service Desk > General | Low |  |  | Closed | Riou, Marc [rima] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2025-04-01 09:09 | 2025-08-20 11:21 | 0 | 1 |  | Not relevant to technical support
TCK58779 | User | Service Request | BYO --- SE6---GASPAR ---SCHNEIDER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-04-01 08:47 | 2025-04-02 10:54 | 1 | 0 |  | c'est OK pour nous, on a validé le montage avec Chloée et le séparateur apparaîtra en tant que XDSE + n de plan et mentionnera l'ancien type (D006)
TCK58747 | User | Service Request | Old gauge and Seal assembly | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-31 19:35 | 2025-04-01 10:55 | 1 | 0 |  | I don't find anything with the drawing number or part number but... the DT8 replaces the D108, + we know it's HC276 as you said and LRS1, I made you a conversion with a new code. It needs to be validated with the customer though. MEP5D61B20-316L+AORP382+ARIN+D108HCLRS1 MEP5461B22 + AORP382 + ARIN2606 + DT82N60066001 I have no view on potential options, and the minimum mounting range will be 10bar on this.
TCK58738 | User | Service Request | Option 0060 for MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-31 17:47 | 2025-04-01 14:50 | 1 | 0 |  | here is the code of the window of option 0060: 10309155 we just mounted today the last ones on stock but we receive some next week.
TCK58610 | User | Service Request | Datahseet | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Soskic, Orhan [soo] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-28 09:27 | 2025-03-28 15:13 | 1 | 0 |  | yes, it's this code in M18: T 08406.2030 VERSCHR. KOMPL. NG 6 M18X1,5 INOX 10318686
TCK58596 | User | Service Request | installation and mounting pieces | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-28 07:58 | 2025-03-28 08:59 | 0 | 0 |  | you can order it, here are the codes: 10296201 ZWI.STUCK G DIN 16281-6 MS N 05012.000 10310391 JOINT EN837 G1/2 6,2/17,5X2 PTFE N 01890.0106
TCK58531 | User | Service Request | SOGEBI --- VDAV --- 11119340 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 15:11 | 2025-03-27 17:27 | 1 | 0 |  | pas de nouvelle étude, par contre c'est un code de chez nous avec des options allemandes, on le codifiera plus comme ça, tu peux demander la réouverture / re-création en disant que c'est OK avec moi et ils vont créer le code proprement.
TCK58527 | User | Service Request | VS: [External] Spørgsmål/tilbud | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Soskic, Orhan [soo] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 14:51 | 2025-03-27 17:19 | 1 | 0 |  | there you go, the drawing is attached in the ticket.
TCK58523 | User | Service Request | TED-6.B.3.B33.R Article no.: 11260681 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 14:30 | 2025-03-27 17:17 | 1 | 0 |  | it seems like a good option from the requirement you have, either TED5 or 6 depending on the system and barrier used and he can set 2 separated setp or 1 + the hysteresys. I would send him the one you defined + the datasheet so he can adjust the exact code depending on his setup and preferences.
TCK58514 | User | Service Request | AD gauge datasheet | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 13:18 | 2025-03-27 17:11 | 1 | 0 |  | it'll be a MEX2 D20 xxx, but the code doesn't say the range.
TCK58483 | User | Service Request | MR5 Dual Scale | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 09:18 | 2025-03-27 09:58 | 0 | 0 |  | dual scale is possible but please make sure there is no further marking to add (we have very limited space on those). K33 is not existing yet so there will be a creation fee (we print the scale specifically for those dials).
TCK58477 | User | Service Request | MV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-27 08:05 | 2025-03-27 09:50 | 0 | 0 |  | on peut mais elle sera a créer (V22) car je ne la trouve sur aucun MV existants (y compris en X22). par contre en équivalent, le bar/psi existe sans surcoût (au lieu de kPa/psi)
TCK58431 | User | Service Request | MEX5 Option 0745 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-26 15:19 | 2025-03-26 17:02 | 0 | 0 |  | it's really not a common option, I have no idea why it appears on the website but I understand it's not in the P2F ;) anyway, this option is available, codes are open and it's onpen in role 1 un the configurator. and we have stock of the components. I don't see anything that can block, you can go ahead. just to know, is there any reason they don't want the polycarbonate one ?
TCK58419 | User | Service Request | AT.D.2.5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-26 12:12 | 2025-03-26 14:01 | 0 | 0 |  | Enclosed the datasheet from BTIN.
TCK58407 | User | Service Request | Enquiry For D2002L20-DSX200-LRS1 / Theersak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-26 10:19 | 2025-03-26 14:04 | 1 | 0 |  | DSX200 has no meaning for me and D200 could be different DTs, but based on the odds, it's probably a: DT2.2L20xxxxx01 please check for the details of the cleaning ring, bottom part and process connection with the customer.
TCK58189 | User | Service Request | MEX5 D30 B77 + D820-10L.42CE.B2001 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-23 14:34 | 2025-03-24 09:47 | 1 | 0 |  | sorry, the minimum for this mounting is 10bar (or -1...9). you'll have to check the D821 I think, we won't have flush options in DN25 for this pressure. [Image][Image]
TCK58186 | User | Service Request | MEX3D21B74+DANC01 DN38 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-22 07:19 | 2025-03-24 09:42 | 1 | 0 |  | the Seal Navigator contains the most up to date informations, you can go ahead and order it.
TCK58088 | User | Service Request | DIONISIO- CAF - MFE pressure Gauge ¦ AMLD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-20 13:28 | 2025-03-20 17:07 | 1 | 0 |  | pas de soucis, pose une réunion dans mon agenda, on verra comment on peut prendre le problème.
TCK58087 | User | Service Request | AROE-DELTA-AMFD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-20 12:13 | 2025-03-26 17:32 | 1 | 0 |  | cette option est bien dispo chez schneider parker etc... par contre je ne trouve rien de déjà créé, il va falloir faire la demande pour AMFD02N + option anti tamper head unit ci-dessous les options dispos chez schneider pour t'aider a être plus spécifique sur le moyen de protection: [Image]
TCK58025 | User | Service Request | EF6 work on altitude from 4500m to 5100m | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-19 14:34 | 2025-03-19 16:42 | 0 | 0 |  | the EF6 cell is linked to the atmospheric pressure through a ventilation hole. so at 5000m, the atmospheric pressure will be about 0.5bar so if what you measure is not linked to atmospheric (closed pressurized air system), the EF6 will signal 0.5bar more of difference between the atmospheric and the pressurized system. it'll work fine otherwise.
TCK57983 | User | Service Request | confirm model code for customized stem length - model RT2E | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-19 06:27 | 2025-03-19 09:31 | 0 | 0 |  | the basic interest of working with sliding connection is the standardization and adaptability. if your customer wants special length, it'll be more expensive, and longer to deliver. on top of this, he'll be trapped with us because he won't find those length easily with another manufacturer. it's a lot of drawbacks when a 400mm will do the job exactly just as well as a 385mm. we are not willing to do this if not absolutely necessary as it represent no advantages for us or for the customer. if after presenting those facts to the customer he still insist, I'll need the qty requested to make the request to the PM and we'll need to start authorization, price determination and code creation pro
TCK57938 | User | Service Request | WEP5 - BAES-DIONISIO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-18 10:53 | 2025-03-18 14:29 | 1 | 0 |  | voici les réponses que je peux te donner: 1.- For WEP5, please send the table with the battery duration according to frequency time transmission. We know the LoRa documentation mentionned up to 5 years and that with the 1minute min frequency, we count on a month but it's also very linked to the type of acknoledgment setting and the quality of the network (communication with the gateway). I'll work with mamour to give a more complete answer. 2.- Is there a future version of temperature with LoRa transmission. We are talking about new LoRa projects and a temperature one was mentionned, it's not past the M1 phase of concept yet. 3.- Are Baumers pressure and temperature gauge homologate by IPQ (Portuguese Institute of Quality)? https://www.ipq.pt/ We never had the requirement for this, our products are conform to all types of norms and requirement (mostly european) and are the validations were done by private or public european quality services (based in England, france, switzerland and germany for the most part). what would you want to have certified by the IPQ ?
TCK57927 | User | Service Request | 转发: ZPN202CGSM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-18 10:33 | 2025-03-18 14:58 | 0 | 0 |  | the code you have is old, the microswitch CGSM existed before 2000 and was replaced (including in the nuclear documentation SEPTEN). the CGSM was replaced by the switch CGM. the new code would be if I refer to your old code: ZPN202CGM.
TCK57818 | User | Service Request | Dual Scale for MCX7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-17 09:22 | 2025-03-17 14:58 | 1 | 0 |  | no problem, your code is possible but the dial isn't existing, you'll need to request the creation.
TCK57438 | User | Service Request | [BHFR material data] Changement de boitier dans la nomenclature de l'article 11095217 | Service Desk > General | Low |  |  | Closed | Riou, Marc [rima] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2025-03-14 16:21 | 2025-08-20 11:21 | 0 | 1 |  | Not relevant to technical support
TCK57434 | User | Service Request | TR: Demande d'achat de pièces détachées Ref trescal : 3955781 et 2649357 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 16:02 | 2025-03-14 16:28 | 0 | 0 |  | le capteur est de chez nous et on peut le remplacer, par contre leur raccord adaptateur est un adaptateur pour serrage manuel et il ne vient pas de chez nous. On utilise aussi ce type de raccords sur nos postes de calibrations mais ils ne seront pas compatibles.
TCK57429 | User | Service Request | RE: Certificat CE 2014/68/eu pour MEX/DANC | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 15:21 | 2025-03-17 16:30 | 1 | 0 |  | les DANC ne remplissent pas les conditions minimum de volume et de pression pour etre concernés par la DESP.
TCK57416 | User | Service Request | TR: consultation n° 34353871 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 12:22 | 2025-03-14 14:01 | 0 | 0 |  | je peux pas en tirer beaucoup plus que toi mais je commencerais par proposer un MS5 mini-maxi.
TCK57403 | User | Service Request | TR: BAIT / Customer: ELTECH / ED701R204R4A2510/9007/0080 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 10:33 | 2025-03-17 15:20 | 1 | 0 |  | le ED701 était un transmetteur de pression avec une version hygiénique. le problème est que le "R" qui désigne la connexion napparaît dans aucune doc, donc je suppose un remplacement par un PBMN flush ou hygiénique pourrait marcher mais il faudra définir le connexion. voici les caractéristiques principale du code ED701 et la datasheet pour référence: 4...20mA Connecteur DIN 43650, IP 65 0...5bar abs huile silicone
TCK57381 | User | Service Request | CAPILLAIRES | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 07:48 | 2025-03-17 14:52 | 1 | 0 |  | notre préconisation de rayon de courbure mini est de 75mm, c'est ce que nous respectons pour les tests et le transport. la préconisation trouvée par l'UNIE (<35mm a remplacer, <50mm ne pas toucher) semble raisonnable à appliquer sur des montages non conforme en place même si le remplacement reste la solution la plus sure. le montage en dessous des 75mm de rayon crée la possibilité d'une déformation plastique irréversible ce qui induit une fragilité intrinsèque qui dans léventualité d'un choc mécanique (ou de vibrations, ou de séisme par exemple) pourrait entraîner une rupture, la dissipation du gaz et donc un appareil hors d'usage. Bien sur, les risques avec la protection torsadée en inox sont réduits car il repartissent les efforts et empêchent de forcer en un seul point mais notre recommandation de rayon mini reste la même à 75mm pour une utilisation sure.
TCK57368 | User | Service Request | Bimetal Thermometer -70 to 300 deg.C / GUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-14 04:51 | 2025-03-17 15:35 | 1 | 0 |  | here is what I would offer, and if the range is important, you can make an SPR for a special range. we don't have antiparalax mirrors on those types and the external adjustment Order code: TBHI130-264.140.15T/0751_9003_0280 TBHI old Code / SAP code: + Order code: AGFB-0257N.1M210.T AGF old Code / SAP code: (18-14mm)
TCK57356 | User | Service Request | TBH100.212.152 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-13 15:43 | 2025-03-14 10:30 | 1 | 0 |  | just to make sure as your TBI code is not clear on this, you have an 8mm diameter stem, right ? if yes, I think the installation with the T 08416.0200 is fine and pretty standard, the thermowell doesn't usually go all the way through the pipe. what you need is to ensure the minimum immersion length of the TBI (you can find it on the datasheet) is immersed in the fluid. the only thing I would have a remark on is the case will be slightly separated from the connection but if there is no vibrations, then it's really not an issue.
TCK57308 | User | Service Request | Sensotec -BAES- MEX5 for NH3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-13 10:00 | 2025-03-13 15:40 | 1 | 0 |  | si c'est pour du spécifique et qu'il existe déjà en DR, je suggérerais d'utiliser directement les DR existant. sinon ça va être une galère à codifier.
TCK57278 | User | Service Request | BAIT - Water System - MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-12 16:53 | 2025-03-12 17:19 | 0 | 0 |  | in both case the limiting factor will be the tube in the gauge, and in conformity with EN837-1, it resists 130 of the full scale without damages. in this case as the same tube is used for -1...9bar and 0...10bar, you can consider you're ok up to 13bar (test requirement for EN is 15min without any change in the conformity and integrity of the gauge).
TCK57277 | User | Service Request | leakage test of EF6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-12 16:51 | 2025-03-14 15:33 | 0 | 0 |  | all our EF6 are tested for leaks, before being calibrated, we test the product for leaks with air (below 100bar) or with a water + anti-freez mix (over 100bar and with anti-freez as the calibation is made over the whole range of temperature resistance, so down to -40C) the test is to pressurize and close the system and check for a pressure drop (with a ATEQ system).
TCK57270 | User | Service Request | BAIT / Customer: Scandiuzzi Steel Constructions / Gauge for Hydrogen | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bonifazi, Daniele [boda] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-12 16:01 | 2025-03-13 15:41 | 1 | 0 |  | Answered H2 connection question via teams call.
TCK57243 | User | Service Request | Connecteur 10163310 - CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-12 12:11 | 2025-03-12 16:09 | 0 | 0 |  | chez nous le connecteur DIN 43650 standard (compatible avec le CTX sortie DIN) est le 11125168. je t'envois une photo du produit par teams.
TCK57231 | User | Service Request | TR: rfq D820 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-12 10:29 | 2025-03-12 15:53 | 0 | 0 |  | on a vérifié sur ce D820 et d'autre et sur le plan du revêtement, il couvrira toute la face RF (toutes les parties en contact avec le fluide.
TCK57161 | User | Service Request | 30272155 / MEX8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-11 11:40 | 2025-03-11 15:36 | 1 | 0 |  | c'est compliqué les couples de serrage car ils ne sont pas définis dans les normes. mais avec nos joints alu fournis, le couple de serrage pour un G1/2 sans apport de lubrifiant ou de loctite sera de 70 - 100 Nm. c'est ce qu'on utilise pour le montage de manomètre sur séparateur en G1/2 en atelier.
TCK57136 | User | Service Request | Replacement ¦ ED 522/384 231/065 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-11 09:01 | 2025-03-11 16:28 | 1 | 0 |  | according to the part number, it's a transmitter with clamp connection (size to confirm). 4...20mA 0...6bar DIN43650 Stainless steel calibration at 100C accuracy 0.4% (according to the ancient calculation system) the best replacement would be a PP20H but the size of the clamp is to confirm and the electrical connection in DIN43650 is not available. another option is the Y91333B20R + DANC (size to confirm) but the accuracy of the mounting will be a bit higher (1% normally).
TCK57035 | User | Service Request | TR: BCH - M. Mazier - AROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-10 10:15 | 2025-03-10 14:31 | 1 | 0 |  | nous n'avons pas fait de test exacts mais compte tenu des utilisations des clients et de la rigidité du tube, on dit généralement qu'il est possible de les soumettre ponctuellement a un vide "grossier" qui est généralement définit comme la plage entre 1mbar et 1bar absolu.
TCK57028 | User | Service Request | Step/CAD file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-10 09:54 | 2025-03-10 11:48 | 0 | 0 |  | I'm sorry but we don't have the ARIN26F6 3D drawing available in our library but you can find the dimensions on the datasheet. or you can order the creation of a technical document via sales to have the 3D created.
TCK56956 | User | Service Request | RPPN-7G3.205 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-07 10:30 | 2025-03-10 16:02 | 1 | 0 |  | your setp are possible as below, but we strongly suggest to order it already set, it's the most complicated one to set. Switch 1: RPPN-7G3.205/SETP "https://servicedesk.baumernet.org/wm/" SETP Micro 1: 11 Bar Rising SETP Micro 2: 11.5 Bar Rising Switch 2: RPPN-7G3.205/SETP "https://servicedesk.baumernet.org/wm/" SETP Micro 1: 9.8 Bar Rising SETP Micro 2: 10.3 Bar Rising
TCK56912 | User | Service Request | TR: Alstom ¦ PO7000326575 - SIMPLE MANOMETER DIA 80 0-10 BAR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-06 15:59 | 2025-03-10 12:17 | 1 | 0 |  | En combinant le numéro de DTR client avec le code plan je trouve celui ci, tu peux vérifier les caractéristiques avec le client mais ça devrait être ça. FICHE PRODUIT - MANOMETRE FERROVIAIRE / PRODUCT FILE - RAILWAY PRESSURE GAUGE Internal codeSAP code / 10287862 (Kanban) MFE413344JB2-02 (G3/8 svt EN837 (REGIOLIS)) 11093687 (MFE)DiameterCircuit numberPressure (bar)AccuracyConnectionMounting studLeft locating pin grooveRight locating pin grooveLighttingPointer beneath dial (H+6)Case colourDial colourLettering colourRoundel colourSpecifics markingsFront pointer colourFront pointer black heelRear pointer colourRear pointer black heelTrim colourTrim screw colourInternal drawing nBH parts list drawing nCustomer drawingCustomer part list drawingHomologationRemarksTechnical documentDial screen printing Front guard screen printing ResolutionInternal user's manualCustomer article code
TCK56894 | User | Service Request | RE: RFQ ETE/FLOEX - 2501054 - BAUMER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-06 11:57 | 2025-03-10 16:41 | 1 | 0 |  | on a validé que le montage est OK avec la R&D: RDN-8B6.234/SETP (ex : 10296524) + AKPLB0N603 (Hp) (10264885) + AKPLB0N608(Bp) (11120162) + D821-10N.42DC.B20 (11160482) + LRS1 Setp : 1,6 bar à la baisse hystérésis 0.6 b
TCK56849 | User | Service Request | 转发: Fw:回复: 3100206323-到货差异 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-06 08:13 | 2025-03-10 11:51 | 1 | 0 |  | the material of our tubes was standardized to the 316L (1.4404) for all tubes. before the 316Ti (1.4571) was used on some high pressure ones. all tests and validations were performed to ensure compliance to the EN837.
TCK56664 | User | Service Request | FW: BAUMER Pressure switch / TA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-04 10:31 | 2025-03-04 11:26 | 0 | 0 |  | I can bring a bit more information, it's an MR5 or MS5 with 1 switch and a bottom connection, but what is weird is that it's not one of our diaphragm seal and I don't have the DN to offer a replacement.
TCK56582 | User | Service Request | H2 cone and thread RFQ_0025-00402 --- AROE ---- 03.03.2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-03 12:00 | 2025-03-03 15:14 | 1 | 0 |  | pour le MEP3 on ne sera pas prêt tout de suite car on est tributaire de l'inde pour les raccords, mais on va lancer des demandes échantillons et lancer des validations. par contre en 1/4 conique fileté je peux te proposer le MEP5 pour l'instant, on a déjà fait des protos en prod' et il est possible a créer et a fabriquer. C018: RACCORD MEP5 HP 1/4-28 UNF PAS A GAUCHE avec accessoires (bague et écrou)
TCK56574 | User | Service Request | TR: Creation BAFR // Client : ATI EXPORT // Cde2330434/ Clt : 35647691/ 55145211 // TSF100-231.BB0.12T.20 // Position 10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-03-03 11:20 | 2025-03-03 15:24 | 1 | 0 |  | tu as envoyé la demande sur la mauvaise adresse, tu peux faire les demandes de création chez: FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org>
TCK56501 | User | Service Request | INFO SEPARATEUR D821 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-28 12:08 | 2025-02-28 14:06 | 0 | 0 |  | c'est bien correct, le 1/2" est plus petit que le DN15 (membrane de 32 contre membrane de 38mm) et sa limite de montage est de 10bar contre 6 pour le DN15 donc c'est normal.
TCK56487 | User | Service Request | TR: Création BAFR // Client : SOLYRO // Cde Clt : HA2500669RSR // MEX5-463.B22/0751_0760_0078_9941 // Position 80 | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-28 11:01 | 2025-02-28 12:01 | 0 | 0 |  | Marie Claire, mauvaise adresse, merci de renvoyer la demande a: FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> Best, Nicolas
TCK56475 | User | Service Request | RP2-NR.654.DT2.2L90.0.960.01/SETP_9941, 3 Sets/ Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-28 09:50 | 2025-02-28 11:57 | 1 | 0 |  | [Image]
TCK56454 | User | Service Request | product identification ¦ code: 6685700109896 ¦ Customer request: Products | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-28 07:36 | 2025-02-28 11:51 | 1 | 0 |  | the code 6685700109896 is from Westinghouse and I don't find relation with a Bourdon part number. but on the second line you have a Bourdon code, I would offer this: MIX7A63D20/0771
TCK56452 | User | Service Request | Need details of PG as per attached image | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-28 06:58 | 2025-02-28 11:35 | 1 | 0 |  | yes it is special, this is something we do for mercedes in germany, but we can't sell it to someone else, the mercedes employees should contact mercedes in germany to order one.
TCK56426 | User | Service Request | TR: AVF - LISTE MANOS MDPX et MPB.S | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-27 13:50 | 2025-02-28 11:33 | 1 | 0 |  | pas de soucis, ci-joint les anciennes datasheets pour référence.
TCK56364 | User | Service Request | BAPL ¦ N 05014.0001 ¦ Type of stainless steel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-26 14:33 | 2025-02-27 09:52 | 1 | 0 |  | I would have to dig to find a more specific answer, but I can tell you (and it should probably satisfy your customer) that it is for sure 316L / 316Ti.
TCK56337 | User | Service Request | Optrions 0679, 0678 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-26 10:14 | 2025-02-27 09:44 | 1 | 0 |  | here is the page about the pointers in the presentation of gauges' options. the micrometer (0679) is very precisely adjustable by a system of gears but it's heavy so not suitable for very low pressure ranges. the friction (0678) is adjustable because the pointer is fixed in pressure between 2 aluminum plates and can be adjusted by force. it's less accurate but lighter, so for use on low ranges. [Image]
TCK56170 | User | Service Request | Y91 à connecteur M12 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-24 11:22 | 2025-02-25 11:32 | 1 | 0 |  | les IP sont notés a coté des raccords dans la datasheet.
TCK56164 | User | Service Request | [BHFR Material Data] FW: Our inquiry / prijsaanvraag 22513751 | Process Instrumentation > Material Data Management | Low |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2025-02-24 10:59 | 2025-02-27 09:56 | 1 | 0 |  | yes, we can replace this, I suggest a Y91333B20R + D820-10L.42CH.B2001. I went with standard when information was missing (so confirmation with the customer necessary). and we didn't renewed our marine approval since several years (because of lack of requests for it).
TCK56055 | User | Service Request | adv -- question GAMME MANO DIF mx/mz | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-21 09:48 | 2025-02-21 10:09 | 0 | 0 |  | c'est bizarre, on a jamais détendue en -1...+ sur du différentiel, seulement sur du relatif. après si létendue de 25bar diff est possible c'est OK mais il faudra créer le cadran (X26)
TCK55927 | User | Service Request | Liquid detection sensor | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lim, Janice [ljan] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-20 02:53 | 2025-02-20 09:52 | 1 | 0 |  | this on the picture was an LSM (even the previous version I think). it can be replaced in the modern catalogue product line below:
TCK55920 | User | Service Request | GRCH -- ICCARE --- 19.02 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-19 16:46 | 2025-02-19 17:49 | 0 | 0 |  | pour moi ce n'est pas du DN80 PN20 sinon ce serait une bride 8 trous. mais un remplacement avec un D820 en 3" classe 150 donne bien une bride 4 trous. il faut qu'il vérifient avec les dimensions de la datasheet et peut être leurs marquages de bride mais pour nous le montage mano séparateur est possible.
TCK55834 | User | Service Request | transmetteur ETTN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-19 10:53 | 2025-02-19 14:59 | 1 | 0 |  | oui pas de soucis sauf en version plongeur câble. il faut demander le digit de réglage dans la codif et préciser létendue de mesure que tu veux.
TCK55799 | User | Service Request | Pressure gauge 0+40 kPa / Exxon | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-19 07:58 | 2025-02-19 14:42 | 1 | 0 |  | It's weird, D11 (40kPa) should not exist, it's below our minimum range and we have no tube for this (maybe the reason it was a class 2.5 on the dial). we'll have to offer the minimum possible which is D12 (60kPa) so the current best option to offer is: MEP5-461.D12/0679/0718
TCK55751 | User | Service Request | Part number question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-18 14:02 | 2025-02-18 14:17 | 0 | 0 |  | this is to separate the "Redesigned" MEP5 from the old version. you'll find the same thing on the MPG6.
TCK55703 | User | Service Request | Pressure switch requirement OEF/GPP-2/PR/504/2024 (Petroexcel) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-18 09:14 | 2025-02-18 09:47 | 0 | 0 |  | I would go for a RP2ERN54
TCK55645 | User | Service Request | WEP5 information, Urgent | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-17 12:00 | 2025-02-17 15:27 | 1 | 0 |  | the best would be to set a live call to help them, I'll be able to find the codes if you give me the barcode on the product but we could do the same I did with US and do a live session where we can help them with the codecs creation, decription and many other things.
TCK55500 | User | Service Request | FW: Atex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-13 16:10 | 2025-02-14 16:10 | 1 | 0 |  | I checked the marking on a TBI and it's the same as on the datasheet, only the global marking with no details. it was handled in germany and it might just be the minimum marking authorized. It doesn't mean we're less good as our overall zoning and classifications are the same. but I would also be in favor of more details. I asked Jorg and we'll have R&D and quality check this out and make updates (the c in the ATEX marking for example is obsolete and should be replaced by an h or nothing).
TCK55469 | User | Service Request | TOTAL - VDAV -OPTION : 0117 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-13 12:15 | 2025-02-17 17:28 | 1 | 0 |  | Yannick, on a pas l'option plombage sur les lunettes plastiques, désolé. il faudrait des grosses quantités pour refaire les moules avec un trou, si besoin total il faudra passer par une SPR. Best, Nicolas
TCK55460 | User | Service Request | TR: BAIT - IMA / ECOR - DRCG101.211.133.535/3273_DANC07-2L2B_LRS1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-13 11:28 | 2025-02-13 12:00 | 0 | 0 |  | non ce n'est pas possible, le DANC07 (DN25) ne se monte pas avec des manos diameters 100 ou plus. tu peux trouver cette information sur la datasheet du DANC ou dans le Seal's Navigator.
TCK55419 | User | Service Request | KUHN --- MEX3 / 0191 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 17:22 | 2025-02-13 11:55 | 1 | 0 |  | l'option est possible mais limitée (seulement possible sur quelques tubes en C). donc pas possible au dessus de 60bar, désolé.
TCK55416 | User | Service Request | AROE --- CUENOT --- mft | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 17:15 | 2025-02-13 12:00 | 1 | 0 |  | il faudra voir comment le codifier mais on peut proposer quelque chose: déviation sur le contact électrique, on a pas du tout la connexion spécifique qu'ils demandent. DOUBLE DIAPHRAGM DIFFERENTIAL PRESSURE GAUGE Fantinelli model Dial Size Connection 1/2" NPT-F (ON TOP) Range bar Accuracy Mounting "D" Case Type Case material Execution "D" Wetted parts V42= connection ON TOP with flusnhings on bottom V43= V47= V58= BOURDON logo on dial V99= CM 33 Connector M20X1,5 side of case Exec.Q
TCK55402 | User | Service Request | Approval? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 15:27 | 2025-02-12 17:01 | 0 | 0 |  | yes, the 316L is approved by the EN 1935 (MOCA) but I don't have an internal compliance document. the BH1, 2, LRS1 and 10 complies under the accidental pollution of the EN1935 (certificate enclosed). otherwise we work more with the internationnally recognised 3A and FDA standards, certificates enclosed.
TCK55401 | User | Service Request | MG5 spares | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 15:25 | 2025-02-14 10:42 | 1 | 0 |  | the window of an MG5 is a bit complicated but here are the components you need to replace: 10308598 VOYANT TRIPLEX MG/MR/MS5 1 PC 10307989 REHAUT 31,8 MG5 check with the sales if you can order it and how, and once the order is placed if you tell me at the right moment I can mount everything together before sending it (will be easier for you and avoid wrong gasket or spacer placement).
TCK55392 | User | Service Request | remplacement d'un E92 avec label peinture | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 14:22 | 2025-02-12 15:57 | 0 | 0 |  | pas de soucis, n'oublis pas les options habituelles de montage et de précision de l'ensemble, ça peut donner ça: Y91-3.3.3.B35.R/2164/0831 + D055-2L14.07/0832/0592/9010
TCK55388 | User | Service Request | Confirm selected model is correct | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 13:45 | 2025-02-17 16:00 | 0 | 0 |  | this is a quite complicated topic but after checking with sales that did it on a project long ago and quality, we accept you order it with a NACE certificate (Q017).
TCK55367 | User | Service Request | [BHFR Material Data] FW: Onze prijsaanvraag 6000275814 [**EXTERNAL**] | Process Instrumentation > Material Data Management | Low |  |  | Closed | Aernouts, Senne [asen] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2025-02-12 10:58 | 2025-02-21 10:04 | 1 | 1 |  | sorry, your email bounced back and forth and ended up on the technical support ticketing system. I validated the Baflex you wanted was ok and requested the codification, when I receive the code, it'll be transfered to you.
TCK55339 | User | Service Request | CUENOT-AROE-MORCILLO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 08:55 | 2025-02-12 10:14 | 0 | 0 |  | en réalité surement, mais la norme ANSI B36.10 défini les maximums en P et T pour un schedule 80 comme 100bar a 400C. on ne peux pas aller contre sans s'affranchir d'une recommandation normée, je ne vois pas vraiment qui fera ça.
TCK55330 | User | Service Request | Need pressure gauge with welted parts in Halar coated | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-12 06:09 | 2025-02-12 10:02 | 1 | 0 |  | to get a coating you'll need a diaphragm seal, so the DT2 will fit and then you can check the materials and coating options on the datasheet but here is a configuration that could fit your need: MEX3D20F20 + DT22H24023401
TCK55227 | User | Service Request | MONTAGE STD770 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-10 14:47 | 2025-02-10 15:40 | 0 | 0 |  | pour moi en 1" tu devra passer en D912 en allant au plus simple. le STD770 est a valider pour montage, j'ai du STD720 ou du STG770 pour l'instant donc a confirmer.
TCK55192 | User | Service Request | TR: BAIT - Alstom - EF6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-10 10:32 | 2025-02-13 10:07 | 1 | 0 |  | Normally it' more up to the installer to know the mounting torques but I checked on several documentation (including wika and parker) and with our guy on the manufacturing line and I find things between 55 and 75N.m. also th fact that they are using loctite usually mean that the low end of the torque range would be enough (so I'm thinking about 60 to 65N.m)
TCK55175 | User | Service Request | RPPN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lim, Janice [ljan] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-10 09:33 | 2025-02-10 17:07 | 1 | 0 |  | both products are from order 1245595 position 170 and 220: below the details: 11190409 RPPN-7B6.205/2261_SETP (11.5bar rising) 11190408 RPPN-7B6.204/2261_SETP (1.5bar decreasing)
TCK55165 | User | Service Request | VANNES YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-10 08:25 | 2025-02-10 15:46 | 1 | 0 |  | les doubles vannes ne sont dispo que sur le D918 et uniquement sur la class 2500. il n'y a pas de raison de mettre des doubles vannes de ce type sur d'autres montages, les vannes standard conviennent.
TCK55161 | User | Service Request | Confirm ambient temperature suitability for -29 Deg Cel. model Mex2 & MMX1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-10 07:43 | 2025-02-10 09:47 | 0 | 0 |  | please find enclosed the note detailling the limits for low temperature use. it means the MEX2 is fine at -30 but not the MMX1, sorry.
TCK55154 | User | Service Request | length of the connector 转发: 回复: 回复: 回复: 回复: WSS型温度计 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-08 10:50 | 2025-02-10 16:04 | 1 | 0 |  | here is the measurment you request for an M20*1.5 fix connector for TBHI (picture below). [Image]
TCK55042 | User | Service Request | morcillo-- cuenot---ara | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-06 12:17 | 2025-02-07 09:35 | 1 | 0 |  | Oui, bien sur, tout dépendra des pulsations mais aucun problème pour la viscosité. Best, Nicolas
TCK55029 | User | Service Request | MEX5-D60.B74/0751 + NW 16KF Flange | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-06 10:53 | 2025-02-07 09:45 | 1 | 0 |  | we don't have NW/KF flanges in our product range. as far as I know it's a type of clamp flanges used in vacuum industry. my suggestion is to either use an adapter like below or check if they could use ou iso clamps (DANC or adapters). [Image]
TCK54991 | User | Service Request | KAESER--CCT---05.02.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-05 17:05 | 2025-02-05 18:22 | 0 | 0 |  | L'impact sera minime sur un 100bar, et la précision ne devrait même pas être impactée. je suis plus inquiet sur un risque de fuite avec la pression mise sur lévent, avec du BH2 ça devrait aller mais surtout il fit bien une mise a latmosphère, puis replace lévent avant le montage et ne coupe pas le téton. et lidéal serait de le mettre en pression au moins une fois une fois installer pour vérifier que le tube ne crée une trop grosse différence de pression qui pourrait causer une fuite. la température pourrait aussi être un facteur aggravant. lidéal quand un montage comme ça est prévu c'est de demander une connexion a 90, comme ça lévent sera quand même sur le haut et il n'y aura pas de problème. Best, Nicolas
TCK54962 | User | Service Request | à propose des TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-05 12:49 | 2025-02-05 13:43 | 0 | 0 |  | Voici les certificats des TED5 et TED6. Best, Nicolas
TCK54942 | User | Service Request | TR: Raccord H2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-05 11:16 | 2025-02-05 15:13 | 1 | 0 |  | j'ai retrouvé des femelles 9/16 UNF mais on a redessiné ce raccord il y a peu et j'ai un plan trop détaillé, je t'ai fait lextrait des éléments importants pour le client (code C069): [Image][Image] [Image]
TCK54937 | User | Service Request | replacement ¦ 11015579 ¦ YTED-H.3.B38.R ¦ Zapytanie ofertowe Baumer - UP/279/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-05 10:22 | 2025-02-05 10:58 | 0 | 0 |  | we didn't renew the ATEX certification on TED, but here is a similar one without ATEX. otherwise an option with a PFMN atex can be offered also. TED-6.H.3.B38.R 11139818
TCK54806 | User | Service Request | MIX7-A60.B22.AKPL.D295-2N70.22P.09/0191+ SAP 11214522 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-03 17:32 | 2025-02-04 10:19 | 1 | 0 |  | D295 replaced D294 and LRS9 replaced LRS3 and the rest seems the same, it's the right replacement.
TCK54788 | User | Service Request | BAPL ¦ Request for help in identifying MG5 and TSSE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-03 15:15 | 2025-02-03 16:35 | 0 | 0 |  | it's too old to find a SAP code or a reference, and I can't tell much from the pictures. What I can say is it's an MG5 single contact inductive ATEX and a TSSE single contact inductive ATEX
TCK54762 | User | Service Request | RP2 Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-03 13:07 | 2025-02-03 14:48 | 0 | 0 |  | I would replace it by a standard (A) or an hermetic (C), those are the 2 with the most silimar cutting power and they have the same setting caracteristics.
TCK54732 | User | Service Request | ATEX Certificate for RP2E / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-03 10:43 | 2025-02-03 14:38 | 1 | 0 |  | please find the certificate enclosed.
TCK54695 | User | Service Request | product identification ¦ code: 0607 ¦ Moulded screened cable M12-8 pin. Length 2 m ¦ Zapytanie 8616/25/ZD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-02-03 07:02 | 2025-02-03 14:36 | 1 | 0 |  | I can't say for the sales if they can sell it as a spare and the price but here is the code of this cable: 10303910 CORDON M12*8 2M BLINDE RKTS82992M
TCK54669 | User | Service Request | TR: DDPMO1049 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-31 14:28 | 2025-02-03 14:33 | 1 | 0 |  | les manomètres de pression d'ambiances pour les chambres hyperbar n'avaient pas de connexion et vu comment il tient, celui ci en a une. je dirais que c'est simplement un mano de pression d'eau externe donc a remplacer par un type MEX en 10 bar avec la double graduation en pieds et m d'eau de mer. mais on ne fera pas la graduation inversée, l'aiguille ira du zéro a gauche au 10m d'eau de mer a droite (om = 0bar et 100m = 10bar)
TCK54535 | User | Service Request | 3d CAD files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-29 17:07 | 2025-01-30 16:25 | 1 | 0 |  | please find enclosed the 3 3D drawings.
TCK54521 | User | Service Request | FW: REQUEST FOR QUOTATION | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-29 14:23 | 2025-01-30 16:22 | 1 | 0 |  | those were manufactured by KHW in germany and are a bit old so it's difficult to come back to an exat part number but it seems it was something like: DRC 100/811.133/-1..3 BAR/1541/1310 we could offer to replace it by a: MEX5D20B76 + 1541 + LRS1. I see no visible options but you can check if there is things to add with the customer.
TCK54490 | User | Service Request | TR: Demande séparateur sur capteur de pression | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-29 10:02 | 2025-01-30 15:56 | 1 | 0 |  | je n'ai jamais vu ce raccord, il semblerais que ce soit de l'APV DN50 mais je ne trouve pas beaucoup d'infos, je doute qu'on puisse proposer rapidement quelque chose et vu qu'on ne connais pas ce raccord, je doute que la demande soit très important. sinon il faudra faire un projet de développement de ce séparateur.
TCK54481 | User | Service Request | WG: Votre demande concernant le manomètre MEX5-D31.B20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Sivri, Berkan [sibe] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-29 09:16 | 2025-01-30 15:42 | 1 | 0 |  | accoring to the informations we have from the customer and considering the fluid is oil, I would suggest a small ASIPD as it'll be more compact and robut, and will easily lower the temperature of the oil to around 30-50C inside the gauge. the only thing to check is the thickness of the oil as the diameter of the capillary for the fluid to travel inside the colling element will be 1mm. if the oil is too thick or contains too much particles, it'll be worth adding a diaphragm seal like a D043 MEX+ASIPD+D043+LRS9 or LRS10 would be a great fit.
TCK54389 | User | Service Request | availability/ replacement ¦ DROE 100/813.133 095/0201/0794/9711/5001 4710/1085 ¦ INQUIRY URGENT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-28 10:55 | 2025-02-04 10:35 | 1 | 0 |  | here is the replacement for your mounting with Bourdon products instead of the Haenni ones. please check with your client if OK with him, the minimum range for mounting with the diaphragm seal required that I changed the range from 25 to 40bar. MR52D3KB26/0794/9711/4710 + D0552L15.07
TCK54383 | User | Service Request | ETTN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-28 10:20 | 2025-01-30 15:13 | 1 | 0 |  | j'ai chécké beaucoup de codes et j'en trouve seulement 2 qui matchent ta demande. vérifie avec le client mais c'est tout ce que j'ai trouvé: ETTN61X3D1/0604_SETP EN 11144177 300mm ETTN61X3D1/0604_SETP
TCK54343 | User | Service Request | Safety glass window | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-27 17:57 | 2025-01-30 14:58 | 1 | 0 |  | Answered by Jorg.
TCK54287 | User | Service Request | RE: Question sur l'impact de variation de pression atmosphérique sur les capteurs de pression | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-27 10:42 | 2025-01-30 14:58 | 1 | 0 |  | voici la réponse de Philippe sur le sujet, désolé, une autre semaine j'aurais pris le temps de chercher des détails mais j'essaye pour l'instant de rattraper mon retard, si d'autres détails sont nécessaires, on se redis par teams. "Notre client est Arésia (anciennement Alkan) pour ces 2 produits. Pourquoi une telle question, il devrait la poser à Arésia et non à nous. Après ces 2 produits fonctionnement à des pressions de 250 bar et 300 bar. La pression atmosphérique ninterfère pas sur la mesure au vue de la précision demandée."
TCK54266 | User | Service Request | FW: **SPAM** Pressure Indicator / Chemicoat | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-27 08:24 | 2025-01-27 10:34 | 1 | 0 |  | according to the description, it's a MIX7 or MEX8 with a blow our back plug and a special dialso we can definitely do that. the only thing I can't be sure about is the connection, according to the text it is 1/4NPT and then it's fine but on the picture with the plastic protection it seems weird for a 1/4NPT so we need to clarify this with the customer. on the dial, the alarm and cut of marks can be replaced by an orange and red sectors for example.
TCK54199 | User | Service Request | Installation information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-24 10:23 | 2025-01-24 11:56 | 0 | 0 |  | with datas gathered from different sources of threads suppliers and internal recomandations here is what we can say: 2 or 3 wraps of PTFE tape around the thread tightening with 20 to 30N.m of torque (2 to 3 turns from hand tight)
TCK54139 | User | Service Request | AORPB504 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-23 11:04 | 2025-01-23 11:22 | 0 | 0 |  | you're right on the rule, the usual setting is the max of the scale of the instrument (so 1bar here) for the setting if it's fitted with one of our instrument I suggest to order it as an assembly with the setting done here. if it's for mounting with something else, it's better if the setting is done by the person doing the assembly but we can still set it here (it's just that as it'll be manipulated during assembly, it's good to do a final test at least after assembly).
TCK53982 | User | Service Request | grch-dow-21/01/2025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 15:30 | 2025-01-27 14:47 | 0 | 0 |  | Ça semble compliqué de le faire avec un ME7 mais on a une autre solution pour le remplacer, a voir si ça te va : MCX7D60N67/9696 + AORPB501 + ARIN2606 + D631.2L2.04DEB20.01 (-100 +160 mbar avec secteur vert de -20 à 20mbar et AORPB réglé a 160mbar).
TCK53979 | User | Service Request | TBI ---MAURYA --- SLB | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 15:24 | 2025-01-21 15:54 | 0 | 0 |  | selon la norme la résistance en température ambiante des thermos doit être définie par le fabricant. mais étant donné que je ne trouve aucune info je vais te donner les limites minimum qu'on est obligé de respecter dans le paragraphe pour le stockage soit -20....60C ambiante.
TCK53975 | User | Service Request | Scale clarification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 14:20 | 2025-01-21 14:35 | 0 | 0 |  | our standard is black/black and otherwise it'll be an X24.
TCK53955 | User | Service Request | Y91 sur ARID | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 11:28 | 2025-01-21 14:20 | 1 | 0 |  | le code que tu donne comme exemple était composé des éléments ci-dessous, 11203048 PBMN-23B76RA14403201010/Q021 11202636 MG5-7D30.B76/4710 11154663 XDSE:53888_01 10265045 ARID-DF5F53 10303605 VASELINE 30CST SP2046 le Y91 proposé peut remplacer le PBMN et tous les autres éléments sont ouverts dans SAP pas de soucis par contre le prix du séparateur est a re-valider avec les ventes et le Q021 sera à commander séparément. 10301518 Y91-3.3.3.B76.R 11202636 MG5-7D30.B76/4710 11154663 XDSE:53888_01 10265045 ARID-DF5F53 10303605 VASELINE 30CST SP2046 (0592 + 2151 +9010 options obligatoires pour le montage Y91 sur séparateur)
TCK53952 | User | Service Request | Y91 Power on to reading time | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 11:05 | 2025-01-21 14:29 | 1 | 0 |  | I don't think it is something we measured but it's pretty immediate. once the power enter the circuit it'll pass the wheaston bridge and immediatly be transform to have an output so pretty much the time it takes electrons to run through the circuit. the display will be taking more time anyway.
TCK53945 | User | Service Request | Blocked material | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 10:49 | 2025-01-21 14:34 | 1 | 0 |  | it was blocked a first time because not sold for more than 5 years and then it was closed in september last year because the price had to be reviewed. the FUD (tubular seal) is open and the PBMN is open in SAP and on the website. you can request the reopening of the code and a new price will be confirmed.
TCK53914 | User | Service Request | Enquiry for PBSN-A-A/M-M-NCDE / Gilang | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-21 07:48 | 2025-01-21 14:25 | 1 | 0 |  | if I read correctly the manufacturing order would be 102071592 which is a: 11014632 PBSN-1.2.B22.R.A2.14.06.2.1
TCK53901 | User | Service Request | INFORMATION REGLAGE MANO CONTACT ORANO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-20 17:12 | 2025-01-20 18:18 | 0 | 0 |  | la clé pour faire les réglages est disponible, elle est emboîtée dans le presse étoupe électrique sur le coté de l'appareil. après ils poussent pour descendre la lame et ils tournent pour emmener les contacts.
TCK53847 | User | Service Request | montage sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-20 11:59 | 2025-01-20 15:56 | 1 | 0 |  | les 2 montages sont OK, après c'est ta fonctionnalité qui va changer entre le PBMN et le Y9 mais ça c'est ta partie. n'oublie pas que pour tous les transmetteurs montés sur séparateurs il y aura l'option monté sur séparateur et l'option précision x2. et sans partie basse de DT2 il faut qu'ils en ai une sur site pour le montage.
TCK53818 | User | Service Request | DEXIS-CLAUDON- VDAV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-20 09:07 | 2025-01-20 10:56 | 0 | 0 |  | oui, je pense aussi, un MR ou MS5 avec un DT2 et après au niveau matières ça dépends de son process, au minimum une membrane en HC, et on peux ensuite revêtir la membrane et ou le flasque inférieur voir même le proposer tout en PTFE si la température le permet.
TCK53792 | User | Service Request | avf-vdad-sep MINI CLAMP1095 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-17 15:29 | 2025-01-20 10:12 | 1 | 0 |  | je ne suis pas sur de comprendre, la ref complète du 1095 c'est 1095... voici un exemple de code, ça te donnera en même temps la limite basse pour un diamètre 63: 11274991 MX35-D20.B22.1095/LRS1_1359_9941
TCK53736 | User | Service Request | Dial Scale For MPG6-A61.V19 / TA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-17 04:39 | 2025-01-17 09:47 | 0 | 0 |  | here is what it looks like: [Image]
TCK53718 | User | Service Request | TBHI over pressure | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-16 15:11 | 2025-01-17 14:57 | 1 | 0 |  | The issue here is that the connection is not really the problem, it'll resist 60bar of pressure. the problem will be the stem, it is very unsusual at this pressure to not use a thermowell for safety and easy replacement without stopping the process. if it is 60bar of a static gas, it'll be ok, but 60bar of a fluid, even if not moving a lot, do you immagin the effort it puts on the stem ? My official answer should then be that according to the usages and good practice, a thermowell should be used.
TCK53679 | User | Service Request | MPG6-A62.H78.D043-2L06.02 & MPG6-A62.H22.D043-2L06.02 / Servo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-16 11:08 | 2025-01-16 12:05 | 0 | 0 |  | it's not the same product so there will be differences of course (our products are just generally speaking slightly different then wika so they need to review the datasheets to check if something is an issue for them but as far as I reviewed your offer is very similar). but your option seems like a valid alternative to the wika one. I just added the welded connection: MPG6-A62.H78.D043-2N06.02/9030 MPG6-A62.H22.D043-2N06.02/9030
TCK53665 | User | Service Request | TR: LOx Tank - BAUMER - PO035850 - Interférences vibrations | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-16 09:57 | 2025-01-16 11:31 | 0 | 0 |  | voici les qualifications que nous avons fait sur nos produits, en vibrations ils sont testés selon la norme EN60068-2-6. c'est vrai que cette norme ne prend pas en compte les fibrations de faible fréquences mais je ne vois pas vraiment de raisons pour que le produit se déstabilise autant (l'instabilité du dernier digit est normale mais les autres c'est étonnant) quand ils disent que le 2ème produit est sur la même barrette et reste stable est ce que cela veut dire qu'il est aussi soumis aux vibrations ? [Image] si cette étude doit aller plus loin, merci de m'envoyer les retours de air liquide sur ma boite perso, il faudra intégrer le spécialiste produit pour aller plus loin.
TCK53655 | User | Service Request | YOKO 11093448 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-16 08:48 | 2025-01-16 09:46 | 0 | 0 |  | je n'ai pas de plan d'ensemble mais j'ai le plan de la bride inférieure si tu veux: ça permettra au client de vérifier les détails pour le montage et ça ne dévoile rien de relatif au moyen de mesure (membrane, voûte, volumes etc...).
TCK53623 | User | Service Request | 0-16Bar/Psi gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-15 15:35 | 2025-01-16 11:18 | 1 | 0 |  | here is the dial you would get with a K24 on a MEX5 and it's going to 16bar and approximately 232psi: [Image]
TCK53603 | User | Service Request | RE: DEFINITION PRODUIT TRES UREGENT PROJET ARAMCO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bredon, Jérôme [jbr] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-15 12:06 | 2025-01-15 12:10 | 0 | 0 |  | ce serait un D920 spécifique mais la bride API est à dessiner, chiffrer et il faudrait trouver la matière, vérifier la soudabilité avec l'HC276 et valider le montage (HP standard, pas diff). et definir, dessiner et chiffre le flushing ring. ce serait surement possible sur un projet a long terme mais aucune chance d'avoir ca demain.
TCK53600 | User | Service Request | DEFINITION PRODUIT TRES UREGENT PROJET ARAMCO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-15 11:39 | 2025-01-15 12:10 | 0 | 0 |  | ce serait un D920 spécifique mais la bride API est à dessiner, chiffrer et il faudrait trouver la matière, vérifier la soudabilité avec l'HC276 et valider le montage (HP standard, pas diff). et definir, dessiner et chiffre le flushing ring. ce serait surement possible sur un projet a long terme mais aucune chance d'avoir ca demain.
TCK53580 | User | Service Request | max. prcocess pressure of TBHI100 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-15 09:01 | 2025-01-17 14:58 | 1 | 0 |  | The issue here is that the connection is not really the problem, it'll resist 60bar of pressure. the problem will be the stem, it is very unsusual at this pressure to not use a thermowell for safety and easy replacement without stopping the process. if it is 60bar of a static gas, it'll be ok, but 60bar of a fluid, even if not moving a lot, do you immagin the effort it puts on the stem ? My official answer should then be that according to the usages and good practice, a thermowell should be used.
TCK53508 | User | Service Request | MPJ6 & MPG6 / Simon GTS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-14 10:47 | 2025-01-14 11:11 | 0 | 0 |  | the dial X26 kPa ANSI exists on the MPG6 and it's the same as MPG6 so to me it's ok (idem 11111958). and 11267185 MPG6-A60.X39/0771_9709_D000 is kPa ANSI and it's open for sales.
TCK53480 | User | Service Request | Dial Scale MPG6 in PP case / TA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-14 07:05 | 2025-01-14 09:16 | 0 | 0 |  | for the S34, the code already exists and the scale is ANSI so no problem 11178570 MPG6-A60.S34 for the Ds, D means kPa EN so you'll have to use X for kPa ANSI. the X22 kPa ANSI exists in multiple MPG6: example 11276229 MPG6-A62.X22/9709_D000 the X19 kPa ANSI scale exists but only with a specific LOGO, it might have to be created without.
TCK53455 | User | Service Request | TR: MABEO INDUSTRIES - Devis 01957Z 0813 - | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-13 15:20 | 2025-01-13 15:53 | 0 | 0 |  | pour le PN/DN et types de face ça semble bon de vue mais ce serait bien de vérifier le marquage sur le coté de la bride pour confirmer. concernant la matière, c'est un peu compliqué le Nickel peut être utilisé pour plusieurs produits, mais si je me base sur ses usages les plus courants (produits alcalins, caustiques type NaOH, KOH) selon la concentration et la température je proposerais 2 solutions: Monel (alliage riche en nickel au très bonnes propriétés de résistances sur NaOH et KOH notamment): D82510L42GHD90 HC276 (large spectre de résistance correcte, peut être suffisant selon la concentration et la température): D82010L42GHD60
TCK53276 | User | Service Request | TSF100-261.1B4.11T.30 + thermowell | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-09 10:47 | 2025-01-09 15:29 | 1 | 0 |  | to go with your application if no specific material or design requirement, I think the best options are as below: TSF100-261.1B4.11T.30 + T084100301 (but a slight part of the minimum insertion length (Ls) will not be in the process, it's not a very big deal but it can impact a bit the accuracy).
TCK53220 | User | Service Request | D055 versus 1080 -- AROE --- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 16:48 | 2025-01-09 17:32 | 1 | 0 |  | Yannick, demande la réouverture de ce code pour des pièces de rechange, étant donné qu'on a pas de remplacement pour le DRO80, je préfère pour l'instant. les membranes des 1081 et D055 ont étés re-testées pour revoir les minimum et pour celui ci en 4bar sur du 1" on a vu avec Christophe et ce sera bon. Best, Nicolas
TCK53192 | User | Service Request | thermostat électronique pour la SNCF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 11:17 | 2025-01-08 16:38 | 1 | 0 |  | j'avais répondu par email la dessus, cétait un produit spécifique mais qui n'a a priori jamais été vendu et ne semble pas avoir passé le stade du projet, on ne peut pas le faire en létat, si le projet est a reprendre il faudra faire un processus SPR mais c'est quasiment un développement complet.
TCK53187 | User | Service Request | Y91 Situation | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 11:02 | 2025-01-08 16:25 | 1 | 0 |  | the current situation if you place an order for like 100units would be 8weeks delivery. but for an opportunity like this, I suggest you contact the PM to determine if it's worth putting something special in place to take the business. (and we have a shortage of the component for the option 2152, for this specific one only it would be more than 8 weeks)
TCK53164 | User | Service Request | DEFINITION SEPRATEURS YOKO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 09:23 | 2025-01-08 10:28 | 0 | 0 |  | Salut Véronique, tu trouvera cette info dans la codif du transmetteur, c'est une caractéristique qui est notifiée dans un digit comme pour les flasques low volume. Best, Nicolas
TCK53159 | User | Service Request | DEMANDE SEPARATEURS HONEYWELL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 09:03 | 2025-01-08 10:26 | 0 | 0 |  | pour les brides ANSI entre 1/2" et 2"1/2, les dimensions pour de la class 900 sont les mêmes que pour la 1500, c'est la raison pour laquelle on a pas de 900. tu peux proposer en 1500.
TCK53137 | User | Service Request | FW: quotation requested | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-08 07:18 | 2025-01-09 09:08 | 1 | 0 |  | this number is the same as on the previous email, I saw there was 2 pictures and I checked both numbers already. I'll close the ticket, but send me the full picture if you get it and I'll be happy to help as much as possible.
TCK53102 | User | Service Request | Facture et avoir Teamex validés | Service Desk > General | Low |  |  | Closed | Lambert, Thierry [tla] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-07 14:52 | 2025-01-07 15:28 | 0 | 0 |  | Erreur destinataire
TCK53096 | User | Service Request | SEPARATEUR + PURGE ORANO --- TVE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-07 14:10 | 2025-01-09 17:15 | 1 | 0 |  | Désolé, on aura pas de possibilité de PTFE en anneau de nettoyage en standard ou déjà existant. si la demande a des volumes, une spr pour voir si on pourrait designer une solution specifique...
TCK53062 | User | Service Request | D820 CL300 1" | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-07 11:26 | 2025-01-07 15:13 | 1 | 0 |  | the 1" on the D820 wasn't existing but recently we had requests for it. Code for the 1" is always the same so it would be a MEX5-D30.F22/0751 + D820-10L2224.G2101 but it's not existing right now in the system (I only have some 1" class 150) so it would have to go through an SPR procedure. I doubt it's very interesting for 1 piece but it should appear in the datasheet in the future.
TCK53060 | User | Service Request | PBMN + D820 (DN50) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-07 11:01 | 2025-01-17 09:10 | 0 | 1 |  | I checked with R&D and we can accept the 600mbar mounting (as usual for those mountings, the accuracy will be x2 and as wnever yet validated this mounting, it could be slightly extended (+/-0.6 for example) but we accept to do it. PBMN-24J12RA11403201000 + D820-10L42CHB20 (PN10/DN50, EM 600mbar)
TCK52963 | User | Service Request | FW: BW CLEAR SKY - Pressure Transmitter | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-06 09:54 | 2025-01-06 15:46 | 1 | 0 |  | the part number is too old to find out in the system but I already received the same request from belgium and this is the answer I sent them with the closest part number I could define from the picture: "For the connection, it's just an intuition given the size (likely 1/2") and that few threads are showing, with no trace of PTFE (the NPT is usually assembled with PTFE tape). We can't see anything about the gasket, so I propose the most standard option, but the customer must confirm if this is okay for them. I've also added the electrical connection, which is an option" Y91-3.3.3.B20.R/2166_2152
TCK52930 | User | Service Request | Tube connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-03 16:57 | 2025-01-06 15:42 | 1 | 0 |  | if they are asking for 1/4" tube connection, the best is to go with the 1/4x25mm. if they are working with either H2 or HP you can mention and maybe show them the 1/4 cone and thread connection, it might be interesting for them in the future but it'll not match what they need right now.
TCK52892 | User | Service Request | FW: Pump research | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-02 12:29 | 2025-01-02 15:13 | 1 | 0 |  | this sure is not a product from our company, it was manufactured in london in the first part of the 20th century (the name Siebe gorman & co Ltd started in 1904 and they moved from London in the 50s) also the design match this period. what I can't really be sure about is if at this time people rebranded other companies products but I don't think so.
TCK52890 | User | Service Request | Enquiry For E910 / Edmond | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-02 11:18 | 2025-01-02 11:49 | 0 | 0 |  | their codes doesn't make sense, so I would offer as below and ask them to check if the electrical connection, accuracy and G1/4 type are ok: Y91-3.2.3.B33.R Y91-3.2.3.B22.R
TCK52867 | User | Service Request | FW: Barflex 4/ Simon | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2025-01-02 03:17 | 2025-01-02 11:44 | 1 | 0 |  | if I'm reading correctly the order number (1977281/10) then here is the code it was: 10270571 EA4A -1000+1600MBAR R + PILE ZODIAC it is phased out but I think you can request the reopening to the PM (Mario Marolf) I think we can still produce it.
TCK52843 | User | Service Request | 回复: 差异102827 ，型号DA0956EV15MW0000产品贴纸型号和订货型号不符 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-27 09:15 | 2025-01-07 17:32 | 1 | 1 |  | We checked this out with Jorg, and we found out that this product from a sub supplier codification changed over the years and we updated it in our definition and SAP code (which you see on the box) But the marking the supplier was doing on the product was never changed. Also, this product was not supposed to still be sold today, so now Ill request a phase out, you can check with the customer if maybe he wants to place a last buy order (I think we could accept it). In the future, a replacement could be defined in our range of products with the MFT, you can send them the datasheet and a proposal for the definition of what they could use as close as possible to the DA09 specifications: DA09 56EV15MW0000 025mbar (180) (02.5kPa linear & 03.2m3/h quadratic) 10bar PN (static) 1/2NPT male 1.4404 Wall mounting No filling No special function Accuracy ? class 2.5 on the dial No Special aspects No contact So I offer to replace with this ? Order code: MFT5-A00.D05K.62 MFT old Code / SAP code: Nominal size: 100 mm 5 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection, 3 back lugs fixing A Type of case: Case without partition of safety 0 Oil filling: No oil filling 0 Pressure unit: kPa D Pressure range: Pressure range for kPa 0 ... 2.5 kPa 05 Static pressure: 25 bar K Process connection: AISI 316L (1.4404), stainless steel 1/2" NPT male 62
TCK52741 | User | Service Request | Transmetteurs de pression | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-19 15:01 | 2024-12-19 16:58 | 0 | 0 |  | tu es sur que c'est de chez nous et pas de la suisse ? je ne trouve rien ni en n d'article, de série ou de n d'ordre de fabrication. par contre ca m'a dit 2 fois que je n'avais pas accès à la division 2100 donc je pencherais pour un produit suisse.
TCK52733 | User | Service Request | Thermostat électronique | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-19 14:28 | 2024-12-19 17:08 | 1 | 0 |  | je ne trouve pas grand chose, il n'y a ni définition ni plan dans l'article mais je peux te donner quelques détails grâce à la liste des composants. Elément sensible: PT1000 Raccord process: 1/2NPT Connecteur: FICHE DIN43650 + PE PG11 REF.C81311N21 Carte câblée avec 3 fils "type CPX" plage de mesure: -40+125C seuil: 105C A LA HAUSSE il semble que c'etait un projet pour la SNCF mais je n'en trouve aucun de produit, il ne semble par qu'il ait passé le stade de proto.
TCK52677 | User | Service Request | datasheet ¦ 10302664 ¦ BRIDE SUPPORT MANO ¦ 32sztuki | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-19 06:56 | 2024-12-19 10:20 | 1 | 0 |  | this is an accessory for mounting the MFE4x or MFE5x, this is only the flange, it can also be delivered with the other accessories to fit the MFE inside (screws and push fixes). please find the pictures of the 10302664 enclosed.
TCK52590 | User | Service Request | TCK51566 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-17 16:21 | 2024-12-17 17:46 | 0 | 0 |  | I found the A009 but it was not easy and for now no product was configured with this option and I doubt you can find a price right away. for this, please structure your code as needed with the A009 option and send the request for a price to sales administration saying I approved it (otherwise they might not be aware of this option and will come back to you).
TCK52519 | User | Service Request | MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-16 16:49 | 2024-12-16 17:35 | 0 | 0 |  | il n'y a pas de définition normative de la résistance en vide des manos à tube de Bourdon. mais étant donné que c'est un tube fin (0...1bar) il ne faudra pas trop tirer au vide pour ne pas lendommager, il résistera un vide grossier s'il n'y reste pas trop longtemps (quelques dizaines de mbar absolut). pour assurer une meilleure résistance et mesure dans le domaine du vide, ils peuvent prendre un pseudo absolut en -1...1.6bar, on doit peut etre même pouvoir faire un -1...1 (c'est un tube qu'on a pour les 30psi).
TCK52425 | User | Service Request | CAPTRAIN --- aroe --- 13.12.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-13 15:50 | 2024-12-17 17:03 | 1 | 0 |  | ci-joint les definitions de ce qu'on peut proposer au plus proche et les deviations: Mano RP Mano CF Mano CG case and ring stainless steel crimped-on ring black anodised U-clamp for panel mounting stainless steel crimped-on ring black anodised U-clamp for panel mounting stainless steel crimped-on ring black anodised U-clamp for panel mounting Nominal size 80 80 100 process connection brass centre back (brass lower back) brass centre back (brass lower back) brass centre back (brass lower back) Thread M 16 x 1.5 L10 (G3/8 std offered) 2 x M 16 x 1.5 L10 (G3/8 std offered) 2 x M 16 x 1.5 L10 (G3/8 std offered) Restrictor screw Orifice 0.8 mm (3mm entry on the connection, smaller after (around 1mm) Orifice 0.8 mm (3mm entry on the connection, smaller after (around 1mm) Orifice 0.8 mm (3mm entry on the connection, smaller after (around 1mm) Bourdon tube 12 bar 6 bar 6 bar Dial Black with stop pin scale and figures white Black with stop pin scale and figures white Black with stop pin scale and figures white range/unit 0 12 bar 0 6 bar 0 6 bar Coloured marks 10 bar 6 bar 6 bar ZA Position adjustment 60 (90) Position adjustment 60 (90) Position adjustment 60 (90) Accuracy class 1.0 class 1.0 class 1.0 movement brass/ german silver (copper alloy and stainless steel) movement brass/ german silver (copper alloy and stainless steel) movement brass/ german silver (copper alloy and stainless steel) Pointer red RAL 3026 (standard red) pointer front red RAL 3026 (standard red) pointer back yellow RAL 1016 (standard yellow) pointer front red RAL 3026 (standard red) pointer back yellow RAL 1016 (standard yellow) Options with indirect lighting by lateral lighting apertures and direct lighting BA9 in the back of the case Light bulb BA9 24V (Direct LED lateral lightning between the dial and the window 24V) or without lightning if customer have external lateral lightning already in place. window reflex-reduced laminated safety glass (standard safety glass) position of installation 50-70 (position adjustment 60) (calibrated at 90) with indirect lighting by lateral lighting apertures and direct lighting BA9 in the back of the case Light bulb BA9 24V (Direct LED lateral lightning between the dial and the window 24V) or without lightning if customer have external lateral lightning already in place. window reflex-reduced laminated safety glass (standard safety glass) Position of installation 50-70 (Position adjustment 60) (calibrated at 90) with indirect lighting by lateral lighting apertures and direct lighting BA9 in the back of the case Light bulb BA9 24V (Direct LED lateral lightning between the dial and the window 24V) or without lightning if customer have external lateral lightning already in place. window reflex-reduced laminated safety glass (standard safety glass) Position of installation 50-70 (Position adjustment 60) (calibrated at 90) MFE412024JB5/0751/97xx(colored mark tbd) MFE422023JB1/0751/97xx(colored mark tbd) MFE522023JB1/0751/97xx(colored mark tbd)
TCK52410 | User | Service Request | ETUDE DAHER -- GRCH -- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-13 12:26 | 2024-12-13 15:05 | 1 | 0 |  | ils ont une référence de MMX2 en bas de la dernière page avec notre nom de fournisseur, donc je partirais sur un MEX2 mais il y a peut être un peu de recherche a faire sur les MEX2 F qu'on leurs a vendu, c'est surement pour du remplacement.
TCK52371 | User | Service Request | MVE7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-13 07:53 | 2024-12-17 10:40 | 1 | 0 |  | it'll be indicated on the back of the gauge on a label, the sames way as for all standard gauges:
TCK52338 | User | Service Request | [BHFR Material Data] MEX5-D30.B19.FUD175/23/1/33/LRS1_0751_0678_1359_0400_9007/0120 - 2300474 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-12-12 11:47 | 2025-01-02 14:58 | 1 | 1 |  | this mounting made by KHW will no longer be possible as several things were not strictly right. the S marking on the gauge is for S2 safety gauge and our MEX5 cannot be S2, also the MEX5+FUD+accessory is not 3A (we don't have a certificate for this). to me we can offer either one of those options but none will be 3A: S1 option (no S marking) MEX5-D30.B19_0751_0678 + FUD175/23/1/33/LRS1_0400_9007/0120 (0400 = 10310395) S3 option (with circled S marking) MEP5-430.B19_0678 + FUD175/23/1/33/LRS1_0400_9007/0120 (0400 = 10310395) after validating with your customer, at code creation you can put in your request that the mounting was validated by Christophe Charpentier (diaphragm seals R&D) and myself.
TCK52299 | User | Service Request | System fill question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-11 18:00 | 2024-12-12 11:30 | 0 | 0 |  | it's an estimation from calculation sheets and I rounded up to be on the safe side but the volume of LRS1 inside the system should be around 10mL.
TCK52292 | User | Service Request | replacement ¦ 11013479 ¦ EDD575 FKCW35V5-AACYY-AA ¦ zapytanie ofertowe | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-11 16:42 | 2024-12-11 17:06 | 0 | 0 |  | we stopped this because it's really not our core competence, we will have no replacement for it. specialized companies will be yokogawa, siemens or honeywell for example.
TCK52286 | User | Service Request | Industrial pressure gauges MEX5................ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Govetosa, Andrea Carmine [gova] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-11 16:14 | 2024-12-11 17:06 | 0 | 0 |  | sure you can answer, I would suggest and my comments on all requested features: MEP5430B39/0078/0765/0763/0771/0792 Type : Industrial pressure gauges MEX5-. Nominal size 100 mm - yes Class 1.0 - yes Measuring range: 0 - 600bar - yes Overload limit: min. 800bar - yes Protection class: IP 66 IP67 (0763) Connection G1/2a radial bottom - yes Medium: Oxygen medical O2 yes (0765) Temperature: -25 - +60C (outdoors) - -20 to 70C ambient T Wetted parts made of stainless steel 1.4571 y316L / 1.4404 Housing with pressure relief S3 with safety blow out back Multi-layer glass triplex safety glass Dial inscription: Oil and grease free EN conform marking Vacuum if low to industrial vacuum OK Oil+fat-free production; dried Standard with O2 cleaning Approved for use with oxygen Standard with O2 cleaning Leak-tested helium leakage rate <10E-7mbar l/s - Helium Lecktest 1.10-7 cm3/s (0792) Thread protection cap yes (0412) with damping , Installation immediately after a piston pump restrictor screw to dampen pulsation (0771), but not vibrations, for dampening vibration you can install it away with a capillary + pipe or wall mounting or add a BH5 filling (its the only compatible with O2 but extremely expensive
TCK52283 | User | Service Request | TR: Relance : AL-aT / Baumer - MAJ Offre 20527752 - Projet PRJ001737 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-11 15:38 | 2025-01-06 17:20 | 1 | 1 |  | Jacques, comme la discussion risque dêtre plus longue qu'une simple réponse, je clos le ticket et je t'ais renvoyé une question par mail. ce sera plus facile de suivre sur un fil de discussion. Best, Nicolas
TCK52137 | User | Service Request | FW: AKP_INQ_20241204_1630_PGNIG | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 17:40 | 2024-12-10 09:24 | 0 | 0 |  | here is the detail of both and what your customer is missing: [Image] your customer is missing: 10306619 x 1 10306641 x 1 if he didn't loose anything, you can also ask him to check the list to be sure he's got everything. then the decision for the sales and price will be up to the sales and also, make sure your customer if confident doing the tighening on it's own, if he's familiar with mounting flanges it should be OK.
TCK52116 | User | Service Request | Gauge Identification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 13:05 | 2024-12-10 14:13 | 1 | 0 |  | this doesn't seem like a vendome product, I'm thinking maybe india, I asked them if it can be their products. We can offer BHFR products to replace (MEX) but I doubt we can find out anything with the markings, markings on the dials are usually customer references.
TCK52114 | User | Service Request | Y91-3.2.3.B20.R/2152_2297_0409 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 12:43 | 2024-12-09 16:02 | 1 | 0 |  | yes the 2 options are available, but we have an issue on the 2152 which is our supplier is not delivering the screw to close the setting port on the side. so all options with settings (standard, 2151 and 2152) have no delivery date right now. but with the option without setting (2153) your configuration is possible: Y91-3.2.3.B20.R/2153_2297_0409
TCK52100 | User | Service Request | TBI100 problem | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Kopeć, Damian [kdam] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 10:48 | 2024-12-10 16:28 | 1 | 0 |  | I can confirm this is not normal, the process to attach the stem to the back of the case on a tbi is a pressing across a hole in the case and I checked our current production and machines, it should not turn. please start a claim process with your sales contact.
TCK52090 | User | Service Request | product recognition ¦ te2 ¦ order no: 179953 . 1 ¦ Request for offer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 10:06 | 2024-12-09 11:33 | 0 | 0 |  | the TE2 are not manufactured at BHFR, and when I check the order I find a completely different product, you need to ask BADK.
TCK52088 | User | Service Request | 80102522 / Anfrage Arrgos GmbH/ Atk. zu AG 20571716/ Thermometer + Mano / = neu AG 20608536 ( KD Logo + BH Logo möglich ) ? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Steinmann, Simone [ssim] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-09 09:58 | 2024-12-09 10:44 | 0 | 0 |  | On a MEX5, it'll be feasible, you'll have the Bourdon LOGO at the Bottom, the customer LOGO on top with the EN Marking and the ATEX on the left and right sides of the center.
TCK52064 | User | Service Request | DANC temperature drift figures | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-08 17:26 | 2024-12-09 10:39 | 1 | 0 |  | over 100C, we'll still respect the MEX3 caracteristics from the datasheet: [Image] so meaning that at 120C, the accuracy will remain within +/-2.4% FS
TCK52063 | User | Service Request | Vacuum chamber for seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-08 15:59 | 2024-12-09 09:53 | 1 | 0 |  | Yes it is safe (the top plug should be removed during the procedure so that the vaccum is the same inside and outside the tube). we don't advise to fit our gauge on diaphragm seals outside of our production and will take no responsability if it's not working afterward. But we know we have distributors doing it and we can't really prevent them from doing it.
TCK52046 | User | Service Request | SOLYRO --- 2411WP067 --- AROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-06 14:17 | 2024-12-06 15:13 | 0 | 0 |  | Vu par Teams
TCK52025 | User | Service Request | Demande de la SNCF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-06 10:27 | 2024-12-06 16:23 | 1 | 0 |  | Salut Cyrille, ce produit n'a pour moi jamais vraiment été vendu ou qualifié, il est bloqué dans le système depuis très longtemps. il faudrait une SPR pour évaluer la possibilité de le rouvrir. Best, Nicolas
TCK52012 | User | Service Request | TR: rfq fitting MEP5/0402/D820/27446-03-00 - BAVE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-06 08:57 | 2024-12-06 09:13 | 0 | 0 |  | Salut priscilla, c'est bon, il y avait juste une confusion pour le plan de l'anneau, le num de plan 27446 est seulement pour le bouchon. les codes ci-dessous sont avec 2 bouchons 1/2NPT MEP5-431.526/0078_Q008_Q003_1201_9941. 0402. D820-10L.2227.G2002 + 53102-01-00 + 2 x 27446-03-00 MEP5-431.527/0078_Q008_Q003_1201_9941. 0402. D820-10L.2227.G2002 + 53102-01-00 + 2 x 27446-03-00 MEP5-431.522/0078_Q008_Q003_1201_9941. 0402. D820-10L.2227.G2002 + 53102-01-00 + 2 x 27446-03-00 Best, Nicolas
TCK51889 | User | Service Request | total --- aroe --- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-04 14:32 | 2024-12-06 10:08 | 1 | 0 |  | ça ne me dis pas grand chose et a Jérôme non plus, ni la ref Total ni la ref fournisseur ne nous donnent plus d'info, il faut partir de la définition et de la req pour proposer quelque chose, désolé. Best, Nicolas
TCK51790 | User | Service Request | Qu'est-ce que c'est que ce type, MR MS ? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-03 15:29 | 2024-12-03 15:49 | 0 | 0 |  | Giorgio, this is an MG5 (you can see the contacts are green if you look closely and only the inductives are green. Best, Nicolas
TCK51751 | User | Service Request | YOKOGAWA COUPLES DE SERRAGE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-03 11:29 | 2024-12-06 16:32 | 0 | 0 |  | Salut Véronique, voici les SP qui sont dans les boites avec le matériel avec la procédure de serrage. pour info ils les ont dans le REQ et avec les produits. Best, Nicolas
TCK51626 | User | Service Request | MDF3BTN4 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-02 08:55 | 2024-12-02 15:10 | 1 | 0 |  | pas de soucis pour moi, il faut faire la demande de prix a Fantinelli et il faudra créer un cadran mais comme c'est du différentiel, faire un 0...600mbar ou un -600...0mbar ne change rien pour nous. fais une demande de prix pour le MFT avec les détails déchelle que tu m'a donné dans ta demande et ce sera bon.
TCK51604 | User | Service Request | 转发: Fw:回复: 3100200491 到货差异问题 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-12-01 14:18 | 2024-12-10 14:14 | 1 | 0 |  | Closed for now, keep me updated if you have any answer. Best, Nicolas
TCK51576 | User | Service Request | MDF3BTN4 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-29 13:34 | 2024-11-29 14:57 | 0 | 0 |  | Salut Véronique, Oui, dans la plupart des cas un MFT pourra remplacer. ci-joint la doc du MDF3 pour faire la conversion. Best, Nicolas
TCK51566 | User | Service Request | Norwegian Defence | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-29 11:33 | 2024-11-29 14:38 | 1 | 0 |  | the MAR5 was a MEX5 based product we use to sell for marine applications. we phased it out but there is a possible replacement with a MEX5 and some options. below is the slide that was presented by R&D to sum up how to do it. So for your replacement according to the elements in the email not from the pictures as the one on the pictures is different, I would offer: MEX5A30A24/0751_0151_0412_0840_0670_A009
TCK51548 | User | Service Request | [BHFR Material Data] product recognition ¦ 25J2013 ¦ 627192/10 -5117 ¦ ZAPYTANIE MWN TRADE 0070616 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] OF-Production Engineering Material Data FRVEN (OF-PE) | Dolbeau, Nadia [ndo] | 2024-11-29 09:42 | 2024-11-29 14:24 | 0 | 0 |  | it was a PBCN-1.3.B59.R.A1.14.02.2.1, but the product is phased out and we don't have capacitive cells anymore (it was supposed to be more robust but actually not really). here is 2 potential replacements the Y91 is analogic and the PP20S is numeric but both can be installed as replacement without any change to the customer process. Y91-3.2.3.B59.R/2164 PP20S-1.1B59R.A114.024100.004
TCK51547 | User | Service Request | DEFINITION PRODUITS POUR YOKOGAWA Arrêt Technique Chaudière D - 42LIS0402A / 42LIS0402B / 42LIS0402C / 42LI0401 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-29 09:40 | 2024-11-29 14:13 | 1 | 0 |  | on a peu déléments et tout n'est pas cohérent (une échelle de 50mbar pour une utilisation a 80 et en dessous une échelle de 1000 pour une utilisation a 50mbar ?? sinon ta proposition de D912 est cohérente pour des brides standard (on a pas de brides ovales) la longueur de capillaire sera a choisir dans la liste. si tu as des #900 il faudra passer en D918. pour le digit du montage avec le transmetteur ce sera F pour les étendues de mesure faibles (pour les 2 premiers) et P pour la dernière (à confirmer une fois les étendues de mesure clarifiées).
TCK51505 | User | Service Request | Electropolishing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-28 16:09 | 2024-11-28 16:43 | 0 | 0 |  | I don't think we specify anything, usually there is limits only for parts in contact with the fluids, not for the cases. I would answer him with this and add that the parts in contact can be Ra < 0.4 and that our mounting if with the right diaphragm seal and filling is 3A approved and the filling 3A approved. I've not seen yet norms for the roughness of the case (but maybe I'm missing something, if it's the case come back to me).
TCK51504 | User | Service Request | Old gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-28 16:09 | 2024-11-28 16:40 | 0 | 0 |  | for me it's an MB40ER2 sold to formerly OLAER (now SAFRAN I think) and it's mounted on planes landing gears. it should be this code 11265076 but I think they can purchase or replace only via the system supplier, we can't seel those gauges to anyone else than SAFRAN.
TCK51502 | User | Service Request | certificats étalonnage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-28 16:00 | 2024-11-28 16:29 | 0 | 0 |  | Salut Cyrille, tu peux demander a tes contacts adv, ils auront plus daccès a tes commandes SAP pour faire les verifs et me demanderont ou a la qualité au besoin. Best, Nicolas
TCK51438 | User | Service Request | MEP3, MEP6, MEP8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-27 21:40 | 2024-11-28 09:10 | 0 | 0 |  | those products are manufactured here in BHFR and you can order MEP3 and 5 right now. For 6 and 8, Jorg sent an email giving the dates at which we'll start codifications and sales (I'll forward it to you).
TCK51304 | User | Service Request | KULLMAN --- grch --- 26.11.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-26 15:25 | 2024-11-27 16:49 | 1 | 0 |  | sur l'article 11188919 la membrane Tantale recouvre 52mm sur une face RF de 68mm j'ai évalué plusieurs options mais aucune n'est viable en terme de prix car si on reprends une membrane plus grande il y a des frais d'outillage pour la presse de détourage, des frais de petite série et on doit redessiner la voûte, Idem sauf la nouvelle voûte si on demande une membrane plus large. et la dernière solution est un D820 tout tantale et ça coûtera extrêmement cher. désolé, je pensais pouvoir trouver une solution mais soit le recouvrement actuel est acceptable par le client selon le diamètre de son joint soit on a pas de solution.
TCK51262 | User | Service Request | 30253473--- Alcaraz Electrónica --- dionision | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-26 10:12 | 2024-11-26 14:39 | 1 | 0 |  | Pas de connexions sur ce produit et pas de remplacement existant, Si cest pour une différence entre la pression dans la chambre et une référence absolue, un MA7 pourrait marcher.
TCK51206 | User | Service Request | Blocage sortie TED - Steritech | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-25 15:24 | 2024-11-25 16:22 | 0 | 0 |  | tu avais déjà eu des retour de l'usine quand tu as signalé ce problème ? (des emails ou une NQ etc...) je ne vais pas pouvoir traiter ça par le système des tickets (il est très bien pour les réponses directes mais pour ce type de problème il faudra surement impliquer de multiples acteurs) je ferme le ticket mais réponds moi avec les éléments précédents de cette affaire sur ma boite Baumer (nbirrien@baumer.com) et je vais voir ce qu'on peut faire au mieux.
TCK51166 | User | Service Request | 1000947171 - capteur PT0401 (QUO-58784) / 20581312 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-25 10:17 | 2024-11-25 16:15 | 1 | 0 |  | les infos de pression et température (12bar et 400C et 10.5bar et 180C dans la datasheet client sont en accord avec le montage). Par contre pour un gaz avec de multiples composants, même si je regarde la compatibilité de chaque gaz, les combinaisons peuvent avoir de multiples effets donc seul l'utilisateur qui gère ce process peut connaitre la compatibilité.
TCK51164 | User | Service Request | solyro --- aroe ---- TBI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-25 10:11 | 2024-11-26 10:56 | 1 | 0 |  | c'est bien le cadran standard sur un TBI 250C, le mouvement du bimetal n'étant pas linéaire le cadran ne l'est pas non plus, mais c'est vrai que ça doit être celui sur lequel ça se voit le plus. ci-dessous un cadran du stock pour TBI100: [Image]
TCK51156 | User | Service Request | PRODEVAL--- MPG --- AROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-25 09:29 | 2024-11-25 15:54 | 1 | 0 |  | il y a un peu a boire et a manger dans la demande de ton client, par exemple "US" quelle norme US ? ou Hazloc qui est une terminologie generique non specifique a une norme et ensuite on revient sur class et div qui viennent de l'ATEX. tu as de la chance, j'ai fais un decryptage complet de notre marquage atex pour un client de Jerome qui devrait clarifier notre compatibilité ATEX/IECEX (et nous ne sommes pour l'instant pas UL) Pressure gauge Marking: [Image] Breakdown of the Elements General marking: II 2 G D IM2 II: This indicates that the equipment is intended for use in explosive atmospheres other than those occurring in coal mines (Group II). 2: This means the equipment is suitable for use in areas where an explosive atmosphere may occasionally occur (Category 2). G: Indicates that the equipment is suitable for use in gaseous atmospheres. D: Indicates that the equipment is also suitable for use in dust atmospheres. IM2: Refers to the inspection module, indicating that the equipment has been assessed for specific conditions (suitable for use in zones 1 and 2 for gas atmospheres, as well as zones 21 and 22 for dust). Gas details: Ex h IIC T6 Gb Ex: Indicates that the equipment is intended for use in explosive atmospheres. h: This means that the equipment utilizes a method of protection by increased safety (h protection). Increased safety techniques prevent the formation of sparks or electrical arcs. IIC: This denotes the gas group. "IIC" means the equipment is suitable for use in atmospheres that may contain more hazardous gases than group IIB (such as hydrogen, for example). T6: This classification refers to the equipment's maximum surface temperature. "T6" means that the surface temperature of the equipment does not exceed 85 C, which is necessary to prevent ignition of explosive gases. Gb: Indicates that the equipment is rated "b" for the level of protection in the zones it is used. Dust details: Ex h IIIC T85C Db Ex: As above, this indicates that the equipment is designed for use in explosive atmospheres. h: Confirms the use of a method of protection by increased safety for dusts. IIIC: Indicates that the equipment is suitable for atmospheres containing combustible dust. T85C: This indicates that the maximum surface temperature of the equipment does not exceed 85 C, which is essential to prevent ignition of dust. Db: Indicates that the equipment is rated "b" for the level of protection in dust.
TCK51153 | User | Service Request | TR: Article bloqué : 11161689 _RPPY-7S3.209.XDSE:52660-04/LRS4_SETP / AR 2291895 ARKEMA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-25 09:20 | 2024-11-25 16:53 | 1 | 0 |  | vu par Teams, ancien fichier d'Hervé pour remplacement des 52660-02 par 53337-01 la demande de validation avec nouveau plan sera fait au client et le code selon nouveau plan créé.
TCK51047 | User | Service Request | low temperature test report of CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-21 16:03 | 2024-11-29 15:09 | 1 | 1 |  | we do not share our internal test reports, but I can tell you the -40C limit is because of the limits of the electronic parts and of the Gasket.
TCK51008 | User | Service Request | ARIN-220L | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-21 11:41 | 2024-11-21 16:46 | 1 | 0 |  | [Image]avec diam = 27 HEX et longueur totale = 45mm [Image] [Image]
TCK50992 | User | Service Request | Création BAFR // Client : LA ROBINETIQUE // Cde Clt : 713686 // RPPN-7A3.207/SETP // Position 10 | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-21 10:38 | 2024-11-21 14:51 | 1 | 0 |  | Bonjour Marie claire, je pense que c'est une erreur, l'email pour les tickets de Nadia et josette est le: FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> Best, Nicolas
TCK50987 | User | Service Request | ZPN--204CHM - 10302278 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Gitau, Kenjie [gik] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-21 09:39 | 2024-11-21 14:49 | 1 | 0 |  | I cannot give pricing, I'm handling the technical support and cannot handle pricing. for prices, please contact your sales administration contact.
TCK50970 | User | Service Request | MTA2-B20.N08 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-21 08:02 | 2024-11-21 14:47 | 1 | 0 |  | Véronique, ce plan 3D n'est pas existant mais il peut etre commandé en tant que document technique au besoin. Best, Nicolas
TCK50940 | User | Service Request | SWAGELOCK 1/4 PULG( AQ), | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-20 15:15 | 2024-11-21 14:54 | 1 | 0 |  | Angel, we are currently working toward supplying those kind of tubular connections, the first ones will be based on parker design and should be available next year. for swagelock we have nothing confirmed yet (we also have to investigate on the compatibility). but if you have some qty and could take the business, don't hesitate to start an SPR for a special product to be created. Best, Nicolas
TCK50928 | User | Service Request | 11107235 - Changement nomenclature | Service Desk > General | Low |  |  | Closed | Pletikos, Maurane [plem] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-20 13:55 | 2024-11-21 14:46 | 1 | 0 |  | Maurane, tu as envoyé ce ticket a l'IT qui l'a renvoyé mais c'est pour Nadia, j'ai essayé de lui transférer mais ça ne marche pas, tu peux lui renvoyer a func-frven-mat-dat@baumernet.org stp ? Best, Nicolas
TCK50850 | User | Service Request | à propos du E763 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-19 16:22 | 2024-11-19 17:27 | 0 | 0 |  | voici la datasheet, on a pas d'equivalent en terme de montage mais pour le reste c'est basé sur un Y91 ou Y92 si il peut faire un montage direct.
TCK50624 | User | Service Request | Manual of storage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-15 23:09 | 2024-11-18 09:54 | 1 | 0 |  | here are the package and storage reccomandations.
TCK50369 | User | Service Request | LRS1 / FDA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-13 17:19 | 2024-11-13 17:45 | 0 | 0 |  | Sure, here you go
TCK50318 | User | Service Request | RPPY-3C3.103 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-13 13:04 | 2024-11-13 15:10 | 1 | 0 |  | the standard is the PG9 as below: [Image] and there is an option for M20 adaptor: [Image]
TCK50232 | User | Service Request | RT2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-12 16:25 | 2024-11-13 11:13 | 1 | 0 |  | it's complicated to be sure but I tried to go with the closest I can from what I visually see. but the customer will have to check: Order code: RT2-YM.43.230C3 RT2 old Code / SAP code:
TCK50175 | User | Service Request | IGC for wetted parts | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-12 12:22 | 2024-11-13 11:06 | 1 | 0 |  | we don't do intergranular corrosion test, we work with normed stainless steel that have specific properties and 3.1 material tests that state the composition and mechanical main properties.
TCK50173 | User | Service Request | Gold Coated Diaphragm Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-12 12:16 | 2024-11-13 11:03 | 1 | 0 |  | here is the code you need for the diaphragm seal: D82010N2259Q27 please mention as an example for the codification and pricing the code 11170633 (D820-10N.2267.Q27) which is the same with a different Class/Rating. the final mounting code will be: MEX5-D60.H21/0751_9941 + AORP0382/SETP + ARIN2606 + D82010N2259Q2709 I changed the AORP for a lower cost option that is able to close at 7bar (max of the gauge) and protect up to 400bar and added LRS9 filling which is the standard for this type of mountings.
TCK50051 | User | Service Request | PRESSURE GAUGES WITH 2 CONTACTS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-11 10:47 | 2024-11-12 12:15 | 1 | 0 |  | those are from the Bourdon Haenni time and can be replaced by a MR5: please still ask them to check the cutting powers and generic features as small changes might have happened since 30years ;) 1] Pressure Gauge with 2 contacts (1MVA) Make- Bourdon Haenni Type- M1256D30B72 MR56D30B72 2] Pressure Gauge with 2 contacts (3.15MVA) Make- Bourdon Haenni Type- M256730B74 (the 1 after the M is probably missing and the 7 makes no sense so I base my code more on the structure of the first code) MR56D30B74
TCK50013 | User | Service Request | RDPNEA8110 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-10 14:32 | 2024-11-18 09:19 | 1 | 0 |  | then it's confirmed, it's this code and we can still provide it: 11050079 RDN-4E8.110 but the Souriau connection adaptor was added on their side and is not included in our delivery.
TCK49999 | User | Service Request | Mise à jour prix CCR2 article 11120399 | Service Desk > General | Low |  |  | Closed | Sa, Antonio [santo] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-08 15:16 | 2024-11-12 10:31 | 1 | 0 |  | Bonjour, erreur de destinataire je pense. Best, Nicolas
TCK49948 | User | Service Request | Absolute PG | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-08 08:19 | 2024-11-08 09:44 | 0 | 0 |  | here are my answer and our MA7 is +/-2%. also you could check the P2F or datasheet, it answers easily 90% of your question. [Image] My Answers: 1) options for O2 cleaning available, 0765 (it'll be clean, calibrate only with AIR and the dial will have the sign "use no oil". 2) helium leak test option available 0792 (1 x 10-7 atm cm3/s) 3) Q003 4) Calibration before and after overpressure is part of the EN qualification so it was done during qualification of the product. in production phase now, they are overpressured to their max over pressure (selected in the code) at the end of mounting before calibration. we don't do 30min vaccum test but with the helium leak test you already ensure this.
TCK49873 | User | Service Request | FW: Counterfeit Certificates of Conformity (CoC) for Titanium Raw Material within the Supply Chain | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-07 12:33 | 2024-11-12 10:27 | 1 | 0 |  | Answered via email
TCK49872 | User | Service Request | ancienne commande à renouveler | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-07 12:29 | 2024-11-07 17:17 | 1 | 0 |  | Salut cyrille, voici tous les details du poste 80: [Image] Best, Nicolas
TCK49859 | User | Service Request | Création BAFR // Client : 3C France // Cde Clt : FDA 104558 // MEX3-F21.B79.DANC07.2H2B.01/0753 // Position 10 | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-07 11:13 | 2024-11-07 17:18 | 1 | 0 |  | Bonjour Marie Claire, ce n'est pas pour moi mais pour FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> merci de leur renvoyer stp Best, Nicolas
TCK49858 | User | Service Request | MPG6DX1B22+FCM12A8T600 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-07 11:12 | 2024-11-07 17:08 | 1 | 0 |  | it's from a project for saudi aramco, you probably have the right reference, but there is big chances that Jerome already have the order directly from them. I'll ask him to confirm.
TCK49829 | User | Service Request | AORP Temperature Limits / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-07 08:35 | 2024-11-12 11:16 | 0 | 0 |  | the temperature of -30.8C is ok to work with the AORP.
TCK49752 | User | Service Request | Incoterms changeover to EXW as of 01.12.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Laboisse, Eric [leri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-06 09:18 | 2024-11-06 17:53 | 0 | 0 |  | Eric, un changement de conditions generales chez Schneider, je ne sais pas pourquoi c'est arrivé chez moi. Best, Nicolas
TCK49739 | User | Service Request | [BHFR Material Data] MMD4 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] OF-Production Engineering Material Data FRVEN (OF-PE) | Dolbeau, Nadia [ndo] | 2024-11-06 08:08 | 2024-11-07 09:23 | 1 | 0 |  | Vu sur TCK49321, pas de datasheet existante.
TCK49738 | User | Service Request | recognise ¦ Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-06 07:53 | 2024-11-06 10:04 | 1 | 0 |  | here are the codes of those 2 components and a picture for reference and we have stock. but only sales will be able to tell you if they can sell it as spares and the costs or MOQ etc... A 05398.9061 A 5398-D8-M12X1-(44)A MUTTER SDI 439 D=8 10323992 A 05399.9040 SDI 439 [Image]
TCK49734 | User | Service Request | Enquiry From Neste | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-06 07:33 | 2024-11-06 09:29 | 0 | 0 |  | yes the mounting is fine but you need at least a small ASIPD inbetween, otherwise the temperature will be too hot for the gauge.
TCK49733 | User | Service Request | MIX7+DT3/DT2 / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-06 07:04 | 2024-11-06 09:23 | 0 | 0 |  | the range is ok and require no overcost even if not existing yet as it's a standard in the configurator / P2F. concerning the rest of the configuration, the dashpot is not possible on the MIX7 with such a low range. otherwise the rest is OK and the fitting with a DT3 is fine.
TCK49683 | User | Service Request | FW: RE: ELCON QUESTION temp switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-05 10:44 | 2024-11-05 16:20 | 1 | 0 |  | Code for the blue ones is 10302465 and you might need to replace also the aluminium part that goes with it (only if lost or damaged) 10302474.
TCK49660 | User | Service Request | TECHNIPIPE - recherche produit | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-05 08:57 | 2024-11-05 10:34 | 0 | 0 |  | Vu en teams 11154626
TCK49638 | User | Service Request | MEX3 + M12 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-04 18:06 | 2024-11-05 10:42 | 1 | 0 |  | the drawing exists, there is no digit for it and it's not in the configurator but with an SPR and the ref of the drawing (57206-14) you can get a price: [Image]
TCK49604 | User | Service Request | diaphragm separator 1005 for homogenizer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-04 12:52 | 2024-11-05 10:44 | 1 | 0 |  | yes, we can still offer 1005 + gauge, please find enclosed the latest datasheet to define your product.
TCK49589 | User | Service Request | PROBLEME DEFINITION TXA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-04 11:14 | 2024-11-05 11:23 | 1 | 0 |  | Oui Véronique, cétait une époque ou on proposait des parties de capillaire nu et l'autre revêtue. il faut donc proposer capillaire = Lcp + Lcn (= 4.5m pourt le 1er) en revêtu je pense et proposer les raccords de la page 4 et 5 de la datasheet des TSF actuels. Best, Nicolas
TCK49562 | User | Service Request | MTX5-F60.N09+ARA10D60N+ARPX06N / APS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-11-04 09:27 | 2024-11-05 11:27 | 1 | 0 |  | yes everything is fine and I see only M/F 1/2NPT for the fittings so it's fine.
TCK49487 | User | Service Request | [BHFR Material Data] Data Sheet on EA4-AP3 and New typ for this. | Process Instrumentation > Material Data Management | Low |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] OF-Production Engineering Material Data FRVEN (OF-PE) | Deschamps, Josette [jau] | 2024-11-01 09:33 | 2024-11-05 11:30 | 1 | 0 |  | the phase out was define on this product so you need Mario Marolf approval for reopening it but I don't see reason that technically block otherwise. once you have the approval to reopen the code from Mario, you can ask the databese for opening and you're fine.
TCK49452 | User | Service Request | Chiffrage TED sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-31 16:10 | 2024-11-05 14:17 | 1 | 0 |  | je ne fait pas les prix (ADV) mais pour le montage ne posera pas de problème, et le D295 sera un des 2 ci-dessous (3" et 4"): 11202176 11204989
TCK49443 | User | Service Request | PROBLEME RACCORDS EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-31 15:00 | 2024-11-06 11:29 | 1 | 0 |  | les montages de filets coniques sur du parallèle ne sont pas bon, je propose de faire plutot un ARIN special avec un des 1/4BSP (a valider avec lui): [Image][Image]
TCK49440 | User | Service Request | replacement ¦ PULSATION DAMPENER EG.CL.2BM.2BF.SG ¦ tabliczka ¦ Prośba o ofertę | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-31 14:25 | 2024-11-05 11:46 | 1 | 0 |  | this is an indian pulsation dampener similar to ARA10 but in 1/4BSP M/F. we don't have this connection on ours but we have 1/2G M/F if acceptable (ARA10D30L/0765)
TCK49430 | User | Service Request | Fw: RE: ELCON QUESTION temp switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-31 13:10 | 2024-11-05 16:21 | 1 | 0 |  | Double with TCK49683
TCK49410 | User | Service Request | GRCH --- ALSTOM Crespin HY00000030994 HY00000031014 ---- 31/10/2024 --- urgent !!!! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-31 09:27 | 2024-11-05 10:25 | 1 | 0 |  | comme regardé ensemble a priori c'est 10263821 et 10263820.
TCK49380 | User | Service Request | TR: RELANCE DP 6000069858 BAUMER pour ARKEMA JARRIE ARTICLE 1106362 MANO P 2,5 1/2GAZ BAS D100 avec PLAN JOINT J98-04 fabrication speciale | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-30 16:25 | 2024-11-05 17:47 | 1 | 0 |  | Pour moi le séparateur utilisé est un 10266178 (plan ci-joint) et voici les articles existants dans SAP en monté sur : 11141719 MEX5-D30.B22.XDSE:52642-07/LRS1 11141781 MEX5-F31.B19.XDSE:52642-07/LRS1 11146106 MEX5-D30.B20.XDSE:52642-07/LRS1 11153120 MEX5-D30.B22.XDSE:52642-07/LRS1_0078 11158619 MEX5-D30.B19.XDSE:52642-07/LRS1 11165726 MEX5-F31.B20.XDSE:52642-07/LRS1 11209530 MEX5-D31.B22.XDSE:52642-07/LRS1 11219409 MR5-7D30.B19.XDSE:52642-07/LRS1 11228188 MEX5-F31.B22.XDSE:52642-07/LRS1 11239413 MEX5-D31.B20.XDSE:52642-07/LRS1 11242166 MIX7-D30.B22.XDSE:52642-07/LRS1 Best, Nicolas
TCK49350 | User | Service Request | PROBLEME RACCORDS EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-30 10:57 | 2024-10-31 14:12 | 1 | 0 |  | le joint plat alu des raccords G ne doit pas rentrer dans un trou, il doit se mettre sur le fond plat du raccord femelle.
TCK49321 | User | Service Request | MMD4 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-30 08:53 | 2024-11-05 14:10 | 1 | 0 |  | j'ai du demander a Jörg pour ça et le MMD4 n'est pas officiellement au catalogue, il faut partir sur un DR 80 pour l'instant.
TCK49307 | User | Service Request | M21+D820 / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-30 07:07 | 2024-11-13 15:15 | 1 | 1 |  | those mountings are acceptable: M21-2P3.B19H+D820-GEL.2217.G20.09 (2 pcs) 2" class 150 4bar diff / 10bar stat M21-2P3.B19H+D820-GEL.2219.G20.09 (2 pcs) 3" class 150 4bar diff / 10bar stat
TCK49247 | User | Service Request | ARIN-220L | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-29 11:21 | 2024-10-29 14:47 | 1 | 0 |  | oui le filetage est du même type mais le 1/2G a un téton contrairement au 1/4BSP ce qui permet de centrer le joint. mais ils peuvent se monter en lieu et place.
TCK49245 | User | Service Request | Block and Bleed Valve | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-29 10:59 | 2024-10-29 14:45 | 1 | 0 |  | yes, the AMFD will do this function, our standard is just slightly more evolved with 2 valves to control both the block and the bleed.
TCK49231 | User | Service Request | TR: Etiquette client sur CTX - JACQUES Apaloo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-29 09:38 | 2024-10-29 15:58 | 1 | 0 |  | c'est techniquement possible, c'est une étiquette imprimée chez nous, on pourrait par exemple avoir le logo du client et un code produit de chez lui. par contre il faut quand même qu'on puisse faire la traçabilité et que les infos critiques soient dessus. l'acceptation et potentiellement la définition du prix et de l'option sont a la discrétion du PM, il faut donc faire une SPR et lui envoyer pour accord puis a Stéphane pour accord prix.
TCK49212 | User | Service Request | MCX5 0 TO 0.25 kg/cm2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-29 06:56 | 2024-10-29 09:25 | 0 | 0 |  | kg/cm2 is the digit F and it exists in 7 codes currently with recent ones,no need for creation. MCX5-XXX.F10
TCK48837 | User | Service Request | TR: MFT question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-28 09:51 | 2024-10-28 16:23 | 1 | 0 |  | j'ai eu Fantinelli, ils n'ont pas ce type de connexion.
TCK48832 | User | Service Request | DEFINITION PRODUIT EDF UTO Z811B328 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-28 09:01 | 2024-10-29 09:36 | 1 | 0 |  | j'ai cherché dans ma boite mail, ticket, docs et projets et je trouve seulement une fois ce code. c'est sur une demande d'obsolescence de westinghouse il y a quelques années et on a répondu après consultation des responsables nucléaires que le produit était inconnu. après on avait seulement le code MA2BH100B sur la demande, avec le code client Z811B328 si UTO a commandé ce code tu peux regarder dans SAP (j'ai essayé et je ne trouve rien en commande de chez nous mais je n'ai pas accès à vos commandes). et la dernière solution et de le redéfinir avec eux en partant d'un MEX5D30B39
TCK47811 | User | Service Request | DE DELAI | Service Desk > General | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-25 13:00 | 2024-10-25 15:52 | 0 | 0 |  | erreur de destinataire
TCK47806 | User | Service Request | Pressure Switch Identification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-25 12:15 | 2024-10-25 14:29 | 1 | 0 |  | yes, this is a UT from Baumer Indian, and I'm as surprised as you are by the Bourdon LOGO, I saw some UT regularly and it was always Baumer Branded. I warned Jorg already, he'll check this out.
TCK47803 | User | Service Request | MEX5-D62.H25/0751 /Del Monte | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-25 11:42 | 2024-10-25 15:24 | 1 | 0 |  | I created a file with the dial scalings but it's in validation since years, anyway, here is the one you looking for: [Image]
TCK47801 | User | Service Request | AKPLBF5X02 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-25 11:09 | 2024-10-25 15:21 | 1 | 0 |  | je pense que c'est possible, la question va être le prix, mais tu peux demander en basant sur ce code 11211184 et on proposera la même connexion avec ton type de protection et en 2m au lieu d'1.
TCK47751 | User | Service Request | TR: MA7 mounted to DS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-24 16:20 | 2024-10-24 17:43 | 0 | 0 |  | Le MA7 est basiquement un differentiel et a donc un souflet avec beaucoup de volume, il faudra au moins un DT2 pour faire ce montage. Best, Nicolas
TCK47748 | User | Service Request | TR: Vietso Petro Project | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-24 15:48 | 2024-10-25 15:34 | 1 | 0 |  | je proposerais des: 53192-25AS330N (316L ANSI B16-5 NPS 2" Class900/1500 RTJ 2x 1/2NPTF) 53192-25AS370N (316L ANSI B16-5 NPS 2" Class2500 RTJ 2x 1/2NPTF) 53192-25AS060N (316L ANSI B16-5 NPS 2" Class150 RTJ 2x 1/2NPTF) un code existe similaire mais en 1"1/2 au lieu de 2, pour exemple: 11196866 les prix et codifs seront à créer
TCK47721 | User | Service Request | Aarti Industries - Project Coral 2241210181 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-24 12:37 | 2024-10-24 17:38 | 1 | 0 |  | I can help you if you have an issue, but I can't make the RFQ for you, it's not a simple one I understand but most data are avalaible and the rest can be set as standard (the type of microswitch for example) otherwise it seems to be RPPY with Diaphragm seals. If you have a specific question or something you want me to check don't hesitate.
TCK47705 | User | Service Request | ENQT 44986 - T-5502-1001 (JOHNSON CONTROLS) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-24 11:00 | 2024-10-24 17:31 | 1 | 0 |  | non, désolé, on a pas ce type d'appareil.
TCK47654 | User | Service Request | WEP5 assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-23 16:44 | 2024-10-23 17:42 | 0 | 0 |  | every gateway on the market today that covers the lorawan region you're in is fine on our side (this one covers all regions) the choice of the gateway and what makes each one specific is the use you want to make of it, this one seems a good choice for indoor non atex zones and if you want to transmit your signal received via LoRa in the gateway to your server via wifi or ethernet.
TCK47608 | User | Service Request | Vietsopetro Project/ MEP5-AORP-ARIN-ASIP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-23 11:46 | 2024-10-23 17:33 | 1 | 0 |  | I see one issue in the mounting, if the gauge is filled, the customer can't open it to set the pointer so you can't have this option. otherwise everything is fine and here is the limits of usage of the ASIP according to the norm (a schedule 80 would fit your need better). otherwise, considering the temperature drop it'll bring to the measuring elements on top, it's impossible to say on this type of ASIP as it depends completely on the fluid and it's properties. it's up to the piping design engineer to define it. sorry I can't help more on this but we can't know the use case better than the customer.
TCK47602 | User | Service Request | TR: Question manomètres | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-23 11:15 | 2024-10-23 14:07 | 1 | 0 |  | le CSA et l'UL sont des normes de produits électroniques donc non applicables aux MTA.
TCK47551 | User | Service Request | TR: Quote - MJX with D825 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 16:31 | 2024-10-24 17:39 | 0 | 0 |  | Erreur dans la demande, géré en direct par Jérôme
TCK47540 | User | Service Request | MIX7 + DT3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 15:27 | 2024-10-22 17:49 | 1 | 0 |  | your mounting is fine but will withstand only up to 200C (with a compatible LRS) we did a special option on DT2 for graphite gasket that could fit your need but it means a min pressure range for the mounting of 1bar and a request for price and availability. the code would be exemple with standard all stainless DT: MIX7-A33.D15 + AKPL (2 metres) + DT22N20026009/1803
TCK47526 | User | Service Request | CLAUGER- ADV--- 22.10.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 14:00 | 2024-10-22 17:31 | 1 | 0 |  | le NH3 est utilisé comme fluide frigo avec beaucoup de nos appareils, on ne garantie jamais les tenue chimiques car on en connais pas les proportions ou les conditions mais on ne vois pas de contre-indication vu les éléments fournis. 11162499 --- MZ7-A30.B26K/0761_0751 la tenue en temperature entre -40 et 40C process est OK (datasheet) la tenue en pression est ok entre 0 et 25bar diff et jusqu'a 25bar statique des 2 cotés. le certificat de conformité EN (contenant la DESP) est joint ainsi que le manuel.
TCK47512 | User | Service Request | RP2's | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 12:39 | 2024-10-22 17:21 | 1 | 0 |  | here are my answers: Compressor control - RP2-NR.354/SETP set to 8 bar falling and reset at 10 bar OK Compressor acknowledgement - RP2-NR.346/SETP set to 7 bar falling and reset at 7.5 bar 7 and 7.52 OK Drive inhibit - RP2-NR.346/SETP set to 5.5 bar falling and reset at 6.5 bar OK Bogie 10% - RP2-NR.345/SETP set to 5.3 bar falling and reset at 5.6 bar Not OK you should consider a 10bar (46) that will allow you something like 500mbar of H in the middle of the scale (5.2 and 5.7 will work). Bogie 90% - RP2-NR.342/SETP set to 0.5 bar falling and reset at 0.6 bar Not OK, 130mbar min Hysteresys
TCK47510 | User | Service Request | à propos des TED Flush | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 11:59 | 2024-10-22 16:56 | 1 | 0 |  | bien sur, en standard c'est du LRS1, la FDS est attachée au ticket. Best, Nicolas
TCK47498 | User | Service Request | à propos du TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 10:56 | 2024-10-22 16:05 | 1 | 0 |  | je suis désolé mais nous ne donnons pas nos qualifications internes car elles font états des tests de nos composants. notre produit est conforme aux normes européennes et à sa datasheet et nous sommes audités régulièrement par les ISO pour notre système d'assurance qualité.
TCK47494 | User | Service Request | TR: OC 2233985 / 50 - BADK | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pétré, Sandra [sape] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 10:33 | 2024-10-22 10:44 | 0 | 0 |  | Voici tous les documents sauf le GMP EC 2023/2006 mais ca ne doit pas etre obligatoire, c'est un guide de bonnes pratiques. Best, Nicolas
TCK47448 | User | Service Request | Vietsopetro Project | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-22 06:08 | 2024-10-23 14:23 | 1 | 0 |  | this mounting is possible and willcut the pressure at the setp of the AORP and then resist up to 400bars. the ARIN being just 2 x 316L 1/2NPT connections, they resist up to 1600bar
TCK47433 | User | Service Request | cmj-ts --- nih--- 21.10.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 16:04 | 2024-10-22 10:37 | 1 | 0 |  | oui, pour moi c'est un fluide hydraulique aéro, mais sur les offres que j'ai faite sur du Skydrol, le 316L était accepté. Best, Nicolas
TCK47421 | User | Service Request | MFT5 Dual Scale | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 15:03 | 2024-10-21 15:44 | 0 | 0 |  | you can configure it in the P2F I think, it'll be a: MFT5-113.x08N.62/9712 but there might be a cost for the dial creation if not existing, you need to ask your sales administration contact a confirmation for the price.
TCK47419 | User | Service Request | MDF3BTN4 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 14:18 | 2024-10-22 10:34 | 1 | 0 |  | Bien sur, la voici: Best, Nicolas
TCK47405 | User | Service Request | AIR LIQUIDE VALIDATION MEP3/5/6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 12:17 | 2024-10-24 17:47 | 1 | 0 |  | les tests des MEP6 sont avec les MEP5 (rapport de validation labo) les plans des raccords sont envoyés ainsi que des 3D Best, Nicolas
TCK47393 | User | Service Request | [BHFR Material Data] Ticket für ABB HC-Schalter | Process Instrumentation > Material Data Management | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] OF-Production Engineering Material Data FRVEN (OF-PE) | Dolbeau, Nadia [ndo] | 2024-10-21 11:15 | 2024-10-24 17:45 | 1 | 0 |  | I'll need a bit more information like the sap code of the product and a bit of history of the creation of the code and what was validated to be able to help, can you please send those information to my personnal baumer mail box ? it'll be easier to treat from there as I will need other people on this subject.
TCK47387 | User | Service Request | TR: FDA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 10:54 | 2024-10-21 13:56 | 1 | 0 |  | je pense que tu parles de la certification 3A qui est aussi une norme alimentaire (relative aux montages aseptiques et au conditions de nettoyage). l'approbation FDA par contre est liée au fluides qui peuvent entrer en contact avec des process sanitaire ou agroalimentaires et c'est notifié dans la FDS du fluide. par contre dans ton cas, il n'y a aucun fluide et pas non plus de montages sanitaires donc aucune de ces normes n'est applicable.
TCK47374 | User | Service Request | MEX5-D30.F29.D824-10L.2227.G70/LRS2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-21 09:45 | 2024-10-21 14:04 | 1 | 0 |  | Yes, it is possible. Best, Nicolas
TCK47304 | User | Service Request | ORGAPHARM-GRCH -18.10.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-18 10:10 | 2024-10-18 16:44 | 1 | 0 |  | pas de soucis pour moi, un -1...1.5 sur un D435 DN20 PN25 avec revetement membrane et partie inferieur en PTFE. alternative en D631 avec la partie basse completement en PTFE.
TCK47263 | User | Service Request | RPPN-7B3.204/0407+AKPL-BF5305+DT2-2L60.0.F60.09 / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-18 04:02 | 2024-10-18 09:21 | 0 | 0 |  | it is ok, I did one slight change, the cappilary internal diameter is the one for standard mounting, not for differential. so here is your mounting code: RPPN-7B3.204/0407+AKPL-2F5305+DT2-2L60.0.F60.09
TCK47239 | User | Service Request | step file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-17 13:26 | 2024-11-05 09:41 | 1 | 1 |  | Sorry it was so long, here is the 3D file of a MEP5 4Xxxxx (OD12 25mm).
TCK47196 | User | Service Request | presotats ¦ R744 ¦ Zapytanie - presotaty na CO2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-17 07:27 | 2024-10-17 16:22 | 1 | 0 |  | PZH is a polish health, and we work with european norms, not local. WRAS doesn't exist anymore since 2021 None of our products are SIL, concerning safety requirement, we work with ATEX, PED and EN837. concerning the use with his product, we can't confirm compliance of our products with the infinite variety of fluids our customer work with, by working with R744, they must have rules from their design engineering team (based on MSDS of the R744 from their suppliers for example).
TCK47166 | User | Service Request | Pressure Switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-16 12:56 | 2024-10-17 16:26 | 1 | 0 |  | this reference is not from Baumer, and we don't have monel switches. Otherwise, with those deviations, I think an RPPY could be a fit.
TCK47146 | User | Service Request | TR: Datasheet glycerine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-16 10:38 | 2024-10-16 11:42 | 0 | 0 |  | MSDS sent
TCK47038 | User | Service Request | MZ7 - 133 - B10 - J | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-15 14:04 | 2024-10-16 11:38 | 1 | 0 |  | in a similar product, I found out the SAP code was blocked with a message. it seems like in 2018 we had an issue and can't supply the necessary belows from our supplier to make the 10J anymore. I'm sorry it must have been blocked from configurators but not changed in the datasheet (I'm warning Jorg he needs to change it).
TCK47017 | User | Service Request | MVE7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-15 10:56 | 2024-10-15 11:38 | 0 | 0 |  | for the 600psi/kPa, there is a code existing that matches: 11092743 MVE7-D60.V27 concerning the other one I think it was this code but it's blocked you need to ask for it to be reopen: 11092744 MVE7-B60.X41/9712_D000
TCK46991 | User | Service Request | FW: RFQ For E912 / Dylan | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-15 08:03 | 2024-10-15 09:16 | 0 | 0 |  | the order number is too old to find something but based on the informations I can get from the pictures, I would offer Y91223B33R but I can't tell if there was not visible features or options,please check with the customer.
TCK46975 | User | Service Request | need article# for spare glass MEP3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-14 16:26 | 2024-10-14 16:43 | 0 | 0 |  | here is the code of the triplex glass for MEP3: 11179238 VOYANT TRIPLEX MEP3 60.9X3.1MM
TCK46943 | User | Service Request | TR: RT2-EP.45.C0XC6, X=320 mm / Henry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-14 12:09 | 2024-10-14 16:29 | 1 | 0 |  | Nathalie, c'est OK, il y a des codes similaires avec des longueurs plus faible donc pas de soucis. Best, Nicolas
TCK46866 | User | Service Request | FW: 1600 Bar High Pressure | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-11 18:20 | 2024-10-16 13:52 | 1 | 0 |  | we never standardized this connection but I have several customer specific drawings of it, the most standard to us is the one below. please tell us is it's what you need and we'll create a Cxxx number for it. 57206-173 (9/16-18 UNF 2B Female)
TCK46851 | User | Service Request | TR: PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-11 13:56 | 2024-10-11 14:51 | 0 | 0 |  | le joint sur son schema est un joint alu pour le montage des connections G1/2 et il n'est pas fourni avec le PBSN (mais on en a en stock si besoin de lui en vendre un sachet) le joint NBR qui est choisi dans la codif est le joint de cellule et il y est forcement mais le client n'y a pas acces, il est à lintérieur.
TCK46835 | User | Service Request | 2 x switch assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-11 11:58 | 2024-10-11 15:06 | 1 | 0 |  | yes you can, both switches are SPDT and are cabled separately so you can cable it with 3 wires and it'll always be closed between 1 way and the common or you can cable it with 2 wires as NO or NC (and you can do this on each switch independantly).
TCK46834 | User | Service Request | DANC -- HASTELLOY --- BOCCARD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-11 11:57 | 2024-10-11 14:45 | 1 | 0 |  | Vu au telephone, la membrane en HC276 est une bonne idée pour augmenter la resistance du DANC a l'HCL.
TCK46824 | User | Service Request | WG: Ihre Anfrage KIT Campus Süd Geb. 20.20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-11 10:59 | 2024-10-11 14:47 | 1 | 0 |  | this one is stopped, but the purpose of this product was to cover small differential with High static and the MFT is our product for this application.
TCK46775 | User | Service Request | I: BAIT - TRIVENETA IMPIANTI - RP2-EL.345/SETP + D055-2L14 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-10 14:13 | 2024-10-10 14:44 | 0 | 0 |  | in G1/2 D055, the membrane is the smallest we have, it'll not work with any RP2. the closest solution we can offer as Anne reccomended, is a DT22L200230/9026 (Pmax 40bar accidental), it is providing you a separation from the process via a membrane and the connection is G1/2. we added a bigger passage diameter but it is not flush. I don't think any supplmier can offer a flush diaphragm G1/2 with tyhis type of pressure switch on top.
TCK46759 | User | Service Request | TR: BAIT - TRIVENETA IMPIANTI - RP2-EL.345/SETP + D055-2L14 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-10 12:08 | 2024-10-10 14:40 | 1 | 0 |  | Anne, comme tu le dis ça ne sera pas possible non plus aucun RP2 ne peut se monter sur un D055 en 1/2 ou 3/4. même en 1" ce sera uniquement les hautes pression et sans trop fortes surpression. Best, Nicolas
TCK46752 | User | Service Request | SNCF --- LEADS 30242875 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-10 11:33 | 2024-10-17 09:47 | 1 | 0 |  | Closed after meeting.
TCK46681 | User | Service Request | TR: BAIT - TRIVENETA IMPIANTI - RP2-EL.345/SETP + D055-2L14 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-09 14:20 | 2024-10-09 15:28 | 0 | 0 |  | Pas possible en D0552L14 ni a moins du 1"1/2. le mieux qu'on peut faire c'est un DT22L200230/9026 (Pmax 40bar accidentelle sur le montage).
TCK46658 | User | Service Request | VS: CRM: MEX5 - Overensstemmelseserklæring (EF) 1935/2004 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-09 11:19 | 2024-10-09 12:01 | 0 | 0 |  | the option 1935 with a MEX + DANC compatible add the 3A logo on the dial and please find enclosed our 3A certificate. be careful to have only FDA compliant fillings (BH1 in the gauge and LRS1 for the mounting with the diaphragm seal).
TCK46622 | User | Service Request | MCX5-D60.N10/9690_9691/ Steve | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-09 03:57 | 2024-10-09 11:50 | 1 | 0 |  | as it is a brass version, it's always complicated to define only with a picture but I think it's more a product like below. the connection is not something we supplied (or at least not something that was standard and could be found again). I also offered a 1/4G thread as if it was an NPT, we should see a couple threads out. Order code: MTA3-D20.N10/9695_9696 MTA old Code / SAP code:
TCK46570 | User | Service Request | TR: Request for 42 pieces | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-08 11:46 | 2024-10-08 14:45 | 1 | 0 |  | MEP5460B19/0191 le triplex est standard sur le MEP5 et le full blow out back aussi donc les options 0751 et 0760 ne sont pas necessaires.
TCK46555 | User | Service Request | Old switch codes | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-08 09:55 | 2024-10-08 10:25 | 0 | 0 |  | I think I had this request a couple times already, it might have been sent to multiple distributors. here is the answer I sent Jorg: not the easiest question but here are the translations (oldest to old to newest): ZPN205SH = RPPNCA3205 = RPPN-7C3.205 ZPN209SH = RPPNCA3209 = RPPN-7C3.209 ZPN203GHHINT = RPPNV(GSHH)Y3203 = RPPY-7W3.203 ZPN204GSHHINT = RPPNVY3204 = RPPY-7W3.204 ZPN201GSH = RPPNDA3201 = RPPN-7D3.201 ZDPN224SIINT = RDPNAY3224 = RDY8A3224 ZDPN234SHH = RDPNWA3234 = RDN-8W3.234 this is assuming: the codes were correct INT means INTrisically safe connection is G1/2 (was the standard at the time).
TCK46552 | User | Service Request | CPX : codes de production | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-08 09:29 | 2024-10-08 09:43 | 0 | 0 |  | tu auras un suivi plus adapté si tu rentre dans le processus de claim (sans retour des appareils) il suffira d'une photo avec les numéros et le NF. je pourrais te trouver les infos mais je ne peut pas assurer le suivi des actions correctives et l'analyse des root causes. maintenant que le service qualité est au complet c'est mieux de respecter le process.
TCK46480 | User | Service Request | Another CAD file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-07 11:53 | 2024-10-08 09:45 | 1 | 0 |  | I don't have it in 1/8NPT, but here it is in 1/4.
TCK46472 | User | Service Request | SDI439 data sheet | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-07 11:35 | 2024-10-09 16:51 | 1 | 0 |  | Jorg will call you to discuss this.
TCK46471 | User | Service Request | Accuracy | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-07 11:24 | 2024-10-08 09:51 | 1 | 0 |  | both 0.5% and 0.6% exist and are available on a MEX5 of this range. but we strongly advise for the 0.6% as it'll include all the feature to respect an accuracy and reading of the best quality by still being conform to EN (fine scale, pointer with edgewise tip and specific calibration). the 0.5% option on the other hand is not recognised by the EN and so the EN marking will be removed from the gauge as not respected.
TCK46467 | User | Service Request | AKKODIS-TVE-07-10-2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-07 11:03 | 2024-10-08 10:11 | 1 | 0 |  | j'ai epluché les catalogues de température Bourdon, Sedeme et Haenni, (j'ai meme regardé nos anciens partenaires en temperature (Jarhbur et Rueger), je ne trouve aucune reference en 1515. désolé, ca doit venir d'autre chose. si tu arrive a avoir d'autres details sur le produits on pourra surement retrouver
TCK46370 | User | Service Request | TOTAL-20541302-GRCH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-04 10:17 | 2024-10-04 15:31 | 1 | 0 |  | Voici la bride d'un D438 3" 150 pour total, ca devrait matcher, le plan est joint pour verif avec le client. 11092547 BR 54025-12AF0A11+2 VANNES 0D10
TCK46368 | User | Service Request | RE: CPX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-04 10:00 | 2024-10-07 12:00 | 1 | 0 |  | MTTF TED5 sent
TCK46330 | User | Service Request | Old switches? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-03 16:09 | 2024-10-08 15:16 | 1 | 0 |  | the TXA was the old gas filled thermometer, and you're right a TXA4 can probably be replaced by a TSF080 now. but all the other codes don't match our configuration code and we don't have access anymore to the MACPAC system used in 2002. I attached the datasheet of the TXA hoping it'll help but there is probably some definition to do with the customer.
TCK46327 | User | Service Request | tsf -- total normandie- gie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-03 14:59 | 2024-10-04 15:12 | 1 | 0 |  | 10261758 pas de point bloquant apparent demander ouverture BHDE
TCK46257 | User | Service Request | Life cycle of MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-02 12:37 | 2024-10-07 10:51 | 1 | 0 |  | It's not a really simple question as there is not a standard value to define that and it's extremely dependant on the use case, but I can give you some values that are related to this. according to the requirement of the EN837-1, all our gauges are tested in stable pressure (long max of the scale applied) overpressure (1.25x the max of the scale according to EN, we do 1.5x) and cyclic pressure (100 000 cycles of 30 to 60% of the scale). on top of thosen standard test and all the other required in the EN, we also tested our sensing element of MPG6 (the same tube weldings and movement than the MEX5 during 300 000 cycles (20 to 80% of then scale). all cycling test are validated at the end by an accuracy check.
TCK46254 | User | Service Request | TED X pouvoir de coupure 127V | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-02 12:11 | 2024-11-05 17:48 | 1 | 1 |  | je ne trouve aucune trace de ce TED spécial je suis désolé, ils n'ont pas plus d'informations ? (une ancienne commande, une désignation complète ou encore mieux un code SAP ?)
TCK46227 | User | Service Request | CPX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-02 09:55 | 2024-10-07 10:33 | 1 | 0 |  | MTTF sent
TCK46199 | User | Service Request | MPG6 and Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-01 17:32 | 2024-10-08 10:19 | 1 | 0 |  | looks like replacement parts for this to me: 11168457 MPG6-DX0.122.XBR:FCBM12A8T600/0816_9711+ [Image] [Image] it might be spare parts for Total/Saudi Aramco, you could check with Jerome Bredon if he's not making an offer directly in middle east ?
TCK46153 | User | Service Request | TR: Kom. Wenzel 30242183 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-01 12:14 | 2024-10-07 10:30 | 1 | 0 |  | Nathalie, ça fait beaucoup de choses qui ne se font plus tu as raison, on a plus cette échelle ni l'approbation marine. on a pas de vrai remplacement, on peut juste leurs dire que comme transmetteur 250mbar ils peuvent regarder le PP20S industriel mais qu'ils n'auront pas l'option marine. Best, Nicolas
TCK46148 | User | Service Request | TBI non-standard ranges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-01 11:53 | 2024-10-07 10:25 | 1 | 0 |  | the lower temperature measurment limit is -70C, you cannot go lower. -120/+100C = span 220C too low, -70C limit -40/+160C = span 200C -30...170C existing already (code 55T) otherwise not impossible but you'll need to pay for the dial creation. -50/+200C = span 250C Possible, you need to pay for the creation of the specific dial (-40...200C should not be done as it'll not match the angle of a 250C sensing element).
TCK46104 | User | Service Request | WG: New ticket has been created - Bug + demande info site Baumer franceBug + demande info site Baumer france | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Widmaier, Andreas [wia] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-10-01 07:35 | 2024-10-07 10:17 | 1 | 0 |  | CC for reference only, no action requested.
TCK46069 | User | Service Request | Patina on the bellow of ZT305CHM0202 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-30 16:43 | 2024-10-10 10:19 | 1 | 1 |  | according to our product specialist, it is verdigris and do not cause any functionnal issue. it xould still be good to check your storage conditions (humidity especially) so that you can limit it's appearance.
TCK46046 | User | Service Request | AW: New ticket has been created - Bug + demande info site Baumer franceBug + demande info site Baumer france | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Clement, Alexis [clal] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-30 13:53 | 2024-10-07 10:16 | 1 | 0 |  | CC for reference, no action needed.
TCK45945 | User | Service Request | Consommation PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-27 15:26 | 2024-10-07 10:14 | 1 | 1 |  | Bonjour Jacques, pour répondre il faut que je connaisse le type (4...20 mA ?) et l'alimentation. Best, Nicolas
TCK45902 | User | Service Request | MCX5, Dial Scale 0 to 2000 Pascal / Alex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-27 11:03 | 2024-10-08 14:46 | 1 | 1 |  | it's not existing yet, it'll need to be created and it'll be a 2500 Pa as standard.
TCK45818 | User | Service Request | #11162064 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-26 11:58 | 2024-10-04 11:58 | 1 | 1 |  | no problem, the code is open you can go ahead.
TCK45801 | User | Service Request | FW: TED 5 Switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Shutt, Oliver [shol] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-26 10:00 | 2024-10-08 14:38 | 1 | 1 |  | You're right, the PFSX was a safety product based on the TED but it was not ATEX, it was SIL2 which is a different kind of safety certification. the YTED would have been the closest call but we didn't renew our ATEX certification as the TED is supposed to be phased out in a couple years. the closest alternative would be an ATEX PFMN for example but I doubt they can accept something that is not SIL and we don't have any other SIL products.
TCK45736 | User | Service Request | siap-grch-qualification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-25 12:59 | 2024-10-04 11:54 | 1 | 1 |  | oui pas de soucis, il faudra pas mettre le setp sous 200mbar.
TCK45687 | User | Service Request | MMX1 CAD / STP file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-24 15:47 | 2024-10-08 14:46 | 1 | 1 |  | I don't have the right connection but here is an MMX1 type F (in M10x100 instead of G1/4).
TCK45682 | User | Service Request | SOLYRO --VDAV--24/09/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-24 14:50 | 2024-10-09 17:24 | 1 | 1 |  | tu peux proposer le code ci-dessous, la norme EN n'autorise que l'utilisation des manomètres de sécurité sur les applications acétylène (sinon il faudrait faire comme pour loxygène et enlever le marquage EN sur un MEX et créer une option etc... pour l'instant ce n'est pas prévu). MEP5430B27/9936/EF04
TCK45593 | User | Service Request | MEP3 / BTIN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-23 17:10 | 2024-10-08 11:33 | 1 | 1 |  | according to the material certificate (an example of 2024 delivery with material certificate page 3 of tube and socket used for the manufacturing of MEP3), and according to what I checked in the NACE, for me it is conform (316L and hardness below 250HB) I asked for a confirmation by Nicolas fernandez and David lecomte.
TCK45558 | User | Service Request | replacement ¦ 11113635 ¦ _PFSX-2.3.B77.R.G1.18.03.2.2.0.1 0 1/++ ¦ Zapytanie ofertowe Baumer - UP/2080/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-23 12:52 | 2024-10-04 11:43 | 1 | 1 |  | the specificity of this product was the SIL2 certification and we stopped it, we don't have other SIL2 product to replace it. otherwise, the product was technically similar to the TED, you can give them the datasheet but I doubt it'll be acceptable for them.
TCK45534 | User | Service Request | TR: Confirm the price adder for option 0655 in model MPJ6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-23 10:44 | 2024-10-04 11:39 | 1 | 1 |  | Salut Anne, oui ça se fait mais ce sera l'ancien design. tu peux demander le prix. Best, Nicolas
TCK45478 | User | Service Request | Tool/key for gauges with contacts | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-20 16:57 | 2024-10-04 16:14 | 1 | 1 |  | Andy, the SAP code of the key is 10310128, check with your commercial contact for price and min qty maybe. Best, Nicolas
TCK45447 | User | Service Request | TR: PBMN-24B76RA14474202000 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-20 11:32 | 2024-09-20 12:01 | 0 | 0 |  | le PP20S n'est pas encore validé en montage séparateur donc un PBMN sera bon, le seals navigator valide le montage avec les options habituelles et une precision x2 donc 0.5%
TCK45442 | User | Service Request | ORANO MANO DIFFERENTIEL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-20 10:59 | 2024-09-20 11:54 | 0 | 0 |  | il faudra prendre un X05K (arc de 180) et demander une échelle spéciale -25...0 mbar (-250...0mmCE)
TCK45432 | User | Service Request | ZPX205SH, ZPN204GSHHINT, ZPPN234GSHH & ZDPN234SHH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-20 09:02 | 2024-09-20 10:37 | 0 | 0 |  | Old product datasheet sent for help with conversion.
TCK45370 | User | Service Request | TR: Y91 for oxygen application | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-19 10:55 | 2024-10-08 11:23 | 1 | 1 |  | notre spécification technique pour le nettoyage O2 est un document interne confidentiel, mais je peux te confirmer que nous respectons bien la norme ASTM G93-3, ci-dessous un extrait de la première page de notre spec qui donne les référentiels respectés: [Image]
TCK45254 | User | Service Request | the option code 0202 for ZPNE and ZT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-17 18:19 | 2024-09-18 09:51 | 1 | 0 |  | ZPN-E204CHM-202 _ZPN-E204CHM/SETP 11015193 (code closed, to be recreated with setp value included). ZT--305CHM-0202 ZT--305CHM/2249/SETP 10302321 (code closed, to be recreated with setp value included).
TCK45239 | User | Service Request | TR: Confirm the IC price | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-17 15:47 | 2024-09-18 10:03 | 1 | 0 |  | Answered on parent TCK45153
TCK45165 | User | Service Request | stp.files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-17 08:16 | 2024-10-08 10:25 | 1 | 1 |  | I'm sorry but this one is not yet existing, but you can order it via sales as a technical drawing, please check the price with your sales contact.
TCK45153 | User | Service Request | Confirm the model code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-17 07:10 | 2024-09-17 11:51 | 1 | 0 |  | Either the replacement with the same product and add the ATEX option for the CE marking and certificate DP 100/213.13B.535/0078 (11181066). for this solution, please check if you really want the red Index 97xx, if yes we need to check and make you an offer with it. -70 is not excluded on the datasheet and as there is no filling fluid it should work. A mounting of MEX5 + DANC DN38 + LRS2 is possible but limited to -60C
TCK45032 | User | Service Request | pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-13 10:35 | 2024-09-13 10:56 | 0 | 0 |  | Solved directly
TCK45009 | User | Service Request | D4212N20229G20 / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-13 05:49 | 2024-09-13 09:11 | 0 | 0 |  | D421 is an old name, it's D425 now: https://www.bourdon-instruments.com/fr/en/product-overview/diaphragm-seals/seals-with-flange-connection/d425/p/37214
TCK44976 | User | Service Request | BHFR Logo devenu inutile L311 | Service Desk > General | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-12 14:56 | 2024-09-12 15:14 | 0 | 0 |  | thanks a lot for warning us, I don't know if we can close the Logo and make a note for the replacement but it's very professionnal of you to warn us.
TCK44952 | User | Service Request | WG: 2258576 Bestellung 844321 / Quinta | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-12 11:45 | 2024-09-12 15:12 | 1 | 0 |  | Fantinelli that manufactures the MFT informed us this summer about a change in their product range. the tables and datasheets are under revision to integrate this change with Jorg but it's note yet finalised. I hope the alternatives offered will be ok.
TCK44918 | User | Service Request | MDJ-MABEO-CHEMIN-12.09.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-12 09:00 | 2024-09-12 10:29 | 0 | 0 |  | 11045861 MIX7-D30.X14/9709_D186
TCK44893 | User | Service Request | TR: C4B: THP Instrument (Differential Pressure Switch - Baumer) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-11 17:11 | 2024-09-11 17:39 | 0 | 0 |  | c'est 2 colliers en Inox comme ci-dessous qui viennent se fixer avec des écrou sur la plaquette de montage arrière. Plan joint
TCK44820 | User | Service Request | AORP0384 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-11 11:16 | 2024-09-13 12:03 | 1 | 0 |  | Direct request for a special product to Schneider or Parker or Astava
TCK44807 | User | Service Request | ORANO --- 11271527 ---- Probleme fct. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-11 09:44 | 2024-09-11 09:58 | 0 | 0 |  | Ok pour ce que tu dis, le montage RP2 + D820 DN50 est OK et les -30mbar ne posent pas de problème. par contre il serait bien de nous envoyer l'appareil défectueux après remplacement,t pour confirmer ce qui n'allait pas.
TCK44749 | User | Service Request | Confirm if product is available from BHFR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-10 12:16 | 2024-09-10 16:55 | 1 | 0 |  | I don't think this is from Bourdon Germany and I don't think we have anything to compete with this, sorry.
TCK44743 | User | Service Request | Differential Gauges and Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-10 11:21 | 2024-09-10 16:51 | 1 | 0 |  | this will be an MFT5 + diaphragm seals (Monel maybe with PTFE coating) and sales need to request the offer for the full mounting directly to our MFT5 supplier.
TCK44733 | User | Service Request | TR: Y913 GEO 50Bar (11079781) / Cameron | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-10 10:23 | 2024-09-10 16:48 | 1 | 0 |  | je lui avais déjà confirmé que le 750bar ne sera plus dispo mais a priori il n'avait jamais été commandé, au besoin il faudrait faire un nouveau développement produit mais pour l'instant on le décline juste.
TCK44724 | User | Service Request | MAN Energy Solutions France--- VDAV--10/09/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-10 09:34 | 2024-09-11 16:33 | 0 | 0 |  | c'est un encore plus ancien que les MAR, et on aura plus la collerette carrée avec ces entrax. le mieux est de proposer la nouvelle version du MEX avec la colerette et la peinture special marine comme un MEX5E30B20/0751_0151_0412_0670_0840_9695_9696_9699_A009 Apres bien sur je ne vois pas la connexion ni le type de verre.
TCK44659 | User | Service Request | MX35 Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-09 14:29 | 2024-09-09 17:49 | 1 | 0 |  | Basically yes but it also have a lot of specific parts. there is no datasheet for it and it should only be used when you need to go a bit lower on a diaphragm seal mounting that require 100mm gauge. the drawing and external dimensions will all be the ones of the MEX5.
TCK44619 | User | Service Request | FW: Baumer - Barflex4 / GCD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-09 10:45 | 2024-09-10 13:53 | 1 | 0 |  | [Image]
TCK44596 | User | Service Request | ALKIAN --- BOGUET --- MANOMETRE 10263591 -- URGENT !!! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-09 09:22 | 2024-09-10 10:14 | 1 | 0 |  | Yes, only external seal. Advised for tracking with official claim.
TCK44573 | User | Service Request | MX/MZ overrange limitation | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-09 08:04 | 2024-09-10 17:19 | 1 | 0 |  | in this picture you can see the 2 "golden" pins that'll block the movement in case of over or under pressure. but as it's only for safety and the belows are resistant, it can stop quite far from the max diff pressure on the dial. [Image]
TCK44506 | User | Service Request | Enquiry for Y913 GEO 50 bar, 400 bar, 750 bar / Cameron | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-06 10:48 | 2024-09-06 11:21 | 0 | 0 |  | Customer specific drawings sent.
TCK44484 | User | Service Request | cira-roucout-mdj | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-06 09:15 | 2024-09-06 09:42 | 0 | 0 |  | pour de la pression relative sur des manomètres, absolument. par contre sur un différentiel, non car ce serait faux, on ne peut pas marquer relatif quand ça ne l'est pas.
TCK44473 | User | Service Request | FW: Pressure gauge DRO80/311.111 / Mandy | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-06 08:32 | 2024-09-06 09:38 | 0 | 0 |  | as I can't see the connection or measure it, it can be one of those 3 codes: 10320119 1/2G 10343940
TCK44464 | User | Service Request | Material check for specification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 17:57 | 2024-09-06 09:31 | 0 | 0 |  | 2 capillaries and 2 flanges in the code in SAP
TCK44463 | User | Service Request | TR: Question about FLEX.S112 G1/2 FEM. 1M | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 17:22 | 2024-09-06 09:25 | 0 | 0 |  | [Image]
TCK44462 | User | Service Request | TR: Anfrage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 17:21 | 2024-09-06 09:20 | 0 | 0 |  | Nathalie, c'est un mex avec un tube en queu de cochon, tu peux propposer un MEX5D30B29. et tu peux dire a Giorgio qu'il fasse un effort c'est la reference la plus standard de notre portfolio. Best, Nicolas
TCK44444 | User | Service Request | new code? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 15:20 | 2024-09-05 15:52 | 0 | 0 |  | Sure, here you go: RT-NAA.402.120E0 11011368
TCK44439 | User | Service Request | MESURE-BOILLOT-VDAV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 14:38 | 2024-09-05 15:53 | 0 | 0 |  | Here is the code from the order reference: 10287896 MFE422323JB2-01 / 6346 23 010-02
TCK44435 | User | Service Request | TBHI / TBI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 13:54 | 2024-09-11 16:36 | 0 | 0 |  | (-50...550C based on a 0...600C) is possible, and we validate it from R&D and production but you need to ask for the creation of the code and there might be a price for the creation of the dial.
TCK44421 | User | Service Request | M31 With Range 0 - 400 mbar / Vidara | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 12:51 | 2024-09-05 14:07 | 0 | 0 |  | Dial not existing yet, construction possible.
TCK44411 | User | Service Request | VERIFICATION PRESSOSTATS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 11:25 | 2024-09-06 10:11 | 1 | 0 |  | I'm leaning towards these two references, and anyway, the scales need to be approved by the client because I can only find 11 and 43, not 10 and 40. RDE-4E3.131 RDN-4E3.111
TCK44387 | User | Service Request | RP2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 09:43 | 2024-09-05 12:09 | 1 | 0 |  | parts in contact are in 316L so no problem
TCK44386 | User | Service Request | Enquiry for RP2N / Japan Marine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-05 09:43 | 2024-09-05 12:07 | 1 | 0 |  | 10297002 _RP2-NL.6XX/2178/2264/9704/2500KPAG replacement suggested: RP2-NL.655/SETP_2264
TCK44315 | User | Service Request | availability ¦ DRO 100 411.133 105 ¦ 0-40 Bar ¦ Customer request: Products | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-04 10:32 | 2024-09-05 12:00 | 1 | 0 |  | DRO100411133105 replaced by MEX5D31B27
TCK44310 | User | Service Request | Résistance Y91 au Skydrol Hyjet V | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-04 10:08 | 2024-09-05 14:26 | 1 | 0 |  | Hydraulic fluid based on phosphate esters, so according to the recommendations I've found on several seal manufacturer websites, EPDM seems to be the best option. 316L and ceramic are fine.
TCK44227 | User | Service Request | PLANS BRIDES YOKO / 20592654 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-03 11:28 | 2024-09-03 16:35 | 1 | 0 |  | Drawings to be ordered, or datasheet information available.
TCK44207 | User | Service Request | PLANS BRD912 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-03 09:32 | 2024-09-03 11:25 | 0 | 0 |  | Client changed his mind
TCK44202 | User | Service Request | Urgent drawing request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-03 09:01 | 2024-09-03 09:10 | 0 | 0 |  | Drawing sent raw after on the MJX7 a template was enough.
TCK44148 | User | Service Request | reading accuracy of ZT/ZP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-02 14:52 | 2024-09-02 17:05 | 1 | 0 |  | the scale is just an indication, always calibrate or check compared to a measuring instrument.
TCK44128 | User | Service Request | RDD PBA3131 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-02 13:53 | 2024-09-02 17:23 | 1 | 0 |  | l'équivalent actuel sera un: Order code: RDN-4A3.131 RDN old Code / SAP code:
TCK44074 | User | Service Request | PROJET BOCCARD-- PJ24000141-- VDAV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-02 10:12 | 2024-09-06 10:21 | 1 | 0 |  | MEP5-431.B79/0753_0763/DANC01-2L2A.01 et la tous tes accessoires sont OK: 11109175, 10341806, 11112854
TCK44058 | User | Service Request | GE renewable energy--- BURAK--- AROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-02 09:34 | 2024-09-03 12:03 | 1 | 0 |  | RGB (0, 179, 0) approx RAL 6038
TCK44035 | User | Service Request | Enquiry on RT2E / Eriks | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-09-02 08:12 | 2024-09-02 16:57 | 1 | 0 |  | old code 10298028 RT2-EP.45.C0XC6/SETP/X=270MM to be replaced by: RT2-EP.45.C04C6/SETP 45C rising (250mm)
TCK43937 | User | Service Request | TR: Item information : Pressure Indicator(Safran P/N: 004207-00000) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-30 16:32 | 2024-09-02 11:30 | 1 | 0 |  | for replacement, see with the plain supplier or component manufacturer: [Image]
TCK43932 | User | Service Request | Pression d'épreuve hydrolique pour Air liquide | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-30 15:24 | 2024-08-30 15:45 | 0 | 0 |  | Vu par teams, en prod tous les transmetteurs sont testés en surpression admissible.
TCK43883 | User | Service Request | BAPL ¦ Replacement ¦ ED701 K20 .413.146.11 ¦ OSM BIERUŃ zakład w Głubczycach | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-30 09:29 | 2024-08-30 09:47 | 0 | 0 |  | It was a G1" flush connexion, 250mbar DIN43650 so I would suggest this: [Image]
TCK43868 | User | Service Request | MPF6- D60.X24/9709 / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-30 08:20 | 2024-08-30 14:10 | 0 | 0 |  | MPG6 -> MPG6 new version in PBT (UL 94 V0 fire resistance) MPE6 -> phase out MPJ6 -> stay the same and is also UL 94 V0 MPF6 -> phase out.
TCK43858 | User | Service Request | TR: Product availability of "TXA5-B" | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 17:15 | 2024-08-29 17:24 | 0 | 0 |  | TXA5 replaced by TSF100.
TCK43852 | User | Service Request | RE: Need IC price & decoding | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 16:11 | 2024-09-02 09:32 | 1 | 0 |  | 9704 Phased out and 0833 replaced by 1999 New codes: RDN-8C6.221/2231_1999_SETP RPPN-7C6.209/2231_1999_SETP
TCK43847 | User | Service Request | qualified class pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 15:37 | 2024-08-29 16:11 | 0 | 0 |  | CE certification sent + K3 RCCE
TCK43838 | User | Service Request | TVE-EDF-BEROUDIAUX-29.08.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 14:00 | 2024-08-29 14:59 | 0 | 0 |  | identification visuelle et pland des potentiels filetages envoyés pour aider a l'identification.
TCK43837 | User | Service Request | TESTS DE CONFINEMENT STATIQUES EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 13:58 | 2024-08-29 15:00 | 0 | 0 |  | normalement, la pression interne et externe ségalisera, en cas de pressurisation trop rapide, voyant polycarbonate et ventilation max du boitier conseillés.
TCK43792 | User | Service Request | BAIT / Customer: Kiwa Cermet Italia / Repair request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Grandini, Alessio [gale] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 09:35 | 2024-08-29 10:46 | 0 | 0 |  | Change window: 10308596 WINDOW GLASS 4MM THICK 100MM (glass) 10308593 SAFETY LAMINATED WINDOW NS100 (triplex 0751) 10308595 Window POLYCARBONATE (PC 0753) or order a new gauge.
TCK43791 | User | Service Request | BRIDES ADAPTATION YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-29 09:31 | 2024-09-02 12:19 | 1 | 0 |  | BRD912 4LGE B0M 111 BRD912 4LGG B0M 111 les adaptations et Assomptions sont: vu les PN/DN, norme EN1092 supposée faible PN donc laminé proposé a la place de forgé (plus standard) PN20 pas existant, proposé en PN25 finition standard circuit de tracage 2 vannes SW1/4 OD 10 car utilisé sur les dernieres affaires et conforme REQ
TCK43690 | User | Service Request | RE: RECLAMATION CLIENT ARINDEX | Service Desk > General | Low |  |  | Closed | Petit, Annie [anpe] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2024-08-28 08:55 | 2024-08-28 14:25 | 0 | 0 |  | ticket closed, cc for information only
TCK43656 | User | Service Request | Picture for material on stock | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-27 17:01 | 2024-08-28 10:16 | 1 | 0 |  | Pictures of samples, box and paperwork sent.
TCK43647 | User | Service Request | TR: RECLAMATION CLIENT ARINDEX | Service Desk > General | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Jud, Noel [jnoe] | 2024-08-27 16:08 | 2024-08-28 14:25 | 0 | 0 |  | ticket closed, cc for information only
TCK43610 | User | Service Request | Drawing request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-27 13:14 | 2024-08-27 14:15 | 0 | 0 |  | it can be ordered separately as technical drawings (PRODUCT DRAWING 90000279) you can check with sales for price. if you need the list of codes and tags etc... it might be code (SPECIFIC CUSTOMER DRAWING 90000023).
TCK43602 | User | Service Request | MSDS for BH1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-27 11:28 | 2024-08-27 11:48 | 0 | 0 |  | BH1 MSDS attached
TCK43576 | User | Service Request | BOCCARD --- VDAV --- PJ24000141 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-27 08:44 | 2024-08-28 11:24 | 1 | 0 |  | codes créés et envoyés
TCK43538 | User | Service Request | RE: BAIT - IMA / ECOR / Idroinox - MEX5D30B77/0751/0078+Y91+ARID+DANC | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-26 15:38 | 2024-08-26 16:43 | 0 | 0 |  | Not a request, in cc for information only
TCK43528 | User | Service Request | TR: Customer request: productInfo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-26 14:46 | 2024-08-26 16:09 | 0 | 0 |  | PBSN + DANC DN 38, so FDA compliant but not 3A as the PBSN is not mentionned on our 3A certificate
TCK43517 | User | Service Request | TR: Pressure transmitters - approvals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-26 14:19 | 2024-08-26 18:08 | 1 | 0 |  | TED non UL, certificates review planned with mario to clarify UL products
TCK43446 | User | Service Request | PROJET LEBAS --- grch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-26 08:39 | 2024-08-26 16:35 | 1 | 0 |  | ATEX 2IIGD, no possibility to change this and no other similar products available with better zoning.
TCK43399 | User | Service Request | MEX5 / Vodpak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-23 11:32 | 2024-08-23 14:46 | 1 | 0 |  | On Bar Bourdon scales, the min step of the scale will be printed on the dial as standard. example: MEX5-D30.A20
TCK43304 | User | Service Request | STORENGY - AROE - byo - 22.08.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-22 09:42 | 2024-08-23 15:26 | 1 | 0 |  | Application validée et expliquée par teams.
TCK43188 | User | Service Request | TR: valeurs MTTF pour CPX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 17:31 | 2024-08-20 17:40 | 0 | 0 |  | Document MTTF CPX envoyé
TCK43185 | User | Service Request | WEP5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 17:17 | 2024-08-20 17:33 | 0 | 0 |  | that's good questions and we're thinking about a test kit, but we don't have anything yet. the basic need for them to demonstrate is a gateway they can link to a computer (LoRa gateway compatible with the right regions but I think africa is working on the same wavelength as europe for LoRa). this is to receive the signal and transmit is as a payload. then they need a LoRaWAN network server to unpack the payload (we internally use chirpstack which is online and open source) from this they'll be able to receive a understandable payload and display simple graphs. so it's pretty cheap to set up but they'll need someone that knows a bit about connecting wireless devices or lora and is good with computer and have at least very basic skills in programming (for example to translate the code received into data, we can send them the code in python but they need someone that can at least basically understand it and who knows where to enter it). also obviously we'll be here as support if they need anything.
TCK43182 | User | Service Request | FW: POLA-0296-DIVE-2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Prince, Vinny [lpr] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 16:40 | 2024-08-21 11:57 | 0 | 0 |  | both articles can be ordered and the SAP codes are open.
TCK43129 | User | Service Request | RE: RECLAMATION CLIENT ARINDEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 11:53 | 2024-08-20 14:53 | 1 | 0 |  | Handled by sales and quality.
TCK43118 | User | Service Request | RPPN7-G3.205 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 11:10 | 2024-08-20 15:40 | 1 | 0 |  | Configuration validated by R&D: RPPN7 G3 205 : Micro 1 : 8.75 Bar Rising Micro 2 : 9.25 Bar rising Best, Nicolas
TCK43104 | User | Service Request | TR: PO No. 4241110089 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 10:52 | 2024-08-20 14:42 | 1 | 0 |  | Montage OK au Seals Navigator.
TCK43094 | User | Service Request | Calibration Certificate AORPB | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 10:32 | 2024-08-20 14:40 | 1 | 0 |  | Yes we can and it'll be with certificate 90000267.
TCK43070 | User | Service Request | TR: Copper and zinc banned for MEX3 and DANC, D055 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-20 08:46 | 2024-08-20 17:11 | 1 | 0 |  | nous pouvons donner les certificats matière des parties en contact mais le 316L contient du cuivre (entre 0 et 0.75% selon la norme) ci joint l'exemple pour les certificats d'un flasque et d'une membrane de DANC.
TCK43007 | User | Service Request | Claim CL165446.1 | Service Desk > General | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-19 13:13 | 2024-08-20 14:13 | 1 | 0 |  | Ticket closed, handled by quality dept. Best, Nicolas
TCK42954 | User | Service Request | BAES-wurth- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-19 09:15 | 2024-08-20 15:39 | 1 | 0 |  | le fabricant de la boite avec les embouts et le mano achète juste le mano chez nous.
TCK42952 | User | Service Request | 20585185-grch-ut03 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-19 09:08 | 2024-08-23 16:22 | 1 | 0 |  | Answered on call, the diameter of the entry for the cable is 14.8mm and they have to open the cover to wire it.
TCK42908 | User | Service Request | MEX5+DANC Expected Shelf-life / CSL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-16 13:53 | 2024-08-22 09:48 | 1 | 0 |  | Same for expected shelf life, we never did real tests, honestly I tested gauges from the basement that were very old and still working so I'm sure we could at least state a 10 to 20 years but nothing is sure and we don't know how to test shelf life as we would have the final results maybe only in 50years if thats were it fails. Best, Nicolas
TCK42757 | User | Service Request | Y91-3.3.3.B79.R.DANC01-2L2A.01 / Kah Fei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-14 12:27 | 2024-08-20 14:24 | 1 | 0 |  | this mounting is possible and I added the necessary option for the fitting and the accuracy downgrade: Y91-3.3.3.B79R.DANC01.2L2A.01/0592_9010
TCK42619 | User | Service Request | Certificate of Conformity / CSL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-13 09:13 | 2024-08-22 09:46 | 1 | 1 |  | About ADI, the only thing we could use that potentially could have traces of animal parts is the cleaning/polishing grease. Concerning BSE and TSE, we are not able to make tests obviously (we're not a virology lab) but the parts in contact with the fluid are all stainless. what I would suggest is to offer a laboratory cleaning (or another cleaning if they prefer) on the final parts and components. this is the best answer I can give you on this.
TCK42582 | User | Service Request | PRESSURE GAUGE FOR OXIGEN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-12 18:38 | 2024-08-20 11:58 | 1 | 1 |  | Hi Angel, we have nothing that can be directly used at this temperature. they need to lower the emperature on their side before connecting a measuring instrument. Best, Nicolas
TCK42466 | User | Service Request | code 0201, 0202, 0203 & 0205 in Pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-09 15:17 | 2024-08-20 11:56 | 1 | 1 |  | Hi Anil, please find all codes in the table below: Old code New code SAP Code Qty Remark RDPNCA6221-0202RPPN-CA6209-0202RPPE7E6209-0201RDPNEA6131-0201RPPNEA6204-0205RPPNEA6206-0202RPPNEA6204-0205RP2NL655-0201RP2EL654-0203RPPNEA6104-0201 Best, Nicolas
TCK42350 | User | Service Request | Y913 ¦ N CDE/NOrder: 1218515/10 ¦ odp ¦ Zamówienie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-08 11:08 | 2024-08-20 11:45 | 1 | 1 |  | Hi Anna, the SAP code ordered on the order 1218515/10 was 11105926 Best, Nicolas
TCK42344 | User | Service Request | Question D820 with manomètre | Service Desk > General | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-08 09:47 | 2024-08-20 11:43 | 1 | 1 |  | Hi Giorgio, it'll not work, the diaphragm is flush on the D820 and so DN25 being the smallest one, the diaphragm is too small for this range. you need to check the ones with internal diaphragm like D821 DN25 or D435 DN25 that would work. Best, Nicolas
TCK42141 | User | Service Request | BAIT - IMA / ECOR / Idroinox - MEX5D30B77/0751/0078+Y91+ARID+DANC | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-06 10:59 | 2024-08-20 11:40 | 1 | 1 |  | Request for price handled by sales dept, the technical request ticket is closed. Best, Nicolas
TCK41909 | User | Service Request | need article no. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-01 16:02 | 2024-08-01 17:06 | 0 | 0 |  | Hi Sascha, we advise to take the ASIPD long model, it's more standard, less expensive and more efficient. you can find the datasheet on the website and the prices in the P2F. Best, Nicolas
TCK41887 | User | Service Request | TR: Fiche technique 11204619 / FESTO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-01 11:43 | 2024-08-20 17:20 | 1 | 1 |  | on a pas de datasheet mais les caractéristiques dont l'IP sont les mêmes qu'un MIT3 sauf pour la précision (Classe 2.5) et bien sur pour tout ce qui est dimensionnel.
TCK41878 | User | Service Request | TB063 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-01 09:45 | 2024-08-01 11:22 | 0 | 0 |  | Hi Sascha, we never did a TB with a specific range, for this number of units I think it could be accepted but you need to go through an SPR. please fill one and send it to me or Jorg. Best, Nicolas
TCK41877 | User | Service Request | MEX5 With Dial Scale 0 to 10 mH2O / APS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-01 09:39 | 2024-08-01 11:27 | 0 | 0 |  | Hi Bee Tin, yes, it exists, and it'll be X15/9709_Dxxx, you can take the example below as a reference: 11173855 MEX5-D60.X15/9709_D000 Best, Nicolas
TCK41869 | User | Service Request | 转发: material check | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chen, Weili [chw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-08-01 09:00 | 2024-08-01 17:04 | 1 | 0 |  | Hello Weili, this is possible, I advise the option 0770 (0.4mm instead of 0.3mm because it's already on a visquous oilt and the diaphragm sseal mounting already have a pulsation dampening effect (otherwise you might have a serious impact on response time). MEX3-D23B22/0770 /DT22H600.230.01 Best, Nicolas
TCK41834 | User | Service Request | picture of option logo of company wit logo Bourdon | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-31 22:49 | 2024-08-22 14:13 | 1 | 1 |  | No answer from the customer on necessary details.
TCK41782 | User | Service Request | Diaphragm Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-31 10:15 | 2024-07-31 15:45 | 1 | 0 |  | Hi Daniel, I miss some details about the gauge, and we don't have duplex diaphragm but I suggest to offer something like this: MExxD6xxxx+ D438L100N60227G0L0A09909 a capillary can be added directly in the code of the D438, I offered with 2 valves. Best, Nicolas
TCK41750 | User | Service Request | Question D820 with manomètre | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-30 15:47 | 2024-07-30 16:56 | 0 | 0 |  | Hello Giorgio, for the MEX + D820 you requested it's OK, see below. concerning the PP20S it's not a product integrated yet in the Seals Navigator or authorized for the fittings with diaphragm seals for now, please change it for a PBMN and then it'll be fine. Best, Nicolas
TCK41716 | User | Service Request | replacement ¦ E910110035010 ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-30 12:19 | 2024-07-30 15:32 | 1 | 0 |  | Hi Anna, I don't have any record of a code like this but as it starts with E91, I think it can probably be replaced by the Y91 which is the Atex standardised version of this product. https://www.baumer.com/ch/en/product-overview/process-sensors/pressure-sensors/y91/p/36421 Best, Nicolas
TCK41706 | User | Service Request | BAIT / Customer: Kiwa Cermet Italia / Product Repair | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Grandini, Alessio [gale] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-30 10:38 | 2024-07-30 10:48 | 0 | 0 |  | Hello Alessio, those are old products and they were installed for some time already. with a rubber wrench and if you maintain the case on multiple points to avoid forcing on the connection/case welding you can open and close it but then you'll need a replacement glass anyway. you can check with your sales contact if you can order some replacement windows, here is the SAP code to request. 10308596 WINDOW GLASS 4MM THICK 100MM Best, Nicolas
TCK41637 | User | Service Request | CLEMESSY | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-29 13:52 | 2024-07-29 14:57 | 0 | 0 |  | Salut Véronique, oui tu peux, pas besoin de raccords tournants, les 2 séparateurs sont libres. [Image] Best, Nicolas
TCK41523 | User | Service Request | RT2-EL.45.C03C3/SETP / Martin | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-26 11:11 | 2024-07-26 11:26 | 0 | 0 |  | Hi Bee Tin, yes, no problem. Best, Nicolas
TCK41480 | User | Service Request | RP2-YM.345/SETP / Kah Fei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-26 03:56 | 2024-07-26 09:07 | 0 | 0 |  | Hi Bee Tin, it's on the limit but yes. Best, Nicolas
TCK41465 | User | Service Request | TR: SBB & HP Fittings | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-25 14:42 | 2024-07-25 16:56 | 1 | 0 |  | Hi Anil, we have nothing like that in our range of products, and the manifolds are BTIN ones as far as I can see. if I accept, it means creating special codes, finding suppliers and exotic materials for no chances of getting it as we will source it with low qty. I think BTIN who has full machining capabilities and already does accessories would be a much better choice to get the market. Best, Nicolas
TCK41447 | User | Service Request | MA2 100B 1/2G 0.400B | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-25 11:42 | 2024-07-26 16:47 | 1 | 0 |  | Bonjour Veronique, j'ai peur qu'il faille redefinir un produit, dans nos tables de code OTAN, je retrouve des MA2, mais liés a differents produits (MEX, MMA, MMD) mais aucun code ne matche et tous sont bloqués. il faut repartir du besoin pour definir un produit, on ne pourra pas remplacer. CODIFICATION MACPAC concaténation toutes colonnes utiles pour recherche de plans par exemple ! CODIFICATION SAP et/ou info suppression MMA4 A30B20)MMA4 A30B74)MMA4 A30B24) Best, Nicolas
TCK41432 | User | Service Request | Différence entre 8991 pour RP2 et 8990 pour RPPN7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-25 10:14 | 2024-07-25 16:45 | 1 | 0 |  | Salut Yannick, non, la différence est si nous fermons le plombage ou non, par exemple si le client doit recevoir l'appareil, faire le réglage du setp puis plomber et mettre sur ligne, alors seulement un plombage "préparé" est la solution. si tu veux l'utiliser directement et que tous les reglages sont fait en production ici, alors on peut l'envoyer directement plombé. Best, Nicolas
TCK41430 | User | Service Request | Remplacement gamme --- telemecanique XMLA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-25 10:04 | 2024-07-26 09:25 | 1 | 0 |  | Salut Yannick, pour un nombre de pièces comme ça je ferais une SPR de toute façon mais il faudra la faire pour le raccord G1/4F. Order code: RP2-NL.H46/SETP RP2 old Code / SAP code: pour le 1 seuil écart fixe, je verrais avec un UT de BTIN comme tu le disait pour équilibrer les prix (sinon faire les 2 avec l'écart réglable pour avoir des qtés intéressantes si ça va au client). Best, Nicolas
TCK41388 | User | Service Request | Projet H2 -- ARAC Projet Technocampus Hydrogène --- ADF --HSAN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-24 16:28 | 2024-07-26 10:59 | 1 | 0 |  | Yannick, pour le pressostat, voici ce que je propose, finalement: vu la soudure laser, que la membrane est en 316L plate (sans déformation plastique) de 50 micron et vu les "faibles"pressions et température, un simple RP2Y ou E devrait faire l'affaire. Order code: RP2-YM.355/0829_SETP RP2 old Code / SAP code: Best, Nicolas
TCK41348 | User | Service Request | Minimum Scale per division for MCX5 & MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-24 11:28 | 2024-07-26 11:58 | 1 | 0 |  | Hi Bee Tin, please see hereafter the number of divisions (based on the bar model) for MEX if you take the option for the accuracy 0.6% (if you need this type of dicision, it makes sense only for high accuracy). Pressure range (kPa) Resolution requested (kPa) Resolution with option 0,6% (kPa) Product Comments 10 Best, Nicolas
TCK41325 | User | Service Request | M31 ORANO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-24 11:05 | 2024-07-24 11:23 | 0 | 0 |  | Bonjour Véronique, c'est normal que ce ne soit pas au P2F car 0.25bar avec une statique de 16bar donne un cadran sur 100 et une précision de plus de +/-3% nous ne conseillons pas ce montage, il vaut mieux demander le prix pour un MFT avec contacts: Order code: MFT7-A00.N10N.32/4021 MFT old Code / SAP code: Best, Nicolas
TCK41304 | User | Service Request | demande tech -- ARKEMA- HSAN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-24 09:16 | 2024-07-24 10:34 | 0 | 0 |  | Vu en Teams, DPC 100/813.153/0..16 BAR/9700 Best, Nicolas
TCK41281 | User | Service Request | FW: RFQ TECH/2024/0059 :PRESSURE SWITCH / Kah Fei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-24 03:35 | 2024-07-24 09:08 | 0 | 0 |  | Hi Bee Tin, Absolutely no problem, you just need to ask for the code creation including the setp. Best, Nicolas
TCK41254 | User | Service Request | BHFR | Service Desk > General | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-23 14:49 | 2024-07-24 09:02 | 0 | 0 |  | Hello Giorgio, please find enclosed the drawing you asked for. Best, Nicolas
TCK41207 | User | Service Request | TR: Question about Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-23 09:54 | 2024-07-23 10:01 | 0 | 0 |  | Nathalie, je ne vais pas avoir plus d'infos que toi, ce n'est pas une reference mais un bout de spec avec en plus des incoherences. je commencerais par proposer un MEX8-DS3.B41 (en gardant l'option ouverte pour la double graduation). a confirmer que: 3/8PT veut dire 3/8NPT entre 150 et 160mm Best, Nicolas
TCK41161 | User | Service Request | BAIT - IMA / ECOR / Idroinox - DRCG101.211.133.535/3272DANC01-2L2B_LRS1 ATEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-22 15:57 | 2024-07-22 16:41 | 0 | 0 |  | Hi Francesco, we don't have atex certification on the DRCG101, please refer to my answer on ticket 41159 to know what to offer for the same functionnality with ATEX certification. Best, Nicolas
TCK41159 | User | Service Request | BAIT - IMA / ECOR / Idroinox - DRCG101.211.133.535/3272DANC01-2L2B_LRS1 ATEX | Applications & Software > SAP | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Brunner, Urs [brun] | 2024-07-22 15:56 | 2024-07-22 16:42 | 0 | 0 |  | Answered on ticket 41116. Best, Nicolas
TCK41116 | User | Service Request | BAIT - IMA - DRCG101.211.133.535/1091_3272_LRS1 ATEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-22 14:15 | 2024-07-22 16:03 | 0 | 0 |  | Hi Francesco, sorry this is the only gauge with electrical 4...20mA output. but a solution could be found with a MEX5 + Y91 + ARID + DANC, maybe with a slightly higher range and it would have the same capability. MEX5D30B79 + Y91333B79R + ARIDDF5F53 + DANC01.2L2A.01 Under review with R&D for feasability, and accuracy of the transmitter *2 considering the mounting. Best, Nicolas
TCK41022 | User | Service Request | Y-SEN-13796/2,5 ¦ czujnik ciśnienia Y-SEN-13796 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-19 14:24 | 2024-07-22 15:53 | 1 | 0 |  | Hi Anna, this is not a Baumer / Bourdon sensor, but we have railway dedicated sensors, We could offer them a PP10R but if it's not spec'ed in since the start of the project, I don't know if they can replace with it and also I don't know the pressure connection they're using. otherwise, the range (2.5bar), the electrical connection (DIN43650) and the railway approval EN50155 are available with the PP10R. Best, Nicolas
TCK41005 | User | Service Request | listing MFT UTO-19-07-2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-19 11:34 | 2024-07-25 09:24 | 1 | 0 |  | Salut Yannick, ça ne va pas se faire en une journée, mais j'ai prévenu les acteur et je leurs envois la liste des MFT UTO. On clôture le ticket et le sujet sera pris par la qualité et le nucléaire (en parallèle du sujet global MFT qui est chez le PM). Best, Nicolas
TCK40989 | User | Service Request | SOLYRO-TBI-RECLAMATION-19.07.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-19 09:56 | 2024-07-19 10:25 | 0 | 0 |  | Vu en Teams, gestion via claim. Best, Nicolas
TCK40955 | User | Service Request | FW: RFQ : MIX/CEI + RSI amplifier / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 15:47 | 2024-07-18 16:06 | 0 | 0 |  | Hi Bee Tin, I'll stop you there, I'm here for technical advices, not to do the job of the sales guys and treat the requests from your customer. please define a product according to the customer request, and if you have a technical question I'll be happy to help. Best, Nicolas
TCK40936 | User | Service Request | Ticket update - Urgent! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 13:02 | 2024-07-22 14:00 | 1 | 0 |  | Handled through both ticket number (1 closed and one ongoing for the 3D drawing for an answer begining of this week). Best, Nicolas
TCK40917 | User | Service Request | Serial no. provided for pressure gauge. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lim, Janice [ljan] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 11:03 | 2024-07-18 16:58 | 1 | 0 |  | Hello Janice, most of the time I'm able to trace it back to a reference but not this time. the tracability system on the differentials is still the most recent, if it happens on gauges I'll be able to trace it easily. with the pictures I can still give you some information to start defining the replacement: it seems like an MX7 P60 XXX as I don't see the front I cannot have more details on the filling or range. Best, Nicolas
TCK40913 | User | Service Request | YOKOGAWA QUESTION TECHNIQUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 10:24 | 2024-07-18 16:15 | 1 | 0 |  | Salut Véronique, c'est une norme asiatique, elle est en générale considérée comme équivalente à l'ANSI en terme de designs de brides. on peu lui proposer de la classe 150 ou 300 en ANSI et lui donner les datasheets pour confirmer le dimensionnel par rapport à sa demande. Best, Nicolas
TCK40910 | User | Service Request | TR: pulsation dampener | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 10:03 | 2024-07-18 15:56 | 1 | 0 |  | Normalement non, il faudrait des ARIN, mais un code existe, demande au responsable produit si tu peux toujours le prendre et si le prix est bon: 10265033 ARA-10.L20H Best, Nicolas
TCK40896 | User | Service Request | YTED ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 09:19 | 2024-07-18 15:02 | 1 | 0 |  | Hi Anna, here is the details of this order, we stopped the ATEX version but the same one without atex can be configured on the website and ordered, and the option 0602 is no longer an option but an accessory: 11223048 TED-6.3.3.B24.R/0605 11044646 PROFSI3-B25091-ALG-YT ATEX [Image]
TCK40881 | User | Service Request | MEP5-463.B15.AORPB504/SETP.ARIN2306.ASIP-D1L3/0078_0678_0693_9941 / Neste | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-18 06:23 | 2024-07-18 14:55 | 1 | 0 |  | Hi Bee Tin, it's a bit of an inbetween, if we want to be very careful, both needs it considerin the design temperature. but first the AORP will reduce a bit the temperature and second the operating temperatures would not require it. we cannot take the decision on this, I would offer both options to the customer. Best, Nicolas
TCK40829 | User | Service Request | [BHFR Material Data] FW: Aanvraag PBSN | Process Instrumentation > Material Data Management | Low |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-07-17 12:05 | 2024-07-24 11:42 | 1 | 0 |  | Hi Kees, it was not easy to read but here is the code of the product: 11102339 PBSN-1.2.B76.R.A1.14.03.2.1 Best, Nicolas
TCK40807 | User | Service Request | FW: ATB 16134960 and 16130549 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-17 10:25 | 2024-07-18 14:46 | 1 | 0 |  | Hi Jurgen, please check this datasheet, we can't go as low in pressure but a 600mbar should be possible :). Best, Nicolas
TCK40772 | User | Service Request | Enquiry For MCX5-460.N07.D920-D66.2219.G20.08/0078_0751 , 1 unit / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-17 04:50 | 2024-07-18 14:36 | 1 | 0 |  | Hi Bee Tin, do you know who offered this ? I don't see a reason for the D920 (for transmitters) the connection is wrong, the LRS8 is for vacuum for transmitters, N09 is our lowest possible range for fitting with a gauge etc... I would change it for this, please check if it fits your request: MCX5460N09 - D820B6N2219G2009 Best, Nicolas
TCK40603 | User | Service Request | 3D / Step files - Urgent | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-15 10:02 | 2024-07-23 16:42 | 1 | 1 |  | Hi Paul, please find enclosed trhe drawings. Best, Nicolas
TCK40544 | User | Service Request | Request for material number | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-12 14:44 | 2024-07-12 16:32 | 0 | 0 |  | Hi Jorg, I send the request to Sylvain and you can check with him when you're here next week . Best, Nicolas
TCK40540 | User | Service Request | RE: Nuclear | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-12 14:02 | 2024-07-12 16:30 | 1 | 0 |  | Hi Paul, we have nuclear cleanliness and some gauges were tested for seismic tests and approved for non classified area. we also have K3 classified pressure and temperature switch. all those products can benefit from the nuclear cleaning. But as far as I know we are not class 1E. Best, Nicolas
TCK40537 | User | Service Request | RE: MPG6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-12 13:42 | 2024-07-12 16:28 | 1 | 0 |  | Hi Paul, you can go ahead, it exists in H88, but you'll have to include the fee for the dial creation in D88 as it doesn't exist yet (to be defined with sales ). Best, Nicolas
TCK40536 | User | Service Request | RE: Old switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-12 13:34 | 2024-07-18 16:42 | 1 | 0 |  | Hi Paul, I should have answered way quicker on this as I had the request last week for the same one from singapore, I just recalled. here is the answer: The closest I can offer as a first approach would be this, please check with the customer to validate the details (405 was changed to the closest range available 415, the length of capillary and stem where decided as standard as possible, all of this should be reviewed with the customer). Order code: RT-NAA.415.230E0/SETP RT old Code / SAP code: Best, Nicolas
TCK40334 | User | Service Request | Re: RP2-NL.343 setpoint question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-11 18:08 | 2024-07-12 09:27 | 0 | 0 |  | Hi Paul, You're perfectly right, 0.25 falling setpoint. and it's the standard but the hysteresis should be at the minimum. Best, Nicolas
TCK40329 | User | Service Request | TR: PFMN sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-11 17:12 | 2024-07-12 09:22 | 1 | 0 |  | Bonjour Cyrille, pas de problème de compatibilité ou de montage tu peux y aller. Best, Nicolas
TCK40328 | User | Service Request | UT03--GRCH---urgent | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-11 16:32 | 2024-07-12 09:41 | 1 | 0 |  | Salut Yannick, les UT peuvent en général se régler entre 20 et 80% de léchelle, donc le setp ne sera pas possible sur un UT. ce sera possible sur des RP2. Best, Nicolas
TCK40325 | User | Service Request | TR: N Cde/Nb order: 102723805 ¦ AKP_INQ_20240710_1415 | Service Desk > General | Low |  |  | Closed | Pétré, Sandra [sape] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-11 15:49 | 2024-07-12 11:33 | 1 | 0 |  | Résolu via teams avec n d'OF
TCK40293 | User | Service Request | TR: Material Safety Data Sheet (MSDS) for the oil in MZ7-P63.B12K | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-11 10:42 | 2024-07-11 14:05 | 0 | 0 |  | Anne, ci-joint la msds. Best, Nicolas
TCK40235 | User | Service Request | TR: BAIT - PP10R-1.3B20R.A144.092600.050/C01 - 11268663 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-10 15:15 | 2024-07-10 17:32 | 1 | 0 |  | Anne, pour ca je pense que le mieux c'est de demander a Mario qui est PM si il peut ouvrir le code demandé. ca sera le plus simple. Best, Nicolas
TCK40171 | User | Service Request | RP2-EL.341 With IEC Ex Approval ? / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-10 05:49 | 2024-07-10 09:57 | 1 | 0 |  | Hi Bee Tin, please find enclosed the IECEx certificate. Best, Nicolas
TCK40094 | User | Service Request | Old switch code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-09 12:28 | 2024-07-10 17:42 | 1 | 0 |  | Hi Paul, This is the limit to find something in terms of dates, this one is from 1998 and I already only have a paper version to find the reference. CDE: 535483 REF: ZPX204GS INT N: 2 UNITS So the replacement will be: RPPY-7C3.204/SETP but the switch GS is not ATEX anymore so I had to replace it (to be confirmed) and the connection was not in the code but it was G1/2 as standard so I offered the same. Best, Nicolas
TCK40068 | User | Service Request | TR: Twoja prośba baumer.com | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pétré, Sandra [sape] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-09 11:20 | 2024-07-09 12:18 | 0 | 0 |  | Bonjour Sandra, prix idem MCD7-430.N05 11272108 Best, Nicolas
TCK40055 | User | Service Request | TR: EF6 problem of Yongji | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-09 10:19 | 2024-07-09 12:16 | 0 | 0 |  | Nathalie, j'ai transféré la demande a Mario, je ne peux pas créer de fiche de modification produit. je te dis des que je reçois une réponse. Best, Nicolas
TCK40031 | User | Service Request | TR: CTX with option 0765 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-09 09:00 | 2024-07-09 09:05 | 0 | 0 |  | Nathalie, 315bar maximum et joint viton obligatoire au dessus de 10bar Best, Nicolas
TCK40027 | User | Service Request | Need model code for temperature gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-09 08:55 | 2024-07-09 09:46 | 0 | 0 |  | Hi Anil, the request is handled also in europe, the customer needs to replace 40. this is a product we sourced with stiko so if you think the request will be handled through India, you can contact BHDE in charge of the contacts with Stiko. Best, Nicolas
TCK39990 | User | Service Request | D820 re-work | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-08 16:53 | 2024-07-08 17:12 | 0 | 0 |  | Hi Paul, it is technically possible as it's gaz connection and we can unfit it but as the diaphragm seal need to be reorder, the over cost of ordering a brand new MEX5 will be lower than sending back, unfitting, cleaning and changing the diaphragm seal and fit and fill again. I would advise to order new ones. Best, Nicolas
TCK39969 | User | Service Request | FW: Order receipt Confirmation - POR177866 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Goldsworthy, Sally [saha] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-08 14:55 | 2024-07-08 15:31 | 0 | 0 |  | Hi Sally, we noticed a couple small leaks when the gauges experiences some temperature changes like in the transport. I would suggest to clean externally and check if it still leaks at standard conditions. if it does, please start a claim process with sales administration. Best, Nicolas
TCK39949 | User | Service Request | MANOMETRE H2 100 % en 20 mBar --- BAES | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-08 13:34 | 2024-07-08 15:26 | 0 | 0 |  | Salut Yannick, pas de soucis a ce type de pression normalement les fuites ne devraient pas poser de problèmes. je proposerais un MCX5xx0N05, je te laisse voir ce que veut le client en terme de montage et connexion. pour les options, l'ATEX et un verre triplex sont des bonnes bases pour toutes les applications hydrogène, lévent ouvert pour la ventilation si ils bossent en environnement couvert. Best, Nicolas
TCK39921 | User | Service Request | Square Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-08 10:49 | 2024-07-08 11:23 | 0 | 0 |  | Hi Daniel, yes, it was a product we did in the past but the standardisation of the round model made us stop it long ago, the one you have probably was one of the last made. we don't have replacement for it. In india they might have it but I don't know if you can order through India. Best, Nicolas
TCK39891 | User | Service Request | TR: EF6 problem of Yongji | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-08 08:25 | 2024-07-08 16:32 | 1 | 0 |  | Hi Nathalie, the first thing is that as it was internal sales, if such limitations are stated in a contract, our responsible persons for the product at BHFR should be informed of it. in this case I can provide a couple more information about this change in order to justify the change and it's impact for the customer. this change was motivated by the poor delivery performance of our previous supplier and we changed for a standard supplier for the group on other product. this means the supplier for the new round collar is already qualified and proved it's reliability and quality. considering the impact for the customer, this piece doesn't changeany of the key characteristics, the connections are still the same, it doesn't have impact on the sensing element, electronic etc... I hope this answer will help you understand and accept this change for the future. if any other information is needed, do not hesitate to ask, we'll be happy to provide answers. Best, Nicolas
TCK39866 | User | Service Request | TR: AR 2203330 --- GIRODIN ---- REGLAGE SETP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-05 16:09 | 2024-07-08 11:36 | 1 | 0 |  | Bonjour Priscilla, ce ne sera pas possible car si le setp est a 16bar a la hausse, le réglage minimum de lhystérésis étant à 2.5bar, ça donne un resetpt a 13.5bar ce qui est en dessous du seuil mini des 10% de léchelle. le plus bas SETP sera de 17.5bar à la hausse avec hystérésis mini de 2.5bar et don un reset à 15bar. sinon le fait de reprendre le réglage est possible. Best, Nicolas
TCK39850 | User | Service Request | dde BXDE /Christian CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-05 14:19 | 2024-07-08 11:38 | 1 | 0 |  | Nathalie, les CTX ne sont pas disponibles avec l'option propreté O2, il faudra passer sur un Y9. Best, Nicolas
TCK39849 | User | Service Request | Barflex EA5-AF0.B72R | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-05 14:19 | 2024-07-24 09:06 | 1 | 1 |  | Bonjour Cyrille, désolé que ça ait pris autant de temps, on ne va pas en dessous de 10bar avec le 0.1%, il faudra prendre un EA4A et prévoir un joint. Exemple: EA4-AF3.B72R 10270450 Best, Nicolas
TCK39765 | User | Service Request | Old product question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 17:44 | 2024-07-05 11:48 | 1 | 0 |  | Hi Paul, we can still offer it but from standard, the flange is available as ABR like below (datasheet enclosed) they can choose the capilaries as AKPL and the NPT connections can be welded (9030). 11010329 ABR2233R0N Best, Nicolas
TCK39723 | User | Service Request | Material Certificates | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 12:29 | 2024-07-04 14:35 | 0 | 0 |  | Hi Daniel, the question is tricky because yes, on most of our processes we're able to track the material certificates but not on everything. so the clear rule is it really needs to be ordered with the product, but if you have the problem, now there is more chances we can help you out. you can ask me directly with the order number and the serial number of the product. Also what I'll be able to give you is an extract of production tracking and the material certificate from the supplier, if you need an official one with Baumer certificate letterhead you have to run the request through sales and quality. if you ask me for example for the certificate for serial n22902400029881 I could give you this and the 2 certificates enclosed: Qty Designation Manufacturing order Serial n1 Serial n2 SAP Code Date of production Tube material n Connection material n 2 Best, Nicolas
TCK39715 | User | Service Request | Pressure Switch / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 11:21 | 2024-07-04 12:15 | 0 | 0 |  | Hi Bee Tin, Switch typer TR seems better suited for cutting 2A at 24Vdc (but limited to 10A at 250V, they ask for 15A but we don't have any option to go higher than 10) otherwise it seems right according to their requirement, you can offer diaphragm sealsq with wetted parts in monel as an option for the seawater, it could really improve the lifespan of the product if it's really raw sea water. For tag PS-102-001 to 004: RP2-NR.654/SETP, for these tags for raw seawater, need to add DT seal ? For tag PS-601-003 to 005: RP2-NR.651/SETP, for these tags for brine, need to add DT seal ? For tag PS-501-201 to 006: RP2-NR.656/SETP Best, Nicolas
TCK39695 | User | Service Request | EQUIVALENCE ROBIVAL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 10:06 | 2024-07-04 14:52 | 1 | 0 |  | Bonjour Véronique, On ne fait plus le différentiel à cellule Barton. daprès le tableau ce sont tous des différentiels HP donc plutôt MFT. La meilleure approche est de définir les MFT (pour moi toutes les infos sont dispo). et ensuite se baser sur la datasheet des contacts ou celle d'un mano à contact pour définir les codes des contacts (4021 pour la deuxième ligne par exemple, pas inverseur mais respecte les points de consignes.) un fois tout défini, tu envois aux ventes qui feront la demande de devis à Fantinelli. Best, Nicolas
TCK39679 | User | Service Request | TR: item inquiry : attached the picture. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 08:47 | 2024-07-04 09:48 | 0 | 0 |  | Nathalie, je pense que c'est celui ci mais il faudrait enlever le voyant, la lunette et le premier cadran et m'envoyer une photo du dos dui premier cadran et de la face du second pour pouvoir confirmer. ou trouver une commande du client pour confirmer avec cette reference. 10323208 SDI 439-AXFLOW 1...3M / 0045 Best, Nicolas
TCK39676 | User | Service Request | RP2-NL.346/SETP. DT1.2L90.0.9N0.01 For Sea Water / Mala | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 08:45 | 2024-07-04 09:42 | 0 | 0 |  | Hi Bee Tin, I would offer a DT2 instead of a DT1 (less expensive). otherwise your mounting is Ok and so is your setpoint (because of the hysteresis it'll set back off slightly below 2bar but this is normal). RP2-NL.346/SETP. DT2.2L90.0.9N0.01 (Setp 2bar rising hysteresis minimum). your customer can set it if he has a calibration station, but it's not super easy, it's a precision job. Best, Nicolas
TCK39675 | User | Service Request | TR: CTX quote | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-04 08:45 | 2024-07-04 09:34 | 0 | 0 |  | Bonjour Nathalie, c'est du specifique tres proche du standard, on ne produira plus le code demandé, mais il peut surement etre remplacé par du CTX. le standard le plus proche sera : [Image] a valider avec le client. Best, Nicolas
TCK39655 | User | Service Request | GALI --- NOBLECOURT ---- Option 0151 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 17:33 | 2024-07-08 17:09 | 1 | 0 |  | Salut Yannick, Ci-dessous les références du primer et des composants du mélange de peinture noir. en PJ les datasheets du primer gris et de la peinture noir. Grey primer: 10303615 PRIMAIRE RAL 7035 MAT URA-NYL 016 Black paint and finish: 10303539 PEINT.NOIR RAL9005 BR 75%ALPHA PUR 10303590 DURCISSEUR U-PUR E 135 10303661 DILUANT URKISOL 900/P Best, Nicolas
TCK39629 | User | Service Request | SNCF --- BELKACEM --- AROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 15:27 | 2024-07-04 15:18 | 1 | 0 |  | Lol bonne question, On ne va pas beaucoup l'aider car on utilise des procédés de propreté et donc pas de lubrifiants sur nos mouvements. normalement, si lintérieur est bien nettoyé il n'en a pas besoin. et si le mouvement grippe, il serait plus sage de le remplacer, la graisse ne fera surement que masquer le problème et le défaut reviendra surement à un moment plus problématique. Best, Nicolas
TCK39592 | User | Service Request | 转发: 回复：堡盟压力开关停产咨询 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 12:13 | 2024-07-03 17:22 | 1 | 0 |  | Hi Zhang Wei, I think it is more a problem in the codes, the H in the first code is more related to the type of sensing element (RPPH) and the A is the code for the standard switch. We changed the way the switches were codified a couple time but we don't change the meaning of a digit, a switch type H is always the manual reopening, closing on fall. Best, Nicolas
TCK39540 | User | Service Request | FW: Request for Thermometers | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 08:15 | 2024-07-03 11:31 | 1 | 0 |  | Hi Jurgen, this type of thermometer is not in the official catalogue but it was supplied from Stiko at the time and they still have this pointed stem and handle bars option. you need to request it to BHDE (Jorg Strufing) they're handling Stiko. Best, Nicolas
TCK39534 | User | Service Request | RP2E Product Life Cycle / Petronas | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 06:37 | 2024-07-03 09:23 | 0 | 0 |  | Hi Bee Tin, this product is not obsolete and there is no short or mid-term plan to phase it out. Best, Nicolas
TCK39530 | User | Service Request | TED6 / CSE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-03 04:47 | 2024-07-03 09:17 | 0 | 0 |  | Hi Bee Tin, we have'nt performed specific tests for humid atmosphere resistance. the closest I can give you is the Ingress Protection rating which is IP67 which means it is weather proof and can resist a short immersion in water. I would say it's a good indicator of wether or not it'll resist humid environment. Best, Nicolas
TCK39504 | User | Service Request | TR: Urgent RFQ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-02 14:32 | 2024-07-03 18:35 | 0 | 0 |  | Nathalie, Suite confirmation directe par email comme c'etait urgent, le AMFD02N origine inde 10264916 est bien Nace. Best, Nicolas
TCK39354 | User | Service Request | PFMN sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 14:43 | 2024-07-01 16:48 | 0 | 0 |  | Salut cyrille, il va falloir faire quelques petits changements mais dans l'ensemble ce serait possible. PFMN-54.BC3R.A1551.33200.4501_ZPX1-007_D920-10L.42GH.B30.09_9010 1: préciser le réglage qui ne pourra pas être en dessous de 1bar pleine échelle ZPX.xxx: le ZPX n'est pas encore intégré clairement dans nos montages il vaut mieux proposer un ASIPD pour l'instant (surement le petit). D920: le D820 sera plus approprié, les D9 sont pour des applications faible volume (typiquement les transmetteurs process Yoko, Honeywell, siemens etc...) 09: pas de problème pour le LRS9 mais ce n'est pas une obligation comme avec les transmetteurs process. 9010: précisions globales du PFMN x2 L'uranus B6 ne devrait pas poser de problème mais le prix sera a confirmer (on en fait jamais). confirme avec ton client selon mes commentaires et il faudra me le soumettre une fois les détails clarifiés. Best, Nicolas
TCK39350 | User | Service Request | Replacement For E923 3F B18R + 0409/ Christine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 14:18 | 2024-07-01 17:10 | 1 | 0 |  | Hi Bee Tin, Sure, the only difference is we stopped the non ATEX version to standardize the ATEX, but it's an upgrade only for your customer. the code exist without the 0409 (which is just an external accessory). 11262645 Y92-3.3F.B18.R and you can request the creation of the code Y92-3.3F.B18.R/0409 if I were you, I would offer anyway the 2 options as it might be expensive with the accessory and the customer could prefer to order it without and find the adaptation by himself. Best, Nicolas
TCK39329 | User | Service Request | 5-point Calibration For TED6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 12:25 | 2024-07-04 11:58 | 1 | 0 |  | Hi Bee Tin, please find enclosed an example of a standard 5 points calibration certificate for TED (made today): Best, Nicolas
TCK39328 | User | Service Request | BAIT - Ecor International S.p.A. / IMA - DRCG 101/211.133.535.DANC01-2L2B.01/3272 - 11269322 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 12:17 | 2024-07-01 16:27 | 1 | 0 |  | Hi Francesco, first of all, on the first picture I don't see anything wrong, the pointer is in the tolerance (larger black mark on the 0 graduation). also during the transport via air shipment, the gauges can experience slight changes in atmospheric pressure and so a drift of the pointer. to correct this, you simply apply the standard procedure at the reception of the gauge, you remove the top plug to equilibrate the pressure, tap on the side of the gauge slightly to ensure the pointer is stable. if by any chance a gauge is still wrong after this, you can start the claim process. Best, Nicolas
TCK39326 | User | Service Request | SEPARATEUR TRES HAUTE PRESSION | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 12:01 | 2024-07-01 16:22 | 1 | 0 |  | Salut Véronique, tu as raison bien sur, on a rien a proposer actuellement la dessus. si jamais total veut vraiment demander un développement la dessus, il faut nous faire une demande complète avec les quantités le marché etc... via la direction industrielle comme pour le WEP5. Best, Nicolas
TCK39314 | User | Service Request | fichier 3D pour le TED-6.3.#.###.# | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 11:08 | 2024-07-01 11:10 | 0 | 0 |  | Salut Cyrille, bien sur, voici le 3D Best, Nicolas
TCK39310 | User | Service Request | TR: IC price & code (DOW)4516278735 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-07-01 10:40 | 2024-07-01 10:59 | 0 | 0 |  | Anne, dans l'ancienne codif des DA, le dernier digit était pour les accessoires de montages (collier et joint pour le DANC par exemple). ce n'est plus le cas car nous ne fournissons plus les accessoires. tu peux donc proposer un: MEX5-A30.B26/0751_0078 + AKPL2F5301 + DANC02-2L2A/1340 (11190331) (les accessoires de montage du clamp ne seront pas fournis, le joint dans ce cas ci) Best, Nicolas
TCK39249 | User | Service Request | NIH---ASIP2---PFMN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 16:20 | 2024-06-28 16:22 | 0 | 0 |  | Oui, que le montage se fasse sur le raccord accessoire ou process, le montage et les efforts sur la membrane seront les memes. Best, Nicolas
TCK39244 | User | Service Request | RPPN8 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 15:06 | 2024-06-28 16:09 | 0 | 0 |  | Véronique, on le fait sur les RPPN7 qui ont les mêmes boîtiers donc pas de soucis pour moi, prendre exemple sur le code 11175509 RPPN-7C3.204/2298_SETP. Best, Nicolas
TCK39232 | User | Service Request | FW: Product confirmation temperature transmitter ZTN & Pressure Switch / Gilang | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 13:50 | 2024-07-04 10:17 | 1 | 0 |  | Hi Bee Tin, this ZTN is very old and doesn't even have the date marking so I was not able to find exactly what it was. but the closest I can offer as a first approach would be this, please check with the customer to validate the details (405 was changed to the closest range available 415, the length of capillary and stem where decided as standard as possible, all of this should be reviewed with the customer). Order code: RT-NAA.415.230E0/SETP RT old Code / SAP code: Best, Nicolas
TCK39219 | User | Service Request | SPIE BATIGNOLLE-- nih - asip | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 13:46 | 2024-06-28 16:21 | 1 | 0 |  | et voici: [Image] Best, Nicolas
TCK39202 | User | Service Request | ARTICLE -- 11259776 - lyondell - danzin -- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 11:03 | 2024-06-28 15:42 | 1 | 0 |  | Salut Yannick, elle n'existe pas en 600, il faudra la créer, ce sera le montage des parties en vert ci-dessous et vérifier que le commerce est d'accord pour ne remplacer que le flasque inf. SAP code Designation Qty Comment Best, Nicolas
TCK39178 | User | Service Request | Need explosion proof certificate for model RPPE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-28 08:30 | 2024-06-28 09:34 | 0 | 0 |  | Hi Anil, please find enclosed ATEX and IECEX for RxxE. seems like they mixed it up with the intrisically safe one on the website. I here a lot about hydrogen in India since the beginning of the year. can you tell me who's the contact person for H2 applications at BTIN ? Also please keep me updated about the H2 market and make sure the customer have access to our page about hydrogen and our flyer https://www.bourdon-instruments.com/fr/en/bourdon-instruments-and-hydrogen-energy-transition/a/hydrogen-energy Best, Nicolas
TCK39148 | User | Service Request | Pressure verification system for oven | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-27 16:35 | 2024-06-28 16:17 | 1 | 0 |  | Hio Jurgen, sorry but even for Total, we don't have solution over 400C, maybe they can reduce it before the connection with a T or something similar. Best, Nicolas
TCK39136 | User | Service Request | FW: BAIT - Offer to make | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-27 15:11 | 2024-06-27 17:46 | 1 | 0 |  | Hi Paul, the size and types of contacts are things we don't have and would necessitate real development. Also we don't have the marine approval and would probably have to be tested and approved by vikings. We can't reply yes on this call but I would still send it to jorg to know if this is the kind of customer/projects we could approach in the future (if they have issues with their current supplier). Best, Nicolas
TCK39082 | User | Service Request | TR: Price required for manifold MFD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-27 10:53 | 2024-06-27 11:58 | 0 | 0 |  | Salut Anne, je dirais exactement le code qu'il a détaillé: 11199988 A.MFD.02.N/6170_1803 Best, Nicolas
TCK39067 | User | Service Request | Y913 Serial number | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-27 09:44 | 2024-06-27 12:07 | 1 | 0 |  | Hi Daniel, yes, the G1/2 doesn't add up, I would offer the code as defined (they can't invent the option number) and ask them to check about what they mean with the G1/2: [Image] Best, Nicolas
TCK39049 | User | Service Request | PBSN Serial Number? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-26 17:23 | 2024-06-26 17:57 | 0 | 0 |  | Hi Daniel, it's a manufacturing order number but yet, here is the result: 11051912 PBSN-1.2.B27.R.A1.44.03.2.1 check it matches just in case. Best, Nicolas
TCK39024 | User | Service Request | HENDSOLDT - H2-NIH - 26.06.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-26 14:24 | 2024-06-26 17:04 | 1 | 0 |  | Désolé Yannick, pour l'instant on a pas de solution en termes dinstruments, on pourrait proposer un montage avec accessoire mais je suppose qu'ils ont déjà vu cette possibilité. malgré tout comme pour tout ce qui concerne lhydrogène, je vais informer de la demande et si tu a plus de détails je suis quand même preneur, comme ça si ça revient on aura des billes pour pousser du développement. Best, Nicolas
TCK39000 | User | Service Request | Thermometer with scabbard up to 30/40 °C ¦ zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-26 11:32 | 2024-06-26 13:52 | 1 | 0 |  | Hi Anna, sure, I would suggest as below, but other options would be possible if they have more specifications like material, IP, accuracy or else: TBH080-212.151.52T + T 09143.0005 Best, Nicolas
TCK38948 | User | Service Request | BAIT - SPINA GROUP I: 74659 - RDO Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-25 16:29 | 2024-06-25 16:57 | 0 | 0 |  | Hi Francesco, I'm the technical support, not internal sales, I see the codes are already defined and you can probably find all prices in the P2F. Best, Nicolas
TCK38851 | User | Service Request | Faulty gauge? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-24 15:56 | 2024-06-24 17:28 | 0 | 0 |  | Hi Daniel, this seems a bit weird but we are on extremely sensitive gauge, and a 6mbar shift is somehow very small, what should be done first is to open the top plug on the case to balance atmospheric pressure, with the planne it can sometime shift a bit. the you put the gauge vertically and slightly tap on the case to ensure the pointer is going to it's stable position. if after this there is still an issue, please tell me and we'll see what to do. Best, Nicolas
TCK38849 | User | Service Request | ED701 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-24 15:43 | 2024-06-25 11:08 | 1 | 0 |  | Sorry, I think I found out the connection "R" and it is as follow but we cant replace it or itll be extremely high price for small qty: [Image] Please offer the standard clamp DN38 on a Y923 or a PBMH Best, Nicolas
TCK38790 | User | Service Request | DR80 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-24 11:30 | 2024-06-24 14:52 | 1 | 0 |  | Hi Jurgen, yes you can go ahead and order DR80 like this one: 11170120 for the price and delivery, you need to ask your sales contact, I don't have those information. Best, Nicolas
TCK38708 | User | Service Request | Oil Refinery Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-21 15:56 | 2024-06-21 17:02 | 0 | 0 |  | Hi Daniel, considering what we have validated with Total, 316L stainless is often a good starting point in oil and gas, then, here are a couple tips to offer special materials. if it's an offshore project, you can offer a Monel option on everything that contains the word crude. foul, sour or Acid here might contains sulfur residus (H2SO4 of H2S) that are toxic and usually needs special protection (hastelloy and potentially gold coating during component separation with H2 presence) I hope this helps. Crude Oil Demulsifier Hydro Carbon Liquid Hydro Carbon Vapour Hydro Carbon Gas Foul Water Crude Oil + Caustic Sour Water Fuel Gas Pilot Gas Reduced Crude Oil Corrosion inhibitor N2 Gas Slop Oil Off Gas Vacuum Residue No. 1 SR ? No. 1 3R ? No. 2 SR ? No. 3 SR ? Pre Flash Vapour HGO LGO Off Gas Condensate Dehydrated LGO Dryer Slop Oil Dryer Foul Water Acid Gas Best, Nicolas
TCK38684 | User | Service Request | FW: request price for MEX5 (X77) from Flowmaster | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-21 11:59 | 2024-06-24 14:34 | 1 | 0 |  | Hi Bee Tin, here is our standard for what you ask: MEX5-D30.K77 [Image] It would also be possible to get mmHg in the negative second scale but then it's an existing special (X77/9712). Best, Nicolas
TCK38679 | User | Service Request | MA7-AIRLIQUIDE-NAVANTIA- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-21 11:19 | 2024-06-21 11:25 | 0 | 0 |  | Salut Yannick, la précision est de 2%, mais l'option 0192 n'est pas disponible sur le MA7 et tu as déjà la plus haute résistance dispo a la surpression (25bar). MA7-435.B18K/0751_0078_0764_0792_9941 Best, Nicolas
TCK38678 | User | Service Request | Urgent drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-21 11:16 | 2024-06-21 11:20 | 0 | 0 |  | Hi Paul, here is the 3D and pdf 3D of the MEX3 bayonett. I reported the drawing was wrong to Jorg last week. Best, Nicolas
TCK38675 | User | Service Request | REGLAGE PRESSOSTAT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-21 10:49 | 2024-06-21 11:17 | 0 | 0 |  | Salut Yannick, je ne vois pas très bien le pointeau sur le contact donc c'est compliqué de dire. mais les vis ne semblent pas avoir été touchées et tout semble en place, le seul truc qui membête c'est la trace sur le contact au niveau du pointeau, comme si il avait été poncé. une photo plus détaillée de cet endroit serait bien. Best, Nicolas
TCK38626 | User | Service Request | Urgent !! | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-20 17:37 | 2024-06-20 18:00 | 0 | 0 |  | Hi Paul, I made a few corrections: MEX5-A60.B31/0751 + AORP-386/SETP + ARIN2606 + AKPL-206303 + D044-2L03.01 SETP: 130bar preferably even 100bar (otherwise between 130 and 165bar, the gauge will be damaged.) max pressure of the complete mounting 400bar. for where to change from NPT to G, it could also be the ARIN or the D044. Best, Nicolas
TCK38625 | User | Service Request | Option 0771 restrictor screw | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-20 17:18 | 2024-06-20 17:53 | 0 | 0 |  | Hi Paul, 0771 is a throttle of a certain size so it fits all type of 1/2" connections of instrument including MEX5, MIX7, MCX7 and MX7 for example. an example: 11248124 MX7-A30.D18J/0771 it would not be this screw on a MEX3 because the max connection is 1/4" (it would be the 0769 with 0.3mm diameter). here is the code of the restrictor screw by itself and you need to check with sales if you can order it and at which price: 10305741 VIS FREIN 0,5 M5 INOX Best, Nicolas
TCK38546 | User | Service Request | KEM-ONE-PROJET AS301-VDAV-20.06.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-20 09:09 | 2024-06-20 10:16 | 0 | 0 |  | Bonjour Yannick, j'ai du lancer les demandes de prix, ce que tu avais trouvé était de chez ASTAVA et ils ne font plus de manifolds monoflage. on va donc partir sur du schneider mais ils ne font que de l'ANSI et je vois que sur la 1'', la norme n'est pas très claire. le plus proche que l'on peut proposer c'est tel que ci-dessous avec le lien pour la page schneider pour les details de dimensions et de design: DB4 NGA NGA SD DB4 NNA NNA SD https://www.as-schneider.com/gb/en/Products-Services/Instrumentation-Valves-and-Manifolds/Monoflanges-VariAS-Blocks/VariAS-Blocks-Flange-x-Flange.html Best, Nicolas
TCK38542 | User | Service Request | Anfrage 24-17530 // Angebot 20581524 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-20 08:44 | 2024-06-20 09:34 | 0 | 0 |  | Hi Anja, it is possible, the placement of the markings might be slightly different but the ranges will be as on the picture. you need the range S75: -30"HG+30 psi/bar it'll be in black and black scale, you can request a scale in red as in the picture but it might become a special dial. Best, Nicolas.
TCK38524 | User | Service Request | doubt gasket DT2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-19 23:24 | 2024-06-20 09:23 | 0 | 0 |  | Hi Angel, it's not really a mistake, we cannot put a constraint on an option to be available depending on a digit in the P2F. this option is only available with the DT1. I would try to use the DT1 with a mounting as low as possible, don't hesitate to send me a request with the customer's request details so that we can check the best option to offer. Best Nicolas
TCK38501 | User | Service Request | AXIMA --- 10276661 --- GRCH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-19 15:15 | 2024-06-19 15:45 | 0 | 0 |  | Salut Yannick, le BH7 n'est plus utilisé, il est remplacé par le BH3 qui couvre une plage de temperature plus grande, est conforme aux exigences avec le NH3 notamment et dont la viscosité est du meme ordre. Best, Nicolas
TCK38491 | User | Service Request | AW: Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-19 14:40 | 2024-06-19 15:04 | 0 | 0 |  | Hi Giorgio, sorry but I don't think it is, It's a bit difficult to be sure with the picture but I'm pretty sure it's this option: [Image] 0745 is the same option but in triplex, we have very few codes created with this options and all of them are also with the ATEX option which is not the case on your picture (no ATEX marking). Best, Nicolas
TCK38452 | User | Service Request | BAES-NAVANTIA CARTAGENA-SDI439 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-19 09:07 | 2024-07-12 12:02 | 1 | 1 |  | Closed because of lack of internal response. please reactivate via email if still relevant. Best, Nicolas
TCK38435 | User | Service Request | TBI130-264.140.XXT/0727_0751_0110_9003_0305, Temperature Range -20 to 120 deg.C / Azzam | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-19 05:20 | 2024-06-19 11:21 | 1 | 0 |  | Hi Bee Tin, except for exceptionnal circumpstances, we will stick to the standard ranges on thermometers, so I can offer you [Image] [Image] Best, Nicolas
TCK38405 | User | Service Request | RE: Pressure Switches | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 15:26 | 2024-06-18 15:46 | 0 | 0 |  | Error in direction.
TCK38403 | User | Service Request | Pressure Switches | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 15:17 | 2024-06-18 15:21 | 0 | 0 |  | Hi Daniel, I'm sorry, you're right it'll be a type G, but it's available only on the RPP series. we'll also need the details of each setp to confirm if it can be set like this. Best, Nicolas
TCK38371 | User | Service Request | TR: Création BAFR // Client : CIRA // Cde Clt : 242217714 // MR59D33B16 + DT2.L220.C.20 + LRS1 // Position 10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 12:32 | 2024-06-18 15:02 | 1 | 0 |  | Salut Yannick, les contacts inverseurs sont considérés comme des doubles contacts avec les limites qui vont avec comme tu le vois ci-dessous dans l'extrait de la datasheet des contacts. (ça pourrait être sujet à débat, je l'aurais plutôt classé dans les contacts simples, mais ça a été tranché comme ça). [Image] Best, Nicolas
TCK38370 | User | Service Request | Spares parts | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 12:32 | 2024-06-18 14:12 | 0 | 0 |  | Hi Paul, here are the codes you need for how to order, prices and if there is minimum qties, please check with sales. 11123916 is the sub assebly containg all parts below, you can either order this or the separate parts. 10308170 BEZEL RING 63 1.4404 10308389 JT VOY. MIX3 PVC NITRILE D.63 10308392 VOY 63 US TRIPLEX Best, Nicolas
TCK38348 | User | Service Request | Minimum Scale For MVE7 and MVX7, 0 ...200 psi / Simon | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 10:09 | 2024-06-18 15:45 | 0 | 0 |  | Hi Bee Tin, we only have the drawing for the MVE in 200psi, we would create the one for the MVX if we have an order, sorry. the MVX should have lss graduations as it's only a class 0.6%. [Image] Best, Nicolas
TCK38339 | User | Service Request | Scale in mmHg For MEX5 & MIX7 / Wandee | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-18 08:44 | 2024-06-18 09:32 | 0 | 0 |  | Hi Bee Tin, yes, it's available on both. Best, Nicolas
TCK38323 | User | Service Request | VDS certs. and FM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 18:17 | 2024-06-18 15:17 | 0 | 0 |  | Hey Paul, we don't have any, the last product on which we had the FM was the Y91 and we stopped it a couple years ago for lack of sales. and I never heard of the VDS. But I don't think FM should apply to gauges, it's related to safety equipment for fire and electrycity, it has only been requested on electronic products. if it's a fire safety requirement, on gauges, our MPG6 are UL94 V0 fire resistant. if it's for electrical safety, most our standard gauges are ATEX and IECEX and if it's about pressure safety, they are EN83, respect the PED and some are S3, highested pressure equipment safety level according to the EN. Best, Nicolas
TCK38321 | User | Service Request | PLAN NVX SEPARATEYUR --- DIN 11864-2 ET MINI CLMAP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 17:07 | 2024-06-20 18:01 | 0 | 0 |  | Salut Yannick, voici les 3 premiers 3D, le 1095 va suivre. Best, Nicolas
TCK38320 | User | Service Request | FW: ATB 16134960 and 16130549 - ArcelorMittal (50109648) - Contact Pg's 0-400mbar & 0-160mbar | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 17:04 | 2024-06-18 09:06 | 0 | 0 |  | Hi Rudi, we don't do it in 160mm anymore but you can offer this in 100mm: DPCE160 816 133 136 / 5021 /4710 (160mbar) DPCE160 816 133 156 / 5021 /4710 (400mbar) Best, Nicolas
TCK38305 | User | Service Request | Subject is not set | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 15:24 | 2024-06-20 10:24 | 1 | 0 |  | Véronique, il est vrai qu'on est descendu jusquà 10mbar, mais seulement dans des conditions précises de montage. pour le montage demandé, 25mbar est le minimum possible. Best, Nicolas
TCK38302 | User | Service Request | validation - cuenot - patrick andre--17/06/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 15:15 | 2024-06-17 17:00 | 0 | 0 |  | Oui, cette solution est idéale, après selon le fluide, un capillaire ou un ASIPD peux suffire (s'il n'est pas trop visqueux ou avec des particules). Best, Nicolas
TCK38285 | User | Service Request | TR: [BAKR] T.P inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 13:41 | 2024-06-17 17:22 | 1 | 0 |  | Voici les montages corrigés et les changements apportés: 5 -> 4: 0.1% not possible, the assembly's accuracy will be 0.5% 41 -> 33: connection G1/2 without membrane for fitting with diaphragm seal. 21 -> 00: filling and packing are not relevant anymore when fitted on a diapgram seal. 01 at the end, adding of a filling fluid for the system. PFMN-64.BC3R.A1153.33200.2000 + D821-10L2212.G2001 (min setting 4bar range) PFMN-64.BC5R.A1153.33200.2000 + D821-10L2222.G2001 (min setting 4bar range) N 05013.0003 (Separate accessory ?) TFRH-5132.6006.4051.0110.0500 + AGW-G1212L2D.600 (makes no sense to have a clamp connection on the TFRH if it's to fit it with a thermowell, please redefine with the customer). TFRH-5132.6006.4051.0110.0500 + AGW-G1222L2D.600 (makes no sense to have a clamp connection on the TFRH if it's to fit it with a thermowell, please redefine with the customer). Best, Nicolas
TCK38273 | User | Service Request | RE: Calibration label request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 12:45 | 2024-06-19 12:18 | 1 | 0 |  | Hi Paul, I had to check internally as it's something we never do. but we have options to print a customer specific label and stick it on the gauge (9911/EA08). concerning the information, we can print the serial number of the gauge, the date of calibration and a date 1 year later. I suggest maybe to also print a tag number from the customer if he's using one. It'll be necessary to make a test one for validation before the order and if they need any other marking or number or information, they need to say it now so that we make sure this is an nformation we have access to to print it. Best, Nicolas
TCK38270 | User | Service Request | TR: Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 12:13 | 2024-06-17 14:47 | 1 | 0 |  | It's a bit difficult to be sure with the picture but I'm pretty sure it's this option: [Image] Best, Nicolas
TCK38269 | User | Service Request | TR: SAP code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 12:12 | 2024-06-17 17:47 | 1 | 0 |  | Nathalie, le produit est faisable, il faut demander a Mario Marolf pour l'ouvrir et lui donner la qté prévue en commande. ce sera un Y92318B79R ou Y933.18.B79.R (exemple de code ouvert: 11255931) Best, Nicolas
TCK38255 | User | Service Request | raccord H 2 --55196503 --- 17/06/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 11:11 | 2024-06-17 18:17 | 1 | 0 |  | Bonjour Yannick, comme discuté au téléphone, voici les points clés a transmettre a McPHy, nous sommes complètement d'accord avec leur analyse, les connexions 1/2NPT ou G1/2 ne sont pas faites pour mesurer de lhydrogène sous haute pression nous sommes en train de travailler sur une offre plus complète de connexions dédiées a lhydrogène haute pression, notamment celle que tu mentionne avec étanchéité conique métal/métal, et filet fin inversé qui semble devenir une référence dans le domaine (Parker par exemple ne jure plus que par ça dans ses présentations) par contre attention cette connexion nétant pas strictement normée, certains paramètres tel que le cône ou létat de surface peuvent varier selon les fournisseurs donc bien faire les vérifs' de compatibilité. 2. suite a vos retours et remarques sur nos filets en NPT, je vais donner encore une fois l'alerte et demander des contrôles d'entrée mais je ne peux rien promettre, seul le dept qualité peut s'engager. 3. une alternative en G1/2 (ou 1/4) mais avec joint 316L pourrait être une alternative viable en basse pression. Best, Nicolas
TCK38219 | User | Service Request | TR: NACE siphon? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-17 09:25 | 2024-06-17 18:07 | 0 | 0 |  | Nathalie, la matiere du raccord femelle est NACE, le capillaire aussi, mais ce n'est pas mentionné pour le raccord male, voici a quoi il est conforme: [Image] donc on ne peut pas faire de certificat Nace sur l'ensemble en létat, il faudrait créer un code NACE et sourcer les 2 connections en NACE (possible via une SPR si il y a des quantités). Best, Nicolas
TCK38158 | User | Service Request | DEMANDE CHANTIERS DE L'ATLANTIQUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-14 14:21 | 2024-06-18 09:40 | 1 | 0 |  | Salut Véronique, est ce qu'on doit fournir les puits, ils n'ont pas des fournisseurs AoM répertoriés ? parce que pour l'instant toutes les exigences que je vois sont relatives aux puits. après, si on doit vraiment fournir les puits, je peux demander a nos fournisseurs si ils sont qualifiés et si ils ne le sont pas, il faudra faire une SPR pour demander la création d'AGFR spéciaux chez un fournisseur agréé. Best, Nicolas
TCK38145 | User | Service Request | Question for quotation // Anfrage S // Angebot 20580805 // similar to art.No.11010918 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-14 12:48 | 2024-06-14 15:34 | 1 | 0 |  | Hi Anja, MG5 - 1D3K.B35/4710/0794/1683/LRS1 you can ask for the price and codification (when you have the order.) Best, Nicolas
TCK38137 | User | Service Request | FW: Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-14 10:57 | 2024-06-14 16:19 | 1 | 0 |  | Hi Bee Tin, for me it's this, you can find all information on the website, check with the customer if there is no option missing or something that I would not see from the pictures: TBX100-212.111.20T Best, Nicolas
TCK38128 | User | Service Request | Natural gas approval for article MMD5-D30.B16 // Gas-Zulassung | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-14 09:57 | 2024-06-14 10:48 | 0 | 0 |  | Hi Anja, I don't know about a specific natural gas approval, can they be specific about the norm they want us to comply with ? Best, Nicolas
TCK37965 | User | Service Request | BAPL ¦ TED5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-13 16:01 | 2024-06-14 10:47 | 1 | 0 |  | Hi Estera, here is all the details about the TED that was ordered at the time, as you can see it was mounted on accessory but as they seem to have removed only the TED, maybe we just need to replca it and not the accessories. I would offer all and ask them what they want to replace and what not, and make sure you offer also the certificates: [Image] [Image] Best, Nicolas
TCK37871 | User | Service Request | 3D CAD / Step file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-13 12:09 | 2024-06-18 15:19 | 1 | 0 |  | Hi Paul, here is the 3D of the latest version of the MPG6 (PBT case). Best, Nicolas
TCK37841 | User | Service Request | Quotation - Pressure gauge for Refrigerant R32 / R410a | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-13 09:52 | 2024-06-14 15:48 | 0 | 0 |  | Hi Anja, we can offer our standard version with all those scales, the right connection and filling as below. the dial creation will need to be paid, please check sales for the price. MMD3.DI2.X29/8907/9714/XXXX/DXXX/LXXX Dial: Scale 1: 0...800psi black Scale 2: kg/cm2 Red Scale 3: R32 blue Scale 4: R410A (red) Dial will be slightly different according to our standards and upon order will be send for validation.
TCK37837 | User | Service Request | TR: COMPATIBILITE D924/EJA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-13 09:27 | 2024-06-14 15:50 | 0 | 0 |  | Véronique, on a le feu vert technique pour ce montage, tu peux passer aux ventes pour la définition du prix. et des conditions pour la gestion de la commande transmetteur etc... Best, Nicolas
TCK37810 | User | Service Request | BURAK-GE RENEWABLE-HNI8305502 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-12 17:25 | 2024-06-13 10:32 | 0 | 0 |  | Salut Yannick, oui, il faut que tu lui propose le standard le plus proche, pourquoi pas en MEX3 (72: -1000...60mbar or even -1000...1500mbar). and from that it's the usual dial creation with the colors and options you need. Best, Nicolas
TCK37806 | User | Service Request | MEX5 dial layouts | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-12 16:28 | 2024-06-13 09:10 | 0 | 0 |  | Answered via email with the dials screenshots. Best, Nicolas
TCK37751 | User | Service Request | TED6 serial 4500221973/10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-12 11:53 | 2024-06-13 09:12 | 1 | 0 |  | Hi Daniel, I don't find anything with the number you gave me, I'm sorry. we will need details to redefine it or pictures of the markings and overall product. Best, Nicolas
TCK37742 | User | Service Request | URGENT - TotalEnergies TRA (21106380) - RFQ 2030950 - ...D438.L100.N.60.633R0L.0A000.08... | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-12 11:07 | 2024-06-12 13:42 | 1 | 0 |  | Hi Rudi, seems good to me, there is similar examples with the same configuration (like 11172614), it's all NPT, the AORP fits the range and 3/4 600 RF is ok on the datasheet, you can go ahead and request the price. Best, Nicolas
TCK37717 | User | Service Request | OPTION 0745 --- kem one --- vdav --- 12/06/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-12 09:21 | 2024-06-12 09:48 | 0 | 0 |  | Salut Yannick, comme dans ton code c'est bon, on a quelques codes ouverts et ils est ouvert aussi au config (le plus restrictif) Best, Nicolas
TCK37690 | User | Service Request | R 1/4 ISO 7 connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 17:56 | 2024-06-12 09:18 | 0 | 0 |  | Hi Paul, not available (only one MEX5 with this connection was sold over the last few years) but there is a digit and a drawing existing. you can ask for an offer for it, but make sure you tell us the expected number because it'll impact a lot the price as we will order especially for this order of yours. the digit is T for 1/4 BSP TR. and here is the drawing: [Image] Best, Nicolas
TCK37685 | User | Service Request | Back flange for MX7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 16:42 | 2024-06-11 16:58 | 0 | 0 |  | Hi Paul, the MX7 back flange is welded to the case, it can't be sold as an accessory. Best, Nicolas
TCK37654 | User | Service Request | 30219400--- pressostat - VDAV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 13:46 | 2024-06-11 14:21 | 0 | 0 |  | Yannick, il faudra plus d'information, tu peux lui donner mon contact si besoin mais il faut faire une petite réunion pour en savoir plus. Best, Nicolas
TCK37637 | User | Service Request | TR: M274.D33.X19_D820-10L.42DH.G20.01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 12:01 | 2024-06-11 14:13 | 1 | 0 |  | Bonjour Anne, le montage est OK. Un cadran X19 et mH2O (ou mWG mais pas conseillé car unité équivalent mais moins standard) est possible mais a creer (avec frais). les unités standard ou existantes sont: [Image] Best, Nicolas
TCK37617 | User | Service Request | ISO / EN 17025 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 10:39 | 2024-06-11 17:05 | 1 | 0 |  | Hi Paul, our lab is not EN17025 certified, but we have the option to take the certificates Q016 when the calibration certificate is done by an external COFRAC qualified laboratory (external). Best, Nicolas
TCK37599 | User | Service Request | DEFINITION D910 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 09:07 | 2024-06-14 16:11 | 0 | 0 |  | Veronique, 10mbar pas possible mais le montage sera possible a partir de 25mbar déchelle. EJX110A-JFS5G-919EN + D910LD03X960 (X = 1/4NPTM to match 5 = 1/4NPTF) Best, Nicolas
TCK37598 | User | Service Request | TR: Viton Seal For DANC01 Seal, DN38, MG5-4D30.F20. DANC01.2L2A.01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-11 08:56 | 2024-06-11 10:01 | 0 | 0 |  | Nathalie, c'est ce code: 10306268 on en a 13 en stock. Best, Nicolas
TCK37563 | User | Service Request | BAIT - Intech - price | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 16:58 | 2024-06-10 17:39 | 0 | 0 |  | Hi Francesco, you can have it in your P2F directly: Order code: TSF160-231.160.20T.20/738C_9003_0775 TSF old Code / SAP code: Best, Nicolas
TCK37562 | User | Service Request | Diaphragm Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 16:44 | 2024-06-10 17:32 | 0 | 0 |  | Hi Daniel, the flush threaded would be ideal but as you said the pressure will be an issue. I would suggest a DT2 with a 12m hole in the connection (9026) and a flushing ring. concerning the materials it also gives him a wide variety of choices to be sure it resists it's process. Best, Nicolas
TCK37513 | User | Service Request | Old product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 12:05 | 2024-06-10 15:06 | 1 | 0 |  | Hi Paul, yes we still do it, it's pressure test devices for mercedes. but as far as I recall it can only be ordered through mercedes for the client (and mercedes will place the order to us). Best, Nicolas
TCK37509 | User | Service Request | TR: Ersatzteile für Druckmessgerät Baumer Barflex 4 0-10 Bar - 30209234 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 11:42 | 2024-06-10 15:26 | 1 | 0 |  | Nathalie, please find enclosed the pictures of the adaptor M16F to G1/4M (10309308). I advise to change the NBR part on the bicone and the BS (10303357 + 10303340) Best, Nicolas
TCK37505 | User | Service Request | YOKOGAWA DEFINITION SEPARATEUR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 11:29 | 2024-06-10 14:54 | 1 | 0 |  | Bonjour Véronique, Il va nous manquer la définition du séparateur que tu veux et la connexion "W" du transmetteur napparaît pas sur les datasheets, il faut que l'on sache ce que c'est ? Best, Nicolas
TCK37485 | User | Service Request | YOKOGAWA D910 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 10:16 | 2024-06-10 14:50 | 1 | 0 |  | Bonjour Véronique, pour un montage différentiel en 10mbar, il va falloir faire une demande de faisabilité complète au BE, peux tu me donner les spécification de la bride et du capillaire demandé stp (en mettant un X pour les connections filetées si tu n'est pas sur). tu peux l'envoyer sur ma boite mail directe, je transmettrais à Christophe et sylvain pour la validation. Best, Nicolas
TCK37480 | User | Service Request | 0114-000-014 // Angebot 20579545 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-10 10:05 | 2024-06-10 14:33 | 1 | 0 |  | Hello Anja, only the switch type M is conform with the ATEX intrisic safety of the RP2. it is a mistake in the datasheet that other switches are appearing, I'll sen it to jorg for correction. Best, Nicolas
TCK37426 | User | Service Request | Old Seal | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-07 17:51 | 2024-06-10 14:16 | 1 | 0 |  | Hi Daniel, your definition of the gauge makes sense but for the seal it's something special. it looks like a DT2 top part with an adaptation flange but I don't find anything like this in the system, it's probably too old. I would suggest to take this from the start with the marking on the flange (should contain the norm, size and material) and redefine a solution with a standard (D820/1 could be a better and cheaper option). Best, Nicolas
TCK37411 | User | Service Request | customer logo ¦ Mat3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-07 13:59 | 2024-06-28 09:13 | 1 | 1 |  | Hi Estera, for 500 units, we can do the customer logo with no overcost, you need to order through BHDE. at the order, please also send the company logo and we'll have a design for validation by the customer. Best, Nicolas
TCK37394 | User | Service Request | SEPARATEUR HORS THR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-07 12:02 | 2024-06-07 16:25 | 1 | 0 |  | Bonjour Véronique, le plus simple c'est un: D924_AF6.42GKB7009 il nous faudra la définition complète de l'EJX pour être sur mais de l'EJX430 à 5 bar ne devrait pas poser de problème. Best, Nicolas
TCK37315 | User | Service Request | MIT3 Gauge error on Bourdon Website | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-06 15:31 | 2024-06-06 17:27 | 0 | 0 |  | Hi Daniel, thanks for telling me, it's reported and will be solved quickly. Best, Nicolas
TCK37294 | User | Service Request | replacement ¦ LinkProgrammer ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-06 13:49 | 2024-06-06 15:13 | 0 | 0 |  | Hi Anne, yes, the link was outdated, so the flexprogrammer replaces it and is compatible with the CPX to set the setp values. Best, Nicolas
TCK37287 | User | Service Request | Question about old article 10346659 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-06 12:52 | 2024-06-07 15:35 | 1 | 0 |  | Anja, here is what we can offer as a replacement: MX35D31B26/0794/9711 + Y91333B26R/2151/0794 + ARIDFFF3 + D055-2L15 + LRS1 Best, Nicolas
TCK37266 | User | Service Request | ASIP-D2L3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-06 10:41 | 2024-06-06 11:57 | 0 | 0 |  | Hi Daniel, sorry it was clearer on the old datasheet, here you go: [Image] Best, Nicolas
TCK37238 | User | Service Request | Wake-Frequency-Calculation | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-06 08:46 | 2024-06-06 11:52 | 1 | 0 |  | Hi Sasha, the thermowells you sent me (T6) are mostly for utility on gas flow, and we can't make the wake frequency calculation on them. but we can offer it on the AGF type. if you order it, we will send you a table with the conditions (velocity, viscosity, fluid type etc...) and with the configuration of the AGF we can tell you if it respects the WFC norm and will resist the conditions. Best, Nicolas
TCK37207 | User | Service Request | TR: erreur site internet | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-05 16:12 | 2024-06-05 17:23 | 0 | 0 |  | Nathalie, je préviens Jorg. Best, Nicolas
TCK37188 | User | Service Request | WG: akp_inq_20240605_1445 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fechtig, Kerstin [fke] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-05 15:24 | 2024-06-05 17:28 | 0 | 0 |  | Hi Kerstin, the only product that can go this high in pressure is the MFP that we buy from stiko. but it's handle by you at BHDE, Jorg might be able to tell you if we want to quote on this. Best, Nicolas
TCK37156 | User | Service Request | Indian Manifolds | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-05 13:05 | 2024-06-06 15:10 | 1 | 0 |  | Hi Daniel, as this is all Indian products, I would send the request back directly to Indian sales. you can send the request to Vijayvargia Anil <avijayvargia@baumer.com>. Best, Nicolas
TCK37113 | User | Service Request | préparation réunion H2 Air liquide --- BFI--- 05.06.2024. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-05 09:18 | 2024-06-06 15:11 | 1 | 0 |  | Préparation faite via teams. Best, Nicolas
TCK37031 | User | Service Request | TR: Y92 with Zero Span adjustment | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 14:05 | 2024-06-04 15:45 | 0 | 0 |  | Anne, ce sera un Y92-3.3G.B22R + 2154 + 2151, c'est possible techniquement. par contre il faut que tu fasse la demande au PM pour creer un nouveau code de Y92 et lui demander le prix des options. pour references, ces Y9 contiennent ces options et ce sera les meme sur un Y92: 11251304 11208047 Best, Nicolas
TCK37021 | User | Service Request | BAPL ¦ SPX ¦ DROG101 ¦ GS143202 Pressure gauge 0-160bar - 1szt | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 13:28 | 2024-06-11 10:56 | 1 | 0 |  | Hi Estera, no problem for the DROG gauge but the tubular seal was not sold since very long and the only thing we have is a drawing. if we want to do it now, we need to find the machining company for the body, order the counter parts to be able to test it, create codification and components (it'll be extremely expensive). I would suggest to offer the DROG101/211.133-135.0156.0771.3272.0053 and to orient the customer toward one of our more standard tubular seals: https://www.bourdon-instruments.com/fr/en/product-overview/diaphragm-seals/in-line-seals/c/38217/products Best, Nicolas
TCK37004 | User | Service Request | WG: MEX8-D60.B16 & DİAPHRAGM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fechtig, Kerstin [fke] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 11:01 | 2024-06-04 11:40 | 0 | 0 |  | Hi Kerstin, the diaphragm seal design in the picture looks like our DT3 but the DT2 will be cheaper and the design diference only makes it more compact. the connection on the drawing and picture looks like a G1/2 so I would suggest: MEX8D30B16 + DT22L200230.01 (be careful to check the dimension of the gauge and the connection as I don't have any measure on the drawing, just visual) an even cheaper solution would be the: MEX8D30B16 + D042 2L03 the design is a bit different (all welded and compact) but it's an alternative considering the data I have. Best, Nicolas
TCK36993 | User | Service Request | 2137317_Anfrage 3D STEP-Model | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Korner, Carla [koca] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 10:15 | 2024-06-04 10:22 | 0 | 0 |  | Hi Carla, please find the 3D in the ticket. Best, Nicolas
TCK36976 | User | Service Request | CUENOT-PATRICKANDRE-04.06.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 09:00 | 2024-06-04 10:17 | 0 | 0 |  | Salut Yannick, il a raison sur le début, si il choisi une surpression max de 6 bar, il peut bien mettre 6 bar en BP ou en HP avec l'autre à 0 et ça résistera. pour ce que ça change, ça a du sens pour nous en terme de codification et pour mettre nos limites de pressions, mais pour lui le plus intéressant est de prendre la surpression de 6 bar, il n'a pas de déclassement et pas de prix plus haut. Best, Nicolas
TCK36958 | User | Service Request | What is the graduation per division for MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Lim, Janice [ljan] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-04 07:33 | 2024-06-04 10:12 | 0 | 0 |  | Hi Janice, here is the standard dial of a MEX5...B19: [Image] Best, Nicolas
TCK36940 | User | Service Request | 3D drawings | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-03 15:54 | 2024-06-03 17:31 | 0 | 0 |  | Hi Kerstin, the different window or range will not have an impact on the drawing so only one drawing covers your request, I link it tin the ticket. Best, Nicolas
TCK36934 | User | Service Request | DASHPOOT --- magyar --- aroe --- 03.06.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-03 15:25 | 2024-06-04 09:51 | 1 | 0 |  | Yannick, le Dashpot sera dispo sur les MEP5/6/8. Best, Nicolas
TCK36890 | User | Service Request | Enquiry Errel | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-03 11:36 | 2024-06-04 11:30 | 1 | 0 |  | Hi Kerstin, the MEX5 is from us but the NW-16KF Flange is in my opinion an accessory added by a third party, I don't find anything in our system and it doesn't mean anything for the R&D team. it seems like it 's a connection norm, it could be this kind of accessory: [Image] Best, Nicolas
TCK36862 | User | Service Request | FW: MPG6 + Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-03 10:08 | 2024-06-03 10:55 | 0 | 0 |  | Hi Andrew, yes, I checked with Jerome and it's spare parts for SATORP, it's already offered locally. Best, Nicolas
TCK36856 | User | Service Request | TR: Neobee M-20 Food Grade -25 to 205 Deg | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-06-03 09:58 | 2024-06-03 10:52 | 0 | 0 |  | Bonjour Anne, nous ne pouvons pas introduire dans nos process des huiles de remplissage a la demande du client. par contre notre LRS10 répond aux critères de haute température et de FDA (MSDS ci-jointe). Best, Nicolas
TCK36788 | User | Service Request | TR: Kindly confirm | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-31 14:07 | 2024-06-03 10:49 | 1 | 0 |  | Bonjour Anne, j'ai aussi eu la question d'Anil, c'est répondu. Best, Nicolas
TCK36780 | User | Service Request | RPPE-8A6.204/SETP / Tia | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-31 11:49 | 2024-06-03 10:48 | 1 | 0 |  | Hi Bee Tin, your code is fine and the setp also, here are your max and min deadband and what it means in terms of set and reset points: RPPE-8A6.204/SETP (2bar falling) min DB: 0.5bar (2bar falling, reset at 2.5bar rising) max DB: 4.5bar (2bar falling, reset at 6.5bar rising) Best, Nicolas
TCK36778 | User | Service Request | RT-NEA.303.E00EJ/SETP / Tia | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-31 11:39 | 2024-06-03 10:44 | 1 | 0 |  | Hi Bee Tin, your code is fine and the setp also, and the thermowell will fit. here are your max and min deadband and what it means in terms of set and reset points: RT-NEA.303.E00EJ/SETP (100C rising) min DB: 5C (setp 100C Rising, reset at 95C falling) max DB: 24C (setp 100C Rising, reset at 76C falling) Best, Nicolas
TCK36763 | User | Service Request | AW: Manometer Kältemittel R404A | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-31 10:08 | 2024-06-03 10:37 | 1 | 0 |  | Hi Francesco, there is something wrong in the codes created in the system, this is being reviewed by Jorg. I add you in the email loop. Best, Nicolas
TCK36744 | User | Service Request | R: BAIT - offer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-31 08:05 | 2024-06-03 10:35 | 1 | 0 |  | Hi Francesco, yes, the RPPE7 should be the right one, for example: RPPE7A3600 Best, Nicolas
TCK36721 | User | Service Request | Ministerie van Defensie - MEX3 mounting bracket "G" | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-30 16:48 | 2024-06-04 09:52 | 1 | 0 |  | Hi Rudi, the case is not the same between the type G (crimped bezel) and the type F (bayonett). they'll need to replace the gauges completely. Best, Nicolas
TCK36708 | User | Service Request | FW: RFQ Bourdan Pressure Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-30 15:07 | 2024-06-03 10:30 | 1 | 0 |  | Hi Anil, yes the code already exists: 11247024 MX7-A30.B15H.AKPL.DT2.2L90.0.260.01/9711 SAP code Designation Qty 11247026 in this code the 9711 is a inscription "made in france" on the dial. Best, Nicolas
TCK36697 | User | Service Request | MPJ ? - Can you help me out | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-30 13:59 | 2024-06-03 10:17 | 1 | 0 |  | Hi Rudi, it'll be an MPJ nothing changed on this one (PolyPropylene case). Best, Nicolas
TCK36678 | User | Service Request | Hydrogen Thermowells | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-30 10:38 | 2024-06-03 10:09 | 1 | 0 |  | Hi Daniel, for the differential transmitters, with H2 it might be interesting to not carry the pressure but a signal, so a 2 transmitters solution could be smart. considering the temperature, we internally don't have dedicated solutions yet but we can check with our suppliers of thermowells. I'll ask them for information. Best, Nicolas
TCK36677 | User | Service Request | BH1 MSDS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-30 10:33 | 2024-06-03 10:05 | 1 | 0 |  | Hi Paul, I sent him the MSDS directly and you're in cc. Best, Nicolas
TCK36647 | User | Service Request | MX7 with DT2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 17:20 | 2024-05-29 18:06 | 0 | 0 |  | Hi Sascha, yes, as it's a differential, you have 2 pressure entry which means that you'll have 2 capillaries and 2 diaphragm seals. in the example below (very close except the option is a marking, not the special dial, you can see the detail of the elements: 11247024 MX7-A30.B15H.AKPL.DT2.2L90.0.260.01/9711 SAP code Designation Qty 11247026 MX7-A30.B15H/9711_D58210264889 AKPL-BF530210356150 DT2-2L90.0.260 Best, Nicolas
TCK36611 | User | Service Request | MPG6 + Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 14:11 | 2024-06-03 17:40 | 1 | 0 |  | Hi Andrew, this is a SATORP project and the offer was already handled locally, you can forget about it. Best, Nicolas
TCK36604 | User | Service Request | Springs for TBA072 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 13:46 | 2024-05-29 17:11 | 1 | 0 |  | Hi Sascha, I can't do that anymore, but I can give you the part number and you can request to order it for free delivered to the stockach office but only the sales can validate this. 10329589 Best, Nicolas
TCK36573 | User | Service Request | 2.Scale / Special | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 10:20 | 2024-05-29 16:28 | 1 | 0 |  | Hi Sascha, this is possible except the second scale can't be 400psi, it can be 400 but we can't print a unit that is not real. otherwise not problem, you order the dial creation (creation of the design) on top of the option 9712 for dual scale. Nathalie can give you an offer if you know the rest of the gauge details. Best, Nicolas
TCK36572 | User | Service Request | TR: Manometer Kältemittel R404A | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 10:17 | 2024-05-29 17:56 | 1 | 0 |  | Nathalie, je cloture sur le ticket. tu peux me renvoyer la demande par email stp ? il va falloir voir avec Jörg, il y a quelque chose qui cloche dans les code. Best, Nicolas
TCK36564 | User | Service Request | TR: onze prijsaanvraag 6000270701 [**EXTERNAL**] | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 09:53 | 2024-05-29 16:23 | 1 | 0 |  | Salut Anne, on a pas de niveau magnétique, les capteurs de niveau ce sera plus BADK. Best, Nicolas
TCK36557 | User | Service Request | TR: Questio 10325918, raccord a braser G1/2 laiton | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-29 09:28 | 2024-05-29 09:50 | 0 | 0 |  | Vu, commander chez Stockach: [Image] Best, Nicolas
TCK36526 | User | Service Request | SUPPORT MARINE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-28 17:25 | 2024-06-05 17:29 | 1 | 1 |  | Veronique, on ne trouve rien, tu peux leurs dire qu'on a pas d'infos sur ce qu'ils demandent. Best, Nicolas
TCK36459 | User | Service Request | ARPX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-28 09:35 | 2024-06-13 16:45 | 1 | 1 |  | Salut Véronique, je suis désolé, cétait long et en plus la réponse est négative, cet ARPX n'a pas la certification ACS mais en plus schneider ne propose pas de modèles ACS. Best, Nicolas
TCK36419 | User | Service Request | TR: CAF - Fwd: RV: CR1 - Mallorca - transductor X.18X02212X01 ¦ OP CAF Construcciones y auxiliar 12062023 20522046 ¦ Ped 1037 PO 1410587 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-27 16:39 | 2024-05-27 16:43 | 0 | 0 |  | Bonjour Jacques, oui, le connecteur est bien inclus. [Image] Best, Nicolas
TCK36389 | User | Service Request | TR: new quote | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-27 14:26 | 2024-05-27 14:47 | 0 | 0 |  | Salut Nathalie, j'ai bien peur qu'on ai pas grand chose à proposer. le TED535B77R/0609/SETP est le remplacement fonctionnel le plus proche mais la différence majeure est qu'il n'est pas SIL2 on a plus aucun produit SIL2) un autre remplacement potentiel serait le PFMN (avec des contacts numeriques mais ATEX). Best, Nicolas
TCK36356 | User | Service Request | Need confirmation for code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-27 10:40 | 2024-05-27 11:17 | 0 | 0 |  | Hi Anil, sure, here is the code: 10285367 MEX5-D61.B26/0751 Best, Nicolas
TCK36351 | User | Service Request | HSAN-RPPN-SAM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-27 10:16 | 2024-05-27 11:14 | 0 | 0 |  | Bonjour Yannick, les points que tu demande ne sont pas possibles, au mieux on peut proposer 8bar à la hausse sur le micro 1 et 15bar à la baisse sur le micro 2. Best, Nicolas
TCK36244 | User | Service Request | WG: DRO100 vs. MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-24 09:22 | 2024-05-24 09:48 | 0 | 0 |  | Hi Sascha, Right, I should've been more extensive with the explanations: The LRS1 is the closest in terms of temperature coverage and type of application and yes there is a slight gap between the 2. The option 0201 would not make sense as the max pressure for a DADL is 40bar anyway. the replacement by new references will always imply changes even if small (for example the grade of stainless for some parts from DR to MEX, the shape of the connection square, the dial design etc...) so for replacement a review of the new code with the datasheets for the client is always necessary. Concerning the type D or F, that's on me, it doesn't show in the english text, but I trust you to correct this in the reference. Best, Nicolas
TCK36163 | User | Service Request | product identification | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-23 11:52 | 2024-05-27 14:30 | 1 | 0 |  | Hi Sascha, Sorry it was a bit long but it's an order from 2013 and it was not really codified in SAP. this is a: DP 460 BF V17.58.82.93 / FP441-S5 Wall mounting Oil filled case Safety glass Metal plate Calibration certificate Dial 100 mm Range 0/0.1 bar Accuracy class 1.6 @ temp.15/25C with seals on same level and same temperature 2 x Diaphragm seals FP441 Process connection flange DN25 PN16 Capillary 5000 mm And it would translate in our system to this code: MFT5-A13.N08X.X2.100/FP441-S5 (DN25 PN16) It's not a configuration so the final code can be subject to small changes (I'm not sure if the seal will be just FP441-S6 for example or if it'll be with the special seals code: XDSE:FP441-S6 or which details will be added to the short codes). I already requested an offer from fantinelli and will forward it to the sales so that they can give you a price for replacement of the part. Best, Nicolas
TCK36152 | User | Service Request | how much angle | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-23 10:37 | 2024-05-30 10:09 | 1 | 0 |  | Hi Zhang wei, I checked some and there is no way a pressure can damage the movement or that the pointer could turn more than one turn under the influence of the pressure and mechanism. BUT, if you're on overpressure (too high differential pressure), the sector could unconnect from the pointer and leave it free. if it happens, the overpressure was too high and the gauge needs to be check before reconnecting the pointer. you can open a return procedure if you want us to check the gauge (with the overpressure, it'll be at customer's expenses I think. Best, Nicolas
TCK36130 | User | Service Request | MR5-6A33.B18. AKPL2F5310. DT3.2L20.0.220.01 / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-23 07:46 | 2024-05-23 14:52 | 1 | 0 |  | Hi Bee Tin, yes, perfectly fine. Best, Nicolas
TCK36108 | User | Service Request | Y91 sur séparateurs en hastelloy | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-22 17:00 | 2024-05-22 17:39 | 0 | 0 |  | Bonjour Cyrille, Ces montage sont possible (j'ai rajouté un liquide de remplissage) et les séparateurs sont au P2F. Y91-3.2.3.B20.R_DT2-2H60.6.620.01_9010_0592 Y91-3.2.3.B22.R/2164_DT2-2H60.6.620.01_9010_0592 Mais pour standardiser et gagner en coûts sans impact pour le client je conseille quand même de passer en G1/2 entre le séparateur et le transmetteur. Il se peut aussi que ça donne un meilleur délai à la commande car on a plus de stock de flasque sup de DT2 en G1/2. Y91-3.3.3.B20.R_DT2-2L60.6.620.01_9010_0592 Y91-3.3.3.B22.R/2164_DT2-2L60.6.620.01_9010_0592 Best, Nicolas
TCK36105 | User | Service Request | stp. files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-22 16:25 | 2024-05-23 15:22 | 0 | 0 |  | Hi Sascha, We only have the drawing of the AMFD (enclosed) but you can still order the creation of .stp drawings via the sales administration team. Best, Nicolas
TCK36093 | User | Service Request | Spring of TBA072 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-22 15:05 | 2024-05-22 17:13 | 1 | 0 |  | Hi Sascha, there is not much I can say, we usually don't have complaints, it would be good to make sure that they remove the hook of the spring before going around the pipe too attach, but it should not be a problem normally. Best, Nicolas
TCK36091 | User | Service Request | DRO100 vs. MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-22 14:53 | 2024-05-23 15:12 | 1 | 0 |  | Hello Sascha, here is the replacement we would advise: MEX5D31B27 + DADL032L2A01 please just check if they need it electropolished ? Best, Nicolas
TCK36010 | User | Service Request | Urgent MEX8 question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-21 17:07 | 2024-05-22 17:07 | 0 | 0 |  | Hi Paul, I would have been in favor of doing it but here is what is possible with the dashpot (0670 is possible also to add for visual purpose but doesn't make much sense in terms of accuracy). MEX8-D90.B16/0122_0655 (class 1.6%) 9 pcs. MEX8-D90.B18/0122_0655 (class 1.6%) 9 pcs. MEX8-D90.B19/0122_0655 (class 1%) 9 pcs. MEX8-D90.B20/0122_0655 (class 1%) 9 pcs. MEX8-D90.B22/0122_0655 (class 1%) 15 pcs. MEX8-D90.B24/0122_0655 (class 1%) 12 pcs. MEX8-D90.B26/0122_0655 (class 1%) 12 pcs. MEX8-D90.B27/0122_0655 (class 1%) 12 pcs. MEX8-D90.B29/0122_0655 (class 1%) 12 pcs. MEX8-D90.B31/0122_0655 (class 1%) 6 pcs. MEX8-D90.B33/0122_0655 (class 1%) 12 pcs. MEX8-D90.B35/0122_0655 (class 1%) Best, Nicolas
TCK36001 | User | Service Request | DPC100 question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-21 16:19 | 2024-05-23 15:07 | 0 | 0 |  | Hi Paul, I'm sorry but it's too low, it's already complicated to do it in class 1.6. Anyway, I suggest a global meeting on this order with Jorg to decide how to handle. you'll have the other answers this afternoon. Best, Nicolas
TCK35989 | User | Service Request | Insurance certificate | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-21 15:02 | 2024-05-21 16:15 | 0 | 0 |  | Hi Paul, please find enclosed our insurance certificate. Best, Nicolas
TCK35971 | User | Service Request | LOCTITE 266 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-21 11:51 | 2024-05-21 17:47 | 1 | 0 |  | Salut Véronique, la loctite 266 n'est pas PMUC mais les centrales n'utilisent pas de montages sur séparateur a ma connaissance ? on a beaucoup de différentes loctite, il y a peut être une confusion appelle moi sur ce point, ce sera plus simple de remettre dans le contexte. Best, Nicolas
TCK35892 | User | Service Request | Temperature Gauge Special type / Aircond Network | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-20 10:48 | 2024-05-21 09:29 | 1 | 0 |  | Hi Bee Tin, the TBX doesn't exist anymore but can be replaced by a TB63 ike below: [Image] On the picture from your customer we see only the upper and gland parts of the sliding connection, it will be necessary to know what the bottom part is (probably still connected to his process). Also a standard 100mm leangth seems to fit (approximately, to be confirmed by the customer) Best, Nicolas
TCK35780 | User | Service Request | FW: QU3549785 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Shutt, Oliver [shol] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-17 15:56 | 2024-05-17 16:33 | 0 | 0 |  | Hi Oliver, the gauge is 10290929 MJX7-461.X19/0679_9709_0693_D000. but there might be an ARPX (probably a 10265178 ARPX-06N) you can check with the customer, the datasheet is enclosed in the ticket. Best, Nicolas
TCK35776 | User | Service Request | BAPL ¦ replacement ¦ E933-18-B27-R/2164 ¦ zapytanie Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-17 15:02 | 2024-05-17 15:42 | 0 | 0 |  | Hi Anna, the E93318B27R/2164 was basically a Y91333B27R/2164 mounted on a Clamp DN38 ISO. you can either relpace it by a Y91333B27R/2164 + DANC01-2L2A.01. or go for a Baumer Sensor with clamp connection like the PBMH. (2164 is M12 4 pins connector). Best, Nicolas
TCK35717 | User | Service Request | Scale 0....160bar DR80 gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-17 08:38 | 2024-05-17 09:25 | 0 | 0 |  | Hi Sascha, All things on the dial are according to the reccomandations of the EN837 and made to ensure the easiest and most accurate reading in regards of the accuracy. then if a change is necessary for a specificity of a customer, you can request a specific dial, there will be a fee for the creation and the procedure for the approval. Encolsed the standard dial (picture). Best, Nicolas
TCK35674 | User | Service Request | TR: MP_6 pointer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 16:00 | 2024-05-17 09:50 | 1 | 0 |  | Bonjour Anne, l'aiguille du MPx6 est telle que ci-joint au ticket, elle est réglable micrométrique et fine mais pas chantournée. le design de ce produit est très établis et jamais une aiguille chantournée n'a été fournie dessus (de mémoire de SAP). Best, Nicolas
TCK35658 | User | Service Request | total --- 6700160917 --- grch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 15:25 | 2024-05-16 17:25 | 0 | 0 |  | Vu en teams, code existant envoyé par Josette en X15. Best, Nicolas
TCK35615 | User | Service Request | RFQ Meca-Fluid (21107824) - Like info | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 14:10 | 2024-05-16 15:47 | 0 | 0 |  | Hi Rudi, I don't find the serial number in our database sorry, but I defined what I could and I have the matching codes below in the system, can you check if you sold it in the past to this customer or a distributor ? 11106018 11106050 potentially more the first one. Best, Nicolas
TCK35602 | User | Service Request | CIRA-MDJ-16.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 13:12 | 2024-05-16 15:24 | 1 | 0 |  | Salut Yannick, ce n'est pas du tout commun mais on a le raccord au dos + tube en monel et le boitier avec fixation murale, tu peux demander ce code. Best, Nicolas
TCK35573 | User | Service Request | TR: 50148036 -- epsa -- 2188451- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 09:41 | 2024-05-16 15:19 | 1 | 0 |  | Bonjour Priscilla, c'est réalisable, il faut juste ouvrir et changer le cadran et celui-ci existe déjà pour l'impression. attention à bien faire une procédure de retour et à penser à changer le code pour refléter les changements apportés au produit. Best, Nicolas
TCK35566 | User | Service Request | TR: OP 20567308 --- air liquide --- QUESTION TECHNIQUE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-16 08:38 | 2024-05-22 16:34 | 1 | 0 |  | Voici mes réponses: Pouvez-vous me confirmer que tous les composants peuvent bien supporter une pression ambiante allant de 700 à 1300 mbara ? TT GAMME BIN: Aucun test spécifique n'a été fait à ma connaissance mais vu la conception ça pourrait entraîner un léger drift du 0 (réglable) mais pas de dommages. Concernant les 2 pressure transmitter calibrés entre 0 et 10 bara, je vois dans la datasheet que la surpression admissible est de 20 bar. Hors, notre pression de design est de 25 bara. Le capteur peut-il supporter cette pression ? Y91 BIN: Le Y91 ne peux pas (il résistera car la pression déclatement est de 30bar mais il risquera de dériver), mais il y a plusieurs solutions telles que prendre un 16bar (proof pressure 32bar), prendre un PBxx de 16 réglé à 10bar, utiliser un limiteur de pression etc... Le test en pression (hydraulique ou pneumatique) à 25 bara n'apparaît pas dans l'offre, est-il bien prévu ? Juste sur MA7 et AMFD --- > test fuite hélium sur capteur électronique BIN: ça concerne plutôt la mise en pression dépreuve, je conseille plutôt de faire un Q023 (avec un transmetteur compatible bien sur voir ci-dessus) Pouvez-vous me confirmer que le test de fuite à l'Hélium sera bien réalisé pour les capteurs de pression au niveau de la cellule et au niveau du raccord et pour les manifold au niveau des internes de la vanne et au niveau des raccords ? BIN: pour chaque code SAP, lintégralité du produit sera monté sur la machine de test hélium pour être testé. Est-il possible de verrouiller les manifolds en position ouverte / fermée ? BIN: Par sur ceux qu'on utilise actuellement mais on peut faire une demande de prix en négoce pour des spécifiques. Pouvez-vous me confirmer que les têtes électriques des capteurs sont en matériau étanche et non inflammable ? Y91 connecteur BIN: si DIN43650: Molex 1210230341 Polypropylène https://www.molex.com/en-us/products/part-detail/1210230341?display=pdf Pouvez-vous me confirmer qu'un emballage spécifique (plastic, sachet déshydratant, mention warning) et individuel est prévu tel que décrit dans notre TS ? merci de me transmettre photo dun emballage produit dégraissé O2 BIN: procédure ci-jointe (photo a venir). Enfin pouvez-vous me confirmer la fourniture de la procédure de nettoyage oxygène, des données de calibration Q008-Q005 , du rapport de test en pression Q002- est-ce possible sur amfd , du certificat CE, et d'un certificat de conformité à la commande Q001 a rajouter sur offre ? Best, Nicolas
TCK35524 | User | Service Request | Matériau boîtier CTX / Y91 et Pbsn | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 16:12 | 2024-05-16 15:10 | 1 | 0 |  | Bonjour Jacques, les parties en contact sont bien en 316L (et la céramique et le joint bien sur). [Image] le boitier est en 304 [Image] ce sera le cas aussi pour le Y9 et le PBSN comme tu as vu sur les datasheets. pour faire une version spécifique en corps 316L, il faudrait passer par une SPR. best, Nicolas
TCK35520 | User | Service Request | MEP-8-15.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 15:39 | 2024-05-16 18:03 | 1 | 0 |  | Salut Yannick, a ce jour on part sur septembre. Best, Nicolas
TCK35515 | User | Service Request | RV: : Mex5 class 1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 15:01 | 2024-05-16 14:33 | 1 | 0 |  | Hi Angel, the ANSI and the EN don't define the dials the same way. psi based dials are according to the ANSI as standard and so the dials are according to the ANSI which doesn't define accuracy classes nor request it to be displayed on the dial. your idea is good, bar/psi will be an EN dial and will display the class on the dial. Best, Nicolas
TCK35494 | User | Service Request | TR: baes-1534-15.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 11:58 | 2024-05-17 12:07 | 1 | 0 |  | Nathalie, notre client est BAFR et pour moi ils ont le droit de commander des séparateurs seuls (sur les deux dernières années ils l'ont fait régulièrement et on ne sait pas forcement ou ça va, c'est la responsabilité de BAFR). le système de remplissage sur un 1534 est similaire, je ne vois pas de raison technique de refuser. Best, Nicolas
TCK35491 | User | Service Request | inquiry/Step file for 1 and 10 pieces of MCX5-D30.N66.D821-10L.42CD.B20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Mörler, Kevin [mkev] | Birrien, Nicolas [bin] | DEFRI - Friedberg | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 11:33 | 2024-05-16 16:50 | 1 | 0 |  | Hello Kevin, I attached in the ticket the 3D of the MCX but I don't have it for the D821, sorry. It can be requested at the order, for this I let you ask the sales administration for the code to use and price. MCX5-D30.N66.D821-10L.42CD.B20 Best, Nicolas
TCK35484 | User | Service Request | Customer request: product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 11:17 | 2024-05-16 14:21 | 1 | 0 |  | Hi Anja, if there is a leak, it needs to be replaced. we can't offer this specific gauge anymore but a MEX3 could be a good replacement, MEP3 if they need a safety and maybe MAG3 if Jorg think it's a good idea (I personaly think, based on the application that a MEX3 could be a good fit). Can a commercial call them to define the replacement (it's too old to find a part number in the system and I don't see the connection or the pressure range). Best, Nicolas
TCK35467 | User | Service Request | MEX8 with Dial Scale -760mmHg to 0 / Azzam | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-15 09:25 | 2024-05-16 14:09 | 1 | 0 |  | Hi Bee Tin, I'm sorry, I was a bit surprised but it's not existing you'll need to create it. otherwise, the X59 (-76...0 cmHg) or the H59 (-30...0 "Hg) are existing. Best, Nicolas
TCK35392 | User | Service Request | MIX7 + DAVA ? possible ? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-14 10:58 | 2024-05-16 14:00 | 1 | 0 |  | Anne, c'est OK, c'est juste assez rare d'avoir des grands diamètres sur des applications agro-alimentaires. Best, Nicolas
TCK35386 | User | Service Request | TR: open questions CTX and SPR proposal PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-14 10:35 | 2024-05-16 13:58 | 1 | 0 |  | Anne, specific range on CTX needs to be validated by the PM and he'll ask for qties. the PBSN can be selected with B26 and then set at 0...20bar. Best, Nicolas
TCK35359 | User | Service Request | TR: Refrigerant gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-14 09:27 | 2024-05-21 09:31 | 1 | 0 |  | Hi Nathalie, the exact dial doesn't exist but is possible. we have the code 10280014 that is exactly the same (except all black and without the marking POE Fluid) for reference. Best, Nicolas
TCK35310 | User | Service Request | Customer request: product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-13 15:55 | 2024-05-16 11:49 | 1 | 0 |  | Hi Anja, this is an old product that was phased out long ago. we could offer a standard MEX3 for the replacement or a MAG3 upon acceptance from Jorg (not sure we want to sell this product for low qty europe application). Best, Nicolas
TCK35298 | User | Service Request | Sil / MTTF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-13 15:00 | 2024-05-16 11:45 | 1 | 0 |  | Hi Paul, please find enclosed the MTBF analysis for the RPPY based on the microswitch (most sensitive part). Best, Nicolas
TCK35255 | User | Service Request | MEX8 Blow Out protection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-13 11:54 | 2024-05-16 11:39 | 1 | 0 |  | Hi Daniel, the MEX8 have a small blow out device on top (like all our gauges). if they want a partially blow out back, then I would suggest the MJX7which have a part blow out back and a solid front. in the future (normally in a couple weeks / months) the MEP8 will be released and be the perfect fit. Best, Nicolas
TCK35202 | User | Service Request | MANOS AU VIDE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-13 09:10 | 2024-05-16 11:26 | 1 | 0 |  | Bonjour Véronique, le tube 0...100bar est costaud, il résistera sans problème à un vide grossier occasionnel. il faut faire attention à ne pas avoir de stop pin à 0 pour que ça ne force pas sur l'aiguille. Best, Nicolas
TCK35194 | User | Service Request | MONTAGES LOCTITES | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-13 08:51 | 2024-05-16 17:43 | 1 | 0 |  | Salut Véronique, le montage à la loctite entre les instruments et accessoires est l'option 0435. ce sera de la Loctite 266. on ne conseille pas de démonter nos montages, mais en cas de nécessité, il faudra démonter en faisant attention de bien prendre les clés sur les carrés ou hexa, puis de bien enlever toute la loctite avant de remonter avec une nouvelle couche. Best, Nicolas
TCK35089 | User | Service Request | 2 questions about CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-11 12:30 | 2024-05-16 11:21 | 1 | 0 |  | Hello Xu Qiang, the rising min time and falling min time will be the same: [Image] the max voltage power supply is 30V DC on a 0...10V output: [Image] Best, Nicolas
TCK35075 | User | Service Request | FW: Product confirmation pressure switch and temperature switch / Gilang | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-10 10:53 | 2024-05-17 16:49 | 1 | 1 |  | Hi Bee Tin, here is the best I was able to do from the picture, the ZT and the ZD at least need to be redefined or we need more information if there is a plate on the side or a marking I don't see. Tag n Part n cde Pos. 20LAC20CP107SHM Best, Nicolas
TCK35074 | User | Service Request | Flush diaphragm D200-2L20-LRS1 / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-10 10:46 | 2024-05-17 11:50 | 1 | 1 |  | Hi Bee Tin, here is the conversion, to be confirmed it's ok with the customer: Order code: MEX5-D30.B20.DT2.2L20.0.000.01 MEX old Code / SAP code: Best, Nicolas
TCK35071 | User | Service Request | replacement of old RDPNAE6156 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-10 09:12 | 2024-05-16 10:57 | 1 | 0 |  | Hello Xu Qiang, depending on the requested overpressure resistance, all three codes below are potential replacement. RDE-5A6.156 RDE-4A6.156 Also I would reccoment to add the SETP option for factory setting of the setpoint. Best, Nicolas
TCK35068 | User | Service Request | : Mex5 class 1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-09 23:12 | 2024-05-16 10:50 | 1 | 0 |  | Hi Angel, as psi is the leading scale, the norm applied is the ANSI which doesn't require the marking of the class. of course the MEX5 you have is still class 1(+/-1% over the full scale) but this is not recognised by ANSI. Best, Nicolas
TCK35007 | User | Service Request | BAPL ¦ special scale for MEX3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-08 09:47 | 2024-05-21 09:34 | 1 | 1 |  | Hi Estera, here is what a Bar Bourdon (fine scale, digit "A") look like on a 6bar MEX3: [Image] Best, Nicolas
TCK35004 | User | Service Request | drawing of TED | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-08 09:29 | 2024-05-16 10:45 | 1 | 1 |  | Dear Zhang Wei, please find enclosed the .stp drawing. Best, Nicolas
TCK34915 | User | Service Request | 4515999649---DOW---GRCH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-07 10:46 | 2024-05-07 12:01 | 0 | 0 |  | Yannick, oui ce sera bien le meilleur montage, je préciserais le réglage de l'AORPB (a 100mbar). Il te faut un ARIN2606 et mettre toutes les connections en 1/2NPT mais sinon c'est ca. Best, Nicolas
TCK34899 | User | Service Request | Faulty Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-07 09:50 | 2024-05-07 11:40 | 0 | 0 |  | Hi Andrew, you're right and it seems there is something wrong here; we think it might be the bezel ring being undertight. could you please send a picture of the top back of the case, it would help us confirm. then if it is what we think, a return procedure needs to be launched and we will take the gauges back, correct this and resend. Best, Nicolas
TCK34889 | User | Service Request | TR: check product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-07 08:42 | 2024-05-07 10:24 | 0 | 0 |  | Nathalie, avec la codif en D le client aura seulement la graduation en kPa, pour avoir la double graduation, ce sera un X et 9712. je n'ai pas trouvé de code existant dans SAP, peut être pas vendu depuis Bourdon Haenni. Best, Nicolas
TCK34866 | User | Service Request | Gauges with D055 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-06 18:03 | 2024-05-07 10:09 | 0 | 0 |  | Hi Sascha, I don't even have to check, the codes you have in yellow are open and ready to be ordered in SAP. I don't see what blocks you. Best, Nicolas
TCK34826 | User | Service Request | danae--projet--AROE--06.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-06 13:45 | 2024-05-06 14:40 | 0 | 0 |  | Salut Yannick, c'est un M11 lancêtre des manos à contacts je te mets en PJ la datasheet: le plus proche pour remplacer sera un M17 4D30 B29 ce sera juste du 150mm au lieu de 160mm. Best, Nicolas
TCK34809 | User | Service Request | TRANSFESA - 20569189-AROE-06.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-06 12:07 | 2024-05-06 14:24 | 1 | 0 |  | Salut Yannick, ci-dessous les détails de ton code SAP, y compris les numéros de plan client, d'homologation etc...: ci-joint le plan client. le plan que tu as joint est un plan similaire. FICHE PRODUIT - MANOMETRE FERROVIAIRE / PRODUCT FILE - RAILWAY PRESSURE GAUGE Internal codeSAP code (MFE)DiameterCircuit numberPressure (bar)Accuracy 7b/0,05 b to 5b and 0,1 b from 0,5 to 7bConnectionMounting studLeft locating pin grooveRight locating pin grooveLighttingPointer beneath dial (H+6)Case colourDial colourLettering colourRoundel colourSpecifics markings Front pointer colourFront pointer black heelRear pointer colourRear pointer black heelTrim colourTrim screw colourInternal drawing nBH parts list drawing nCustomer drawing Customer part list drawingHomologationRemarksTechnical documentDial screen printingFront guard screen printing ResolutionInternal user's manualCustomer article code Best, Nicolas
TCK34741 | User | Service Request | Enquiry For MEX8D60B26/0751+D1012N20260+LRS1 (103265648)/ Servo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-06 07:20 | 2024-05-06 14:12 | 1 | 0 |  | Hi Bee Tin, here is the replacement for MEX8D60B26/0751+D1012N20260+LRS1: MEX8D60B26/0751 + DT12N200260 + LRS1 Best, Nicolas
TCK34681 | User | Service Request | TR: Confirm if pressure unit in mmWC is possible for model MCX7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-03 14:42 | 2024-05-03 15:13 | 0 | 0 |  | Bonjour Priscilla, aucun problème, par contre ce sera noté -40 ... 60 mmH2O (pas mmWC) mais ça veut dire exactement la même chose. Best, Nicolas
TCK34677 | User | Service Request | KEM-ONE-50107095--LANCEMENT --03.05.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-03 13:59 | 2024-05-03 14:04 | 0 | 0 |  | Salut Yannick, pas de soucis pour moi, tu peux y aller. Best, Nicolas
TCK34611 | User | Service Request | TBH080 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 17:42 | 2024-05-03 09:24 | 0 | 0 |  | Hi Sascha, I'm sorry but it doesn't exist, it'll be either 120 or 200C. Best, Nicolas
TCK34574 | User | Service Request | TED ModBus | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 14:01 | 2024-05-02 17:33 | 1 | 0 |  | Bonjour Cyrille, ci-joint la notice. Best, Nicolas
TCK34573 | User | Service Request | MEX3 with Damping fluid and option 0765 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 13:52 | 2024-05-02 15:30 | 0 | 0 |  | Hi Sascha, this is perfectly normal, on O2 circuits, only the BH5 (fluorocarbon based fluid) is possible to use. O2 reacts with organic fluids like glycerine and silicons and can induce risks of explosions that's why you need a special cleaning, packing and fluids, no organics residus should be able to enter in contact with O2. this is regulated in the EN837. Best, Nicolas
TCK34544 | User | Service Request | FW: MGA ENQUIRY 115772 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Shutt, Oliver [shol] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 11:18 | 2024-05-02 15:27 | 1 | 0 |  | Hi Oliver, Andy seems to have received the same request, here is what I answered him, I attach the datasheet also in this ticket. "Hi Andy, please find enclosed the old datasheet but it was fitted with a diaphragm seal apparently (sort of D821 ?) the E713 can be replaced with a similar Y913 which can also be fitted with a diaphragm seal." Best, Nicolas
TCK34543 | User | Service Request | FW: Request For Quote PRESSURE GAUGE --- Neelkanth Cables Factory Dubai DIC | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bolarinwa, Michael [bomi] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 11:18 | 2024-05-02 16:24 | 1 | 0 |  | Hi Michael, his is an old AAE from Baumer technology India, it was discontinued a couple years back. I don't know your policy about Indian products, but they still have the same in 2.5" or 4". otherwise, we have the MAT3 or 5 if the standard dial fits their needs (cheap option but only standard) otherwise for the dual scale you can offer the MEX3 or 5 or even the DR080 (formerly from haenni). I attach the old AA datasheet with the 80mm option to the ticket for reference. Best, Nicolas
TCK34525 | User | Service Request | LRS10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 10:15 | 2024-05-02 15:20 | 1 | 0 |  | Hi Paul, sure, here you go. Best, Nicolas
TCK34523 | User | Service Request | D043 Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 10:14 | 2024-05-02 16:06 | 1 | 0 |  | Hi Paul, the entry hole diameter is 10mm n 1/2G and 1/2NPT male. Best, Nicolas
TCK34510 | User | Service Request | MS5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 09:19 | 2024-05-02 09:50 | 0 | 0 |  | Bonjour Véronique, Désolé, ce n'est pas modifiable (il y a des trous dans le boitier d'un coté et le boitier et la connexion sont soudés). Le coté de connexion est vraiment important pour son installation ? si c'est crucial, il faudra effectivement remplacer le mano complet. Best, Nicolas
TCK34506 | User | Service Request | stp. Datei | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-02 08:55 | 2024-05-02 09:43 | 0 | 0 |  | Hi Sascha, please find the .stp in the ticket. Best, Nicolas
TCK34484 | User | Service Request | FW: ISO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-05-01 15:59 | 2024-05-02 09:42 | 1 | 0 |  | Hi Paul, here you go, 9001 and 9100. Best, Nicolas
TCK34424 | User | Service Request | PROJET --- 55150516 --- verzi--Messiant | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-30 14:11 | 2024-04-30 17:35 | 1 | 0 |  | Salut Yannick, il faudra bien redéfinir avec le client mais je pense qu'un RP2 pourra peut être convenir, les détails sont a bien reprendre (connexion, setp et pouvoir de coupure notamment). https://www.rockwellautomation.com/en-nl/products/details.836T-T301J.html Best, Nicolas
TCK34423 | User | Service Request | TR: Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-30 14:11 | 2024-04-30 17:30 | 1 | 0 |  | Nathalie, pas de remplacement direct (le Flexbar HRT était un transmetteur sur batterie). mais on peut surement leurs proposer un PFMN (ci-dessous la conversion du code avec le vieux convertisseur excel mais ça donne une base) [Image] Best, Nicolas
TCK34418 | User | Service Request | E713 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-30 14:02 | 2024-04-30 17:24 | 1 | 0 |  | Hi Andy, please find enclosed the old datasheet but it was fitted with a diaphragm seal apparently (sort of D821 ?) the E713 can be replaced with a similar Y913 which can also be fitted with a diaphragm seal. Best, Nicolas
TCK34395 | User | Service Request | TR: inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-30 11:40 | 2024-04-30 17:21 | 1 | 0 |  | Nathalie, oui c'est OK, pas de code existant. Best, Nicolas
TCK34394 | User | Service Request | TR: inquiry ZPN201CHM6/SETP L 40 kPa | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-30 11:37 | 2024-04-30 17:19 | 1 | 0 |  | Salut Nathalie, vu le code c'est surement un remplacement, tu peux proposer celui ci qui a le bon setp: ZPN--201CHM6/SETP EN 11113602 0.15 MPA FALLING RB ZPN--201CHM6/SETPZPN--201CHM6/SETPEN112086510.04 MPA FALLING DEADBAND EQUAL 18 KPARL Best, Nicolas
TCK34338 | User | Service Request | demande tech - separateur DANC .. 29.04.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-29 17:04 | 2024-04-30 09:12 | 1 | 0 |  | Salut Yannick, ce sont exactement les mêmes: ci-dessous une nomenclature de DANC01 et de DANC10 DANC01 [Image] DANC10 [Image] Best, Nicolas
TCK34279 | User | Service Request | ARADOS --- étude technique --- VDAV | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-29 11:22 | 2024-05-07 09:22 | 1 | 0 |  | Yannick, on penche pour un D295 daprès les photos, tu peux envoyer la datasheet au client pour faire quelques mesures et vérifier que ça match et définir la taille. Best, Nicolas
TCK34270 | User | Service Request | FW: RE: DECI RFQ 164.0- SAT - BAUMER BOURDON | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-29 10:08 | 2024-04-29 17:25 | 1 | 0 |  | Hi Andrew, we are dealing directly with SATORP or the replacements, it'll not go through distribution. can you send the pictures and datasheet to Jerome Bredon please ? they have less details than you have. Best, Nicolas
TCK34235 | User | Service Request | ASME-BP--- AD--- 26/04/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-26 14:58 | 2024-04-26 15:02 | 0 | 0 |  | Salut Yannick, il va falloir faire la même procédure de SPR que la dernière fois, mais ça vaudrait le coup de prévoir une réunion toi moi et Jorg pour discuter d'une intégration en gamme si ça se justifie. Best, Nicolas
TCK34209 | User | Service Request | Z811B0K6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-26 12:05 | 2024-04-26 16:34 | 1 | 0 |  | Véronique, c'est possible, voici un croquis du cadran et si commande tu peux notifier l'offre A124-2604-87/J de fantinelli. [Image] Best, Nicolas
TCK34201 | User | Service Request | MEP5 & D438 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-26 10:34 | 2024-04-26 11:48 | 0 | 0 |  | Hi Andrew, when I see that I think it's replacement and the code is existing. I found it in SAP, you can offer: 11147282 Our Reference Long description UOM Quantity 30006371 GA PRESS;0-6BAR,MEP5.161.B20,100MM,2IN GAUGE PRESSURE;PRESSURE RANGE: 0-6 BAR;TYPE: BOURDON TUBEDIAL SIZE: 100 MM CONNECTION SIZE: 2 INCONNECTION TYPE: FF FLANGED CONNECTION LOCATION: BOTTOM CONNECTION MATERIAL: STAINLESS STEEL;FFT:CONNECTION PRESSURE: CL 150ALL STAINLESS STEEL FILLABLES3 ACCORDING EN837TYPE OF MOUNTING: CASE AND BEZEL RING AISI 316L 1.44043 BACK LUGS FIXING CONNECTION: 1/2 IN-NPTOIL FILLING: BH1GLYCERIN/WATERLOW VISCOSITY PRESSURE UNIT: BARGTECHNICAL INSCRIPTION ACCORDING TO SPECIFICATION UPDATE DETAILS ANDPN:ORDER CODE: D438-LD0F.N.60.617A0L.0A011.09TOP HOUSING MATERIAL: HOT ROLLEDNACE COMPLIANT 1.4404 STAINLESS STEEL EN10088-3 (TABLE A4)STAINLESS STEEL CAPILLARY TYPE WITH STAINLESS STEEL PROTECTION ANDWHITEPLASTIC ATEX SHEATHOUTLET POSITION: AXIAL OUTLET 0CAPILLARY LENGTH: 4.5 MCONNECTION ON MEASURING INSTRUMENT: 1/2 NPT FEMALEFOR PRESSURE GAUGE DIAPHRAGM MATERIAL: HASTELLOY C276 2.4819DIAPRAGHM COATING: WITHOUT FLANGE STANDARD: EN1759-1FLANGE FACE FINISH: STANDARD FINISHFLANGE MATERIAL: HOT ROLLED NACE COMPLIANT 1.4404STAINLESS STEEL EN10088-3BOTTOM HOUSING COATING: WITHOUT BOLTING: ASME (UNC) TRACING CIRCUIT:WITHOUT DRAIN HOLE (BOTTOM HOLE): NEEDLE VALVE SW 1/4 IN OD 10 VENTHOLE(TOP HOLE): NEEDLE VALVE SW 1/4 IN OD 10OIL FILLING: LRS9 (-40 TO 400 DEG C) SYLTHERM 800W/SEAL;OEM P/N: MEP5.161.B20/0693_9711_0816_0771_0679_0078-D438-LD04.N.60.617A0L0A011.09/9030 OEM: BAUMER BOURDON-HAENNI S.A.S. MODEL: MEP5161120MODEL: MEP5-161.120/9711MODEL: MEP5;DRAWING: SA-JER-IP501-IPEE-190001POS: SH 2POS: 1REV: 2 Best, Nicolas
TCK34182 | User | Service Request | DEFINITION REMPLACEMENT Z811B0K6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-26 08:30 | 2024-04-26 11:42 | 1 | 0 |  | Salut Véronique, je ne pense pas qu'on soit au clair avec la demande du client, à mon avis le problème vient de la. il faut qu'ils définissent si ils sont en vide ou en pression et les bornes réelles qu'ils recherchent. ensuite je proposerais simplement de proposer l'orientation standard du cadran, ce qui sera mieux pour eux car la convention de lecture sera respectée et je proposerais aussi de faire soit une graduation en bar soit une double bar / %, c'est piégeux de bosser uniquement avec une graduation en % Best, Nicolas
TCK34172 | User | Service Request | Ouverture BAFR // Client : EDF UTO // Cde Clt : 20573194 // MEX5-431.B20/9921 // Position 20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 17:39 | 2024-04-26 10:48 | 1 | 0 |  | Ticket ouvert par erreur Best, Nicolas
TCK34163 | User | Service Request | PLANS 2D SNEF 30213904 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 16:53 | 2024-04-26 14:27 | 1 | 0 |  | Salut Véronique, ci-joint les plans demandés: Best, Nicolas
TCK34133 | User | Service Request | FW: Bourdon-Haenni / Pressure | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 13:32 | 2024-04-26 14:24 | 1 | 0 |  | Hi Paul, here is the magic, this is a pressure switch made for the customer ALKAN. but the reference is only for them, so I thing they need to order the replacement part through them. 10297824 MANOCONTACTEUR ALKAN PLAN 60103-01 Best, Nicolas
TCK34130 | User | Service Request | Question | Service Desk > General | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 12:55 | 2024-04-26 14:30 | 1 | 0 |  | Hi Giorgio, you don't need the option for aluminium pointer as it's standard on the MMD3. but otherwise the code is ok, you'll need to send the logo and validate the dial design at creation, but no technical issue. MMD3-F21.K35/9750_D000_L999 Best, Nicolas
TCK34118 | User | Service Request | Mano MEX5 sur AORP 20564850 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 10:57 | 2024-04-26 15:00 | 1 | 0 |  | Salut Véronique, le montage final sera testé à lhélium complet si tu ajoute l'option 0792 (sur un 6 ou un 0.6bar sans distinction). si besoin d'un certificat, tu peux demander un Q002 mention test hélium (voir avec l'ADV pour comment le saisir) et nos montages standard sont fait au ruban PTFE mais on a déjà fait des montages a la loctite, il faudra être sur de bien le préciser à la création du code. Best, Nicolas
TCK34097 | User | Service Request | DEFINITION REMPLACEMENT Z811B0K6 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-25 08:31 | 2024-04-25 09:14 | 0 | 0 |  | Bonjour Véronique, pas de réel problème, par contre je standardiserais avec un 600mbar (0...100% si ils veulent). on peux aussi leurs proposer de réduire le nombre de connectiques et de tubulures en prenant un MFT monté directement sur manifold 5 voies. Best, Nicolas
TCK34067 | User | Service Request | TR: Revisión especificaciones manómetros + certificados | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-24 15:34 | 2024-04-24 17:37 | 0 | 0 |  | Yannick, as discussed, here are the answer we can give to the customer question: the number of the ISO norm regulating the calibration is the 17025:2005, this number is right. the ANSI 56.8 is about leaks so you could see it on an helium leak test maybe but not on a calibration certificate. the sensor used is is a very precise 300bar sensor for which we have a calibration history. the pressure range of the gauge tested is notified in the field range and you can see it in the table as it's tested for calibration between 0 and 160bar (0...100% of the scale). the pattern is in the table, it's 5 points rising and 1 point (mid-scale) falling. enclosed the calibration certificate of 155.0000 and hereafter it's lifecycle management:[Image] I gave you the traceability with the calibration of the COFRAC Lab. Best, Nicolas
TCK34060 | User | Service Request | laa-zekri -- 24.04.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-24 14:50 | 2024-04-24 15:45 | 0 | 0 |  | Salut Yannick, malheureusement pas grand chose, les manomètres duplex sont tombés en désuétude et la demande est tellement faible que la production a été abandonnée par la plupart des fabricants. par contre on peut generalement les remblacer par: Des différentiels (si l'important était la différence de pression) 2 manomètres séparés qui donneront les valeurs des 2 tubes du duplex par un système électronique avec 2 capteurs et un afficheurs (qui pourrait donner les valeurs individuelles des capteurs et leurs différences). Best, Nicolas
TCK34059 | User | Service Request | TR: Need price for model MPG (PBT Material)-D61-D59 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-24 14:21 | 2024-04-24 15:39 | 0 | 0 |  | Anne, il veut simplement le MPG6 (R) qui a un boitier en PBT PolyButylène Téréphtalate. Best, Nicolas
TCK33982 | User | Service Request | Need complete model code & description | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-24 06:29 | 2024-04-24 11:26 | 1 | 0 |  | Hi Anil, all the informations you need are on the drawings, so you can define pretty easily. sually I could have helped you with the complete codes but this was before SAP so you'll have to define the codes from the drawings, sorry. Best, Nicolas
TCK33966 | User | Service Request | TR: Demande de prixTED6*2067 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-23 15:30 | 2024-04-24 17:43 | 0 | 0 |  | Anne, les branchements sont les mêmes, ce sont juste les câbles qui sont différents, le 0604 est préférable car blindé contre les environnements sujets aux interférences magnétiques (si il y a d'autres appareils dans la zone). je t'envoi les photos par teams. Best, Nicolas
TCK33915 | User | Service Request | TR: MIX7-D61.F22/0751_0678_0760_9941 + D820-10N2214G20 + LRS2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-23 11:41 | 2024-04-23 14:04 | 1 | 0 |  | Bonjour Anne, c'est bon dans les 2 techniquement, on a eu des problèmes de fuite avec les évents arrière sur les MIX7/8, donc l'option est possible sur la datasheet du MIX7 mais pas au configurateur car il y a un remplissage. les options sont donc de proposer sans remplissage ou sans évent. ou alors avec un dashpot + l'évent. Pour le prix, oui ça me semble cher, mais à vérifier avec Nadia. Best, Nicolas
TCK33901 | User | Service Request | WG: Angebot 20547162 / Anfrage : Manomètre MEP vs MIX7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-23 10:37 | 2024-04-23 15:01 | 1 | 0 |  | Hi Sascha, Just talked to Jorg and it seems it was already solved, it would be nice not to do the same work twice. here is the possible complete code for what you ask. MEP3-450. B38/0770_0761_0792_0783 0761 = Housing/ring, bottom bore to prevent condensation 0770 = 0.4 throttle screw 0784 = throttle screw secured with loctite 0792 = Helium leak test 1.10 ˄-7 atm Best, Nicolas
TCK33847 | User | Service Request | TR: inquiry of TED-633.B22R/0771/0829 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-22 16:19 | 2024-04-26 14:34 | 1 | 0 |  | Nathalie, after checking with David and in the configurators, your reference is ok: TED-633.B22R/0771/0829 Best, Nicolas
TCK33799 | User | Service Request | A VOIR SI REALISABLE AXIMA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-22 13:37 | 2024-04-22 14:56 | 0 | 0 |  | Bonjour Veronique, le 1/4VCR mâle est un plan existant chez nous et nous avons différents raccords à souder. par contre pour le moyen, long terme, je ne conseille pas ça, il sera plus intéressant de prendre des raccords a souder vers G1/2 par exemple. ce sera moins cher car plus standard, ils pourront démonter et remplacer les manomètres facilement et si besoin ajouter des accessoires. en soudant les instruments, ils vont payer plus cher des connections hors standard et rendrons le démontage le remplacement etc... très compliqué. mon conseil serait donc de prendre des AEKT a souder vers G1/2 et des manos en G1/2. [Image] Best, Nicolas
TCK33785 | User | Service Request | option code 0829 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-22 12:06 | 2024-04-22 15:19 | 1 | 0 |  | Dear Xu Qiang, this cleaning is meant for gas applications as they can have specific requirements. on many gas application, we're talking about clean gas and ou need to avoid too things: dust particles and water/oil from the calibration. so we offer this option which consists of a deep cleaning and drying of the parts n contact after calibration. Best, Nicolas
TCK33783 | User | Service Request | material of MEX5 window glass | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-22 11:49 | 2024-04-22 15:23 | 1 | 0 |  | Hi Zhang wei, our standard glass is mineral glass and even if the glass itself could probably resist 200C, the window gasket will not. that's if you're talking ambient, if the process température is 200C, it's the maximum but the gauge will resist. here is the temperature resistance of our gauge from the datasheet: [Image] Best, Nicolas
TCK33763 | User | Service Request | ATEX + Customer logo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-22 10:01 | 2024-04-22 15:32 | 1 | 0 |  | Hi Paul, you're right, our logo is mandatory with ATEX and the ATEX marking takes a lot of space. but it seems we already did it a couple time, they need to send us their Logo so we can make a test dial and validate if it's ok with them their logo would be over the pointer. the codes would be: 2 pcs. MEX3-D50.B35/0751_0078_9710 2 pcs. MEX3-D50.B38/0751_0078_9710 Best, Nicolas
TCK33741 | User | Service Request | RE: AQ10032 suppression manchon thermomètres. | Service Desk > General | Low |  |  | Closed | Faure, Arnaud [faa] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2024-04-22 08:43 | 2024-04-23 16:43 | 0 | 0 |  | Information well received, thanks you. I close the ticket. Best, Nicolas
TCK33682 | User | Service Request | TR: 10321430 ... blocked app | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-19 15:37 | 2024-04-19 16:25 | 0 | 0 |  | Nathalie, J'ai confirmé avec Nicolas Fernandez, on ne fait plus les TBHA au dessus de 120C. Le PM devrait faire une note a ce propos. je demande le blocage des codes. Best, Nicolas
TCK33600 | User | Service Request | TR: Serial number | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-18 16:48 | 2024-04-19 10:08 | 1 | 0 |  | Nathalie, c'est un 10262827. Best, Nicolas
TCK33585 | User | Service Request | TR: Frage zu Artikel CTX-3.B.3.B34.0 30214018 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-18 15:29 | 2024-04-19 09:33 | 1 | 0 |  | Nathalie, je ne crois pas, sur la datasheet d'avant (2019/2023) il y avait déjà 316L. Best, Nicolas
TCK33556 | User | Service Request | WEP5 range | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-18 13:33 | 2024-04-18 17:50 | 1 | 0 |  | Yes, you're right, the announced max is 10km on open field (we tested with success up to 7km). in a plant like total reffinery with metal piping all around which is the worst case , i's around 500m to 1km. Best, Nicolas
TCK33543 | User | Service Request | BAIT - Intech - Price for combination | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-18 13:03 | 2024-04-18 17:44 | 1 | 0 |  | Hi Francesco, both constructions are ok, just for the fist one, I would take an AORP standard not an AORPB, because it'll be a cheaper option for the same result. Best, Nicolas
TCK33511 | User | Service Request | AQ10032 suppression manchon thermomètres. | Service Desk > General | Low |  |  | Closed | Faure, Arnaud [faa] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-18 10:21 | 2024-04-22 15:38 | 0 | 0 |  | Thanks a lot for the information, I note it. Best, Nicolas
TCK33433 | User | Service Request | GRCH - CYRIUS - 17.04.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-17 13:45 | 2024-04-17 18:20 | 0 | 0 |  | Salut Yannick, on est ok pour faire ce montage maintenant qu'on a les nouvelles membranes, le code sera: MEX3D21X22 + 1095 + LRS1 (+ option de cadran si besoin à valider avec le client). Best, Nicolas
TCK33429 | User | Service Request | REACH statement | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-17 13:26 | 2024-04-18 10:22 | 1 | 0 |  | Hi Paul, we usually provide Baumer Statement on reach and other material compliances (see enclosed). Best, Nicolas
TCK33320 | User | Service Request | DEMANDE NAVAL 1786066 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-16 17:25 | 2024-04-17 13:46 | 1 | 0 |  | Salut Véronique, désolé mais avec ces références je trouve uniquement des semi conducteurs fairchild et une essence qui s'appellent comme ça. rien de chez nous. Best, Nicolas
TCK33319 | User | Service Request | Pressostat Z0370059 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-16 17:16 | 2024-04-17 11:00 | 1 | 0 |  | Salut Véronique, basé sur leur calcul, un CHM pourrait le faire sans problème et un CDM serait sujet à des problèmes de courant trop faible et de tension trop forte. mais cétait un bon remplacement pour le CGSM, je pense plus qu'il était mal défini avec le client depuis le début. Best, Nicolas
TCK33190 | User | Service Request | D631 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-15 17:34 | 2024-04-15 17:46 | 0 | 0 |  | Hi Paul, yes absolutely, ou can do the mounting below (the standard ranges will be 250 or 400mbar but it's possible: [Image] Best, Nicolas
TCK33186 | User | Service Request | MFE-20109235 -- Mr Combat | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-15 16:51 | 2024-04-15 17:11 | 0 | 0 |  | Bonjour Yannick, le client a déjà commandé ce modèle il y a quelques années mais si c'est pour une nouvelle affaire il faudra créer un nouveau plan spécifique avec Philippe Perrichon. FICHE PRODUIT - MANOMETRE FERROVIAIRE / PRODUCT FILE - RAILWAY PRESSURE GAUGE Internal codeSAP code (MFE)DiameterCircuit numberPressure (bar)Accuracy 7bConnectionMounting studLeft locating pin grooveRight locating pin grooveLighttingPointer beneath dial (H+6)Case colourDial colourLettering colourRoundel colourSpecifics markings Front pointer colourFront pointer black heelRear pointer colourRear pointer black heelTrim colourTrim screw colourInternal drawing nBH parts list drawing nCustomer drawingCustomer part list drawingHomologationRemarksTechnical documentDial screen printingFront guard screen printing ResolutionInternal user's manualCustomer article code Best, Nicolas
TCK33110 | User | Service Request | Enquiry for MEX3, 0 to 250 bar / EXIM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-15 10:15 | 2024-04-15 11:26 | 0 | 0 |  | Hello Bee Tin, here are the 2 dials, I suggest more the Bourdon one: [Image] Best, Nicolas
TCK33106 | User | Service Request | TR: AEM / RFQ_Analogue Pressure Gauge_ Bourdon Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-15 09:54 | 2024-04-15 16:57 | 1 | 0 |  | Hi Nathalie, the MZ7 comes with the 2 Aluminum washers but we don't provide it on the MMX1, MEX2 and MEX3. Best, Nicolas
TCK33036 | User | Service Request | cuenot - MX7 - PATRICK | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-12 15:29 | 2024-04-12 17:02 | 0 | 0 |  | Salut Yannick, oui, chaque soufflet pourra supporter 6 bar dans ta configuration donc ils résisteront a une Pmax statique de 6 bar et une différentielle dans les 2 sens de 6bar. Best, Nicolas
TCK32895 | User | Service Request | DEFINIITION VIEUX PRODUITS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-11 10:33 | 2024-04-11 14:47 | 1 | 0 |  | Salut Véronique, On a regardé avec Philippe, pour le ZPX204SHD, il tient bien à 50bar tout comme son successeur non nucléaire le RPPX204. par contre on a aucune doc spécifique, ça a été des produits spéciaux qui n'ont été codifiés comme ça que sur une courte période a une époque peu informatisée (avant 1996) Le ZPC204 ne devrait pas exister même selon les docs les plus anciennes. Mais le C est normalement pour des fortes surpressions (plus fortes que le X, par exemple le ZPC201 tient 80bar avec une plage de 3.5bar) donc logiquement il doit tenir plus de 50bar (normalement a moins 80bar pour ce type d'organes moteurs, celui du haut sur la photo). Best, Nicolas
TCK32790 | User | Service Request | MMD3 vs. MMD5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-10 14:32 | 2024-04-11 09:23 | 1 | 0 |  | Hi Sascha, the MEX5 can resist -40C dry in ambiant and medium, the datasheet will be changed to reflect this (the gasckets of the MMD5 might not resist, or at least not long term). for the reason of why we don't go higher in pressure with the MMD5 I'm not sure but I'll contact jorg. and the solution is to replace it with a MEX5 (better materials and price lower. Best, Nicolas
TCK32754 | User | Service Request | DEFINIITION VIEUX PRODUITS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-10 10:12 | 2024-04-11 09:17 | 1 | 0 |  | Salut Véronique, on a retrouvé une note nucléaire et pendant quelques années avant 1996, des switch SHD existaient avant dêtre rationalisés en SHM. ll faut donc les remplacer par des switchs nucléaires ZPN204SHM et bien leurs renvoyer la datasheet pour vérification, beaucoup de petits détails ont changés depuis 1996 il faut qu'ils re-valident que les caractéristiques sont OK pour eux. Best, Nicolas
TCK32733 | User | Service Request | FKKW35V4MAFBYAA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-10 08:38 | 2024-04-10 10:53 | 1 | 0 |  | Hi Sascha, yes, the EDD575 was a trade product we had in the catalogue during a couple years, but we don't have anything in the group to replace it, they'll have to adress the request typically to siemens, honeywell or yokogawa. please find enclosed the datasheet, it should help the customer to request a replacement. Best, Nicolas
TCK32656 | User | Service Request | TR: Ersatzteile für Druckmessgerät Baumer Barflex 4 0-10 Bar - 30209234 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-09 14:08 | 2024-04-09 18:21 | 1 | 0 |  | Nathalie, voici les info des pieces a renvoyer pour remplacer ce raccord (a voir si commercialement vous voulez le faire et a combien) code du barflex d'origine: 10270406 Joints coté barflex a changer: Bicone: 10309304 Bage BS: 10303340 Joint torique: 10303357 Adaptateur a garder Raccord à clapet : 10303759 (vérifier G1/4M coté adaptateur) 10303760 (vérifier 1/4NPTM coté adaptateur) le coté adaptateur est celui qui n'est pas protégé par le capot/chaînette Best, Nicolas .
TCK32639 | User | Service Request | I: BAUMER / Cliente : IMA LIFE / Quarto riscontro a richiesta chiarimenti manometro analogico con membrana igienica cod. IMA 6250115 e certificazione ATEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-09 12:27 | 2024-04-09 15:09 | 1 | 0 |  | Hi Erald, we don't have the ATEX certification on the DRCG, what we could offer you is a PFMH, you'll find a display, 4...20mA output and more on a clamp diaphragm connection and dedicated to hygienic environments. Best, Nicolas
TCK32543 | User | Service Request | INFOS VIEUX PRESSOSTATS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-09 10:56 | 2024-04-09 14:54 | 0 | 0 |  | Bonjour Véronique, on remonte jusquà 1981 en catalogues, mais je pense qu'on a un gap entre 81 et 97 en 81 on a pas de pressostat et en 97 ils sont déjà codifiés en RPP. mais sur la datasheet de 97, la conversion des ZP en RP est notifiée mais les switch HD et SHD sont des erreurs pour moi ZPX204HD = RPPXxx3204 = RPPx-8x3.204 ZPC204SHD = ? (le RPPC n'existait qu'en 201). Best, Nicolas
TCK32399 | User | Service Request | Inquiry // Old Inquiry from 2016 20191912 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 16:15 | 2024-04-08 18:18 | 0 | 0 |  | Hi Anja, please codify as below to get the replacement MEX5 but with the back flange of the MAR5: MEX5-X30.X24/0751/0151/0412/(SPEC)/9704 (0...16KG/cm Bourdon) (SPEC) : 1 collerette arrière de MR5 Cadran et option de cadran a valider avec le client. Best, Nicolas
TCK32387 | User | Service Request | TR: Anfrage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 15:16 | 2024-04-08 15:26 | 0 | 0 |  | Salut Nathalie, chez nous ce sera plutôt un code MCX5Fx0N64 (-25...40mbar). si c'est refusé on pourra faire un special mais ça pourrait engendrer un coût de création. connexion a redefinir je ne la vois pas. Best, Nicolas
TCK32380 | User | Service Request | TR: Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 14:51 | 2024-04-08 18:13 | 1 | 0 |  | Nathalie, nous avons la certification UL en dessous de 20bar et tu peux utiliser l'option 9711. ce n'est pas ideal et je remonte l'info au PM. Best, Nicolas
TCK32357 | User | Service Request | 3D drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 13:10 | 2024-04-08 15:17 | 1 | 0 |  | Hi Kerstin, the drawing is enclosed and it's normal for the B instead of the 2, it's the same mounting type (so same drawing) just the material is different. Best, Nicolas
TCK32355 | User | Service Request | WG: MR5-7F33.B72.DADL04.2L2A.01/4710_9750_L280 - 2203368 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 13:03 | 2024-04-08 15:15 | 1 | 0 |  | Hi Kerstin, sorry, the minimum for mounting an MR5 with a DADL4 will be -1...5bar. [Image] Best, Nicolas
TCK32309 | User | Service Request | Inquiry IECEx certificate für compact pressure switch ; Kd. 55149185 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-08 09:28 | 2024-04-08 15:12 | 1 | 0 |  | Hi Anja, there you go, the certificate is enclosed. Best, Nicolas
TCK32172 | User | Service Request | Urgent assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-04 16:42 | 2024-04-05 12:04 | 0 | 0 |  | Hi Paul, it'll be this code: MIX7-F30.X15/2155/9704 (0...30"Hg) Best, Nicolas
TCK32165 | User | Service Request | right product code. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-04 16:15 | 2024-04-05 14:35 | 0 | 0 |  | Hi Sascha, Here is what I would offer: On the temp switch, if they want intrisically safe ATEX (Y) then they can't have the cutting power they need, it'll be max 120mA under 24Vdc (higher cutting power available on standard version or Explosion proof). Order code: RT-AYC.415.226E3/SETP RT old Code / SAP code: for the pressure switch I'm missing information and I don't see what to do with the length... you can have it with a capillary if you want. Order code: RPPY-7C3.204/SETP RPPY old Code / SAP code: Order code: AKPL2F5302 AKPL old Code / SAP code: Best, Nicolas
TCK32163 | User | Service Request | Configuration OK ? - Corlido (50136351) - RFQ AOC343460 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-04 15:46 | 2024-04-04 17:51 | 0 | 0 |  | Hi Rudi, on my side your code is OK to mount, the only thing to change is the bolts, you can't have ISO bolts on ANSI flange you need the ASME ones Also we advise the LRS9, so it would be: MPG6-D60.H23.D438-L100.N.20.216G0L.0A050.01/9711_0693 you should also note that it'll be the MPG6 redesign now. Best, Nicolas
TCK32158 | User | Service Request | 6685140407387 / MP100A1-2GAZ400BARS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-04 15:26 | 2024-04-05 14:48 | 1 | 0 |  | Salut Véronique, On ne trouve pas plus de détails à part qu'il a été mis en obsolescence en 2006. En me basant uniquement sur le code MP100... que tu as donné, je suggère de partir sur un MEX5D30B38, mais c'est basé uniquement sur ce code et on manque cruellement d'info dans ce cas, a bien revoir avec le client. Best, Nicolas
TCK32046 | User | Service Request | TR: Inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-03 16:33 | 2024-04-03 16:44 | 0 | 0 |  | Bonjour Nathalie, les CTX sont UL, ci-joint le certificat de conformité. Best, Nicolas
TCK31969 | User | Service Request | FW: Pressure Transducer / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-03 09:02 | 2024-04-03 12:01 | 0 | 0 |  | Hi Bee Tin, the code is too old to be found. The range is weird, -1 to 5 is possible in bar but really weird in kPa and if that's really needed, you'll need to see with BECH for like a PBMN or a PP20. The assumption for NPT comes from the fact that threads appears and it's mounted with PTFE tape. Please start by offering this but the customer must review it: Y91363B77R Best, Nicolas
TCK31942 | User | Service Request | Changement voyant MFE --- RATP- GRCH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 17:06 | 2024-04-03 15:34 | 1 | 0 |  | Salut Yannick, nous avons un voyant plastique utilisé pour la RATP mais c'est du Methacrylate. le code est 10302581 et il est utilisé dans les 2 codes RATP ci-dessous: 10287838 MFE410104JB1-05 / 6333 73 006 02 RATP 10287840 MFE410104JB2-02 / 6333 73 010 02-RATP il faut proposer ça, si il faut absolument du polycarbonate, il faudra faire une SPR et et du sourcing car il n'existe pas encore. Best, Nicolas
TCK31921 | User | Service Request | Siegwerk --- ADRIEN ROEDERER - | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 15:24 | 2024-04-02 16:32 | 0 | 0 |  | Salut Adrien, ça va poser quelques problèmes pour la définition du produit car je ne sais pas si ils ont démonté juste le flasque supérieur d'un DT2 ou s'ils n'avaient que le flasque sup au départ. pour le ZP ce sera remplacé par un RPPN-7A3.204/SETP et au dessous c'est soit un DT2 complet et ils ont laissé la partie basse en démontant ou un flasque sup de DT2, à revoir. Best, Nicolas
TCK31920 | User | Service Request | Drawing 604827.007 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 15:15 | 2024-04-02 15:57 | 0 | 0 |  | Hi Daniel, I don't find anything with this drawing n, but here is what I would offer, to be checked by the customer. Order code: Best, Nicolas
TCK31917 | User | Service Request | [BHFR Material Data] MPG6-D60.F20+AORP0382+ARIN | Process Instrumentation > Material Data Management | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-04-02 14:31 | 2024-04-02 17:39 | 1 | 0 |  | Hi Daniel, as long as you create only 1 SAP n for the full assembly, it'll be only 1 certificate, but be carefull to add the instructions that it applies to all 3 components so that it's clear in production on each manufacturing orders. if you have a doubt when entering the order in SAP you can ask the internal sales here to check. Best, Nicolas
TCK31907 | User | Service Request | RDN8-A3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 13:10 | 2024-04-02 15:47 | 0 | 0 |  | Hi Daniel, you're perfectly right, depending on the rising or falling (and assuming we set the hysteresis at the minimum which is standard if they just ask for a setp) 3 will be fine with the 234. but as you said if it can see more than 5bar differential it would be better to take a 245 for the first one (it's better for the accuracy of the setp at the middle of the scale and for the durability. Best, Nicolas
TCK31896 | User | Service Request | MANO EDF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 12:07 | 2024-04-02 14:01 | 0 | 0 |  | Salut Véronique, oui c'est un vieux duplex, on aura rien pour le remplacer directement mais on a des différentiels vu que de toute façon c'est pour faire une mesure différentielle. à voir avec eux si ils ont vraiment besoin de lire les deux pressions (et donc d'avoir la statique). on peux aussi le remplacer par un mano simple pour la statique et un différentiel. Best, Nicolas
TCK31872 | User | Service Request | FW: PROFESS - Pressure switch for nuclear power plant | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 10:18 | 2024-04-02 10:34 | 0 | 0 |  | Hi Paul, not problem for a Nuclear pressure switch, ours are K3 qualified according to the RCC-E (contacts CHM or SHM). probably someething like a ZPNY206SHM could work. the only thing is we only have single contacts (SPDT) Best, Nicolas
TCK31869 | User | Service Request | tank level indicator AIDA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-02 10:08 | 2024-04-02 10:25 | 0 | 0 |  | Call from Anja, article in stock. Best, Nicolas
TCK31810 | User | Service Request | option code 2151 for Y91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-04-01 13:56 | 2024-04-02 09:32 | 1 | 0 |  | Hi Xu Qiang, it is available: [Image] Best, Nicolas
TCK31773 | User | Service Request | BAPL ¦ TED-5.3.3.B27.R + 5-point calibration certificate ¦ Zapytanie 8202327899 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-29 09:32 | 2024-03-29 10:16 | 0 | 0 |  | Hi Anne, it'll be the Q005: Q005-3.1 CALIBR.CERTIFICATE (PRESS.TR.) 90001355 Best, Nicolas
TCK31759 | User | Service Request | INFO ANCIEN ARTICLE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-28 15:40 | 2024-03-28 16:45 | 0 | 0 |  | Salut Véronique, je ne suis pas sur que ce soit ça mais ça pourrait donner plus déléments à EDF, j'ai retrouvé une référence très similaire sur une commande EDF (200mm au lieu de 220mm). Jespère que ça pourra t'aider et au moins ça pourrait servir de model pour reproposer en 220mm. [Image] Best, Nicolas
TCK31721 | User | Service Request | Angebot 20568284 4100623.0.0 CWG-BAUMER-001: Anfrage Manometer mit Flanschanschluss & geschraubt | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-28 10:41 | 2024-03-28 11:28 | 0 | 0 |  | Bonjour Anja, c'est possible pour le premier, pour le deuxième ce sera en 60kPa minimum et je proposerais ce séparateur en standard comme on a pas plus de détails: Order code: D435-2N20.212.G2001 old Order / SAP code: : (SS) EN1759-1/ANSI B16-5 : (SS) Best, Nicolas
TCK31667 | User | Service Request | FW: Quote request: Y913-057-H26 / Anjie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-28 03:16 | 2024-03-28 11:09 | 1 | 0 |  | Hi Bee Tin, nothing specific there, the customer only did an introduction leaflet. the 0 was left long ago and is not a digit, so the code would be: Y91357H26R as standard we don't ave FM/CSA anymore (not renewed because of too few requests). then the span adjustment is not accessible as standard (only the 0) options are as follow to access this setting. [Image] Best, Nicolas
TCK31662 | User | Service Request | Special connections | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-27 22:58 | 2024-03-28 09:35 | 1 | 0 |  | Hi Paul, we see a lot of requests recently for this type of connection, but there is something weird in what you show, it can't be both 1/4UNF and 7/16UNF ... I have a table with more details or you client to check if it's what he needs, I would suggest this one: [Image] Best, Nicolas
TCK31656 | User | Service Request | manomètre hydrogène prospect --- ROTH 2. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-27 17:01 | 2024-03-27 17:49 | 0 | 0 |  | Salut Yannick, Ce sera OK pour -40C sur des manomètres sec, on est en process pour changer la datasheet des produits en ce moment. Best, Nicolas
TCK31635 | User | Service Request | REFERENCE MANO MFE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-27 14:34 | 2024-03-27 17:57 | 1 | 0 |  | Salut Adrien, ce serait bien les photos du plan que tu m'avais envoyé, mais ça ne change pas ma réponse, comme c'est trop ancien, même si je te trouve le code des années 80, les caractéristiques ont évolué. il faut redéfinir un MFE mais je peux t'orienter un peu déjà: MFE52_120JB2/ voir l'alim avec le client la taille est a vérifier, le 100mm est ce que j'estime à lil. le chrome à l'avant est beaucoup remplacé par du noir sur les références récentes mais j'ai gardé le chrome dans ce que je propose Best, Nicolas
TCK31566 | User | Service Request | Fichier 3D et certificat ATEX des capteurs de pression Hydrogène | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-27 08:24 | 2024-03-27 09:33 | 0 | 0 |  | Bonjour Jacques, ci-joint les plans des Y92 pour hydrogène HP (jusqu'a 400 bar) et BP (jusqu'a 10bar) et le certificat ATEX. D'habitude j'aurais plutôt joint le certificat de conformité CE mais il mentionne spécifiquement Y91. Pour ce qui est du PBMN, il vaut mieux demander a la suisse mais je pense que ce sera le certificat ATEX standard du PBMN HP, et pour le plan il est peut être à faire. Best, Nicolas
TCK31553 | User | Service Request | material of the MEX5 series pointer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-27 05:18 | 2024-03-27 09:22 | 0 | 0 |  | Hi Zhang Wei, We don't have any option for a stainless steel poiter, it would damage the accuracy with the weight and be extremely expensive. I'm not aware of any competitor doing this either. Sorry. Best, Nicolas
TCK31552 | User | Service Request | BAIT ¦ Marval ¦ DP100 certificates | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 20:32 | 2024-03-27 09:53 | 1 | 0 |  | Hi Erald, here is the certificates BASG orders with their DP100 usually (Q001, Q003 and Q008). for the Ra, I would suggest to have it noted on a Q002 as you mentionned, we don't have specific certificate for this. Best, Nicolas
TCK31550 | User | Service Request | Ambient temp question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 18:48 | 2024-03-27 11:59 | 0 | 0 |  | Hi Paul, please see enclosed a note for the approval of use down to -35C under conditions. Best, Nicolas
TCK31540 | User | Service Request | EQUIVALENCE VIEUX PRODUIT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 15:47 | 2024-03-26 16:40 | 0 | 0 |  | Salut Véronique, j'ai l'impression qu'il manquait quelque chose dans le code de départ, le client devra vérifier que la connexion est OK mais sinon c'est bon. Best, Nicolas
TCK31536 | User | Service Request | tapping for damping screw | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 15:33 | 2024-03-26 16:41 | 0 | 0 |  | Hi Xu Qiang, this option means we'll create the thread inside the connexion to allow the addition of a damping screw but we don't provide the screw itself. Best, Nicolas
TCK31508 | User | Service Request | Smink Group (55244287) - RFQ 25/03/2024 - Old Bourdon Sedeme PG | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 13:30 | 2024-03-26 14:51 | 0 | 0 |  | Hi Rudi, sorry, this application is Ultra High Purity and the instruments are really specific, we used to do this in the past but we stopped completely, we can offer anything on this. Best, Nicolas
TCK31488 | User | Service Request | MANO FERROVIAIRE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 11:08 | 2024-03-26 14:49 | 1 | 0 |  | Salut Adrien, pour le premier c'est un 6343 13 006 (code SNCF) donc un: MFE412304JB1-01 / 6343 13 006 10287850 pour l'autre plan, il n'y a pas de références ni SNCF ni de chez nous que je suis capable de retrouver il va falloir définir un MFE avec le client. Best, Nicolas
TCK31474 | User | Service Request | Datasheet For D042 and D043 / Greg | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 10:19 | 2024-03-26 14:22 | 1 | 0 |  | Hi Bee Tin, please find enclosed the datasheet of the D04x, and it'll be a D043 Best, Nicolas
TCK31460 | User | Service Request | WG: Bitte um Angebot, Druckschalter RP2-NL.346 mit Druckmittler | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-26 09:17 | 2024-03-27 17:05 | 1 | 0 |  | Hi Sascha, it should be this one, just check with the cistomer that it's full stainless without coating on the diaphragm as I can't see it from the pictures. this code existing, I think there is a high probability it's for this customer already. 11155737 RP2-NL.346.DT2-2L20.0.000.01 Best, Nicolas
TCK31428 | User | Service Request | Gauge Gasket | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 17:11 | 2024-03-26 09:13 | 0 | 0 |  | Hi Daniel, for this I prefer that you contact India, a bezel ring gasket is pretty specific to a gauge and it influence the IP, the tightness etc... better to replace it with something that was validated. But BTIN can provide it for sure. Best, Nicolas
TCK31422 | User | Service Request | D904 en DN 50 / 1500 RTJ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 16:40 | 2024-03-26 09:10 | 1 | 0 |  | Bonjour Véronique, pour ce sujet avec Total, je préfère que tu demande à Jérôme. Best, Nicolas
TCK31421 | User | Service Request | Stp. File | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 16:37 | 2024-03-26 09:08 | 1 | 0 |  | Hi Sascha, we have the gauge but the drawing for the D055 isn't available, you can still order it if necessary, please contact sales for it. Best, Nicolas
TCK31418 | User | Service Request | CERTIFICAT ACS - séparateur D435 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 16:12 | 2024-03-26 17:21 | 1 | 0 |  | Salut Adrien, les pièces en 316L ne sont pas dans le scope de l'ACS mais les joints oui, ci-joint pour l'exemple le certificat de conformité a la commande et la commande mentionnant l'ACS des joints PTFE de D435 qu'on a en stock actuellement. Donc pas de problème pour le fournir, il faut juste voir avec les ventes comment le saisir pour qu'il vous soit envoyé avec la commande. Best, Nicolas
TCK31404 | User | Service Request | Demande de définition article / RATP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 14:48 | 2024-03-25 15:03 | 0 | 0 |  | Salut Gregory, c'est un plan trop vieux et je n'ai trouvé aucune reference utilisable. par contre, pour moi ça peut être remplacé directement par notre gamme MFE qui a été développé spécialement pour ce type d'affichage des pressions des conduites de frein, et le design, le diamètre et la connexion seront compatibles. tu peux déjà leurs définir un MFE41, en plus ce sera fait selon des normes a jour. Best, Nicolas
TCK31386 | User | Service Request | EDF / COUPLE DE SERRAGE / JOINT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 12:01 | 2024-03-27 17:25 | 1 | 0 |  | Salut Véronique, Apres avoir fait le point sur la fourniture ci-dessous, les joints ne sont pas définis ni fournis par nous, seul le client pourra retrouver des choses dans le système EDF à partir de ces références. Pour le serrage des M16, ça dépends de beaucoup de facteurs sur site comme l'utilisation de lubrifiant ou le joint utilisé mais pour du tout 316L, on serait en valeur recommandée entre 120 et 200Nm, et attention de bien serrer en croix et progressivement. 11137568 ZPN--201SHM.D820GAL42CKB2001.BRIDE Best, Nicolas
TCK31373 | User | Service Request | Step file | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-25 11:16 | 2024-03-25 14:48 | 1 | 0 |  | Hi Paul, here are the closest 3D I have (RT2N and Y are the same in dimensions). but if it's really necessary, the 3D can be also ordered, you can contacts the sales for that but I think it's not very expensive. Best, Nicolas
TCK31300 | User | Service Request | MONTAGE SUR DT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-22 15:51 | 2024-03-22 16:33 | 0 | 0 |  | Salut véronique, ça ne pose pas de problème. Best, Nicolas
TCK31286 | User | Service Request | code 0202 & 0201 in Pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-22 14:09 | 2024-03-22 14:27 | 0 | 0 |  | Hi Anil, here are the right corresponding codes, if closed it's because the setp was missing. Model codes (Old) Model codes (New) SAP code if valid RPPNEA6206-0202RP2NL655-0201RPPE7E6209-0201RP2EL6540203RP2EL6510201RPPNEA6209PIL2RPPNEA6104-0201 Best, Nicolas
TCK31209 | User | Service Request | Replacement For M1256D30B72 / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-22 03:47 | 2024-03-22 09:09 | 0 | 0 |  | Hi Bee Tin, it's hard to tell as there is obviously something wrong on the part number they gave you, but I think it's the best to offer first. But ask them to check if it's the type they need. Best, Nicolas
TCK31208 | User | Service Request | DP100/213.13B/545/0194_SPEZ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-22 02:46 | 2024-03-22 09:07 | 0 | 0 |  | Hi Bee Tin, here you go, just check the SPEZ option with the customer 11220084 DP100-213.13B.545/0194_SPEZ SPEZ: MEDIUM WETTED PARTS POLISHED RA <=0,4m Best, Nicolas
TCK31164 | User | Service Request | Dual-Use-Declaration for MEX3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dorn, Anja [danj] | Birrien, Nicolas [bin] | BG - Baumer Group | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-21 12:31 | 2024-03-27 12:00 | 1 | 0 |  | Document provided, under approval or instructions from JST. Best, Nicolas
TCK31137 | User | Service Request | Q1060 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-21 10:10 | 2024-03-21 15:47 | 1 | 0 |  | Hi Sascha, Q1060 is the former code for the calibration certificate, it's now as standard replaced by a Q008. Best, Nicolas
TCK31136 | User | Service Request | RAC 7/16 20UNF -1/4 FLARE PP10R- SAP 11245809 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-21 09:59 | 2024-03-21 16:12 | 1 | 0 |  | Hi Daniel, this code is for the component connection of PP10R, please see hereafter pictures and enclosed the drawing with the connection it gives on the final product: [Image] [Image] Best, Nicolas
TCK31131 | User | Service Request | The meaning of Microswitch “H” | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-21 09:37 | 2024-03-21 15:39 | 1 | 0 |  | Hi Zang Wei, I see what you mean, and it is not a mistake but a different choice of language and the new one is better. "falling" or "opening on rise" are the same thing because is not mentionned, the standard understanding is the closing so "falling" means "Closing on fall". Best, Nicolas
TCK31074 | User | Service Request | Advise for pressure switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-20 13:11 | 2024-03-20 15:05 | 0 | 0 |  | Hi Anil, here are the translations of the 2 codes: ZPN209SH = RPPNCA3209 = RPPN-7C3.209 ZDPN234SHH = RDPNWA3234 = RDN-8W3.234 to be validated by the customer, there was some slight changes since 30 years even if I offer you the most similar version. Best, Nicolas
TCK31053 | User | Service Request | HTM - BEZIER-MANO PE - QUANTITE 100 - 20.03.2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-20 10:13 | 2024-03-20 11:27 | 0 | 0 |  | Yannick, voici ce que je proposerais: MEX2-F40.X24/9690_9691-0662 X car ce n'est pas du Bar EN, et on sera sur de faire un BAT. 0662 car c'est sans logo mais le 0661 enlève tous les logos et les références à nos produits y compris sur le conditionnement etc..., le 0662 enlève juste le LOGO sur le produit. l faudra joindre la photo a la demande de création. Best, Nicolas
TCK31050 | User | Service Request | MAR5A30A20MN62 / Christine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-20 10:05 | 2024-03-20 11:12 | 0 | 0 |  | Hi Bee Tin, please codify as below to get the replacement MEX5 but with the back flange of the MAR5: MEX5-X30.A20/0751/0151/0412/(SPEC) (SPEC) : 1 collerette arrière de MR5 Best, Nicolas
TCK31020 | User | Service Request | TR: Online request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-20 08:51 | 2024-03-20 09:06 | 0 | 0 |  | Nathalie, I checked and the only thing those numbers could be is order number but it doesn't match either, sorry. Best, Nicolas
TCK31016 | User | Service Request | RPPNEA6209PIL2 / Vidara | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-20 08:34 | 2024-03-20 09:14 | 0 | 0 |  | Hi Bee Tin, your code match all codes below, and it's all the same except for the SETP. the red ones are closed, only the white one is opens and it has a setp of 0.1bar falling, please check with your customer if ok, if he wants a different SETP the you need to create a new code. _RPPN-7E6.209/1990_9941_SETP_Q1061 11105076 _RPPN-7E6.209/1990_9941_SETP_Q1061 11105077 _RPPN-7E6.209/1990_9941_SETP_Q1061 11105527 _RPPN-7E6.209/SETP/1990/9941/Q1061/PILLA 10297721 RPPN-7E6.209/9941_ED02_1990_Q011_SETP 11105073 Best, Nicolas
TCK30977 | User | Service Request | MEX3B20B59+SE40-01+LRS1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 15:19 | 2024-03-19 18:00 | 1 | 0 |  | Salut Véronique, oui ça aide, c'est une ancienne version qui n'a pas duré très longtemps du D04x, je suis surpris par le nom par contre. il faudrait proposer un MEX3B20B59 + D042-2H02.01 Best, Nicolas
TCK30951 | User | Service Request | Y92 operating manual | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 13:38 | 2024-03-19 17:55 | 1 | 0 |  | Hi Estera, here is the manual for the Y9xx family. Best, Nicolas
TCK30902 | User | Service Request | CTX avec tension de 10 V | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 10:05 | 2024-03-22 16:31 | 0 | 0 |  | Bonjour Jacques, on a fait un test sur ligne et à 10V on a encore de la marge. donc pas de problème en fonctionnement normal, attention par contre ca pourrait avoir un effet dans les extrêmes en température si l'application est en très haute ou très basse température par rapport aux limites il faudra faire une SPR et des essais de validation mais en conditions usuelles (en intérieur ou en extérieur en France par exemple (-10...40C) ça ira. Best, Nicolas
TCK30901 | User | Service Request | RE: CCR2 séparateurs | Service Desk > General | Low |  |  | Closed | Leban, Christophe [cleb] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 10:02 | 2024-03-19 11:30 | 0 | 0 |  | Merci pour l'info.
TCK30893 | User | Service Request | RE: CCR2 séparateurs | Service Desk > General | Low |  |  | Closed | Pichon, Sylvain [psy] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 09:53 | 2024-03-19 11:30 | 0 | 0 |  | Merci pour l'info.
TCK30892 | User | Service Request | 转发: 威 和 张伟 的聊天记录 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 09:53 | 2024-03-19 10:56 | 0 | 0 |  | Hi Zhang Wei, Absolutely, according to EN 837, here is the Pitch dimension of a G 1/2 thread: Pitch (P): 1.814 mm Best, Nicolas
TCK30860 | User | Service Request | MEP5-461.B20/9941 / APS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-19 07:24 | 2024-03-19 09:40 | 1 | 0 |  | Hi Bee Tin, As you know the customer is responsible for the choice of the material with his process fluids. 316L though seems reasonnable, we have options in Monel with a better suitability if necessary. the design of our dials (Aluminium white painted with black letters) is conform and was tested according to the EN837 including in temperature, humidity and UV. If you need parallax correction, you can orient the customer toward the MV calibration gauges series with parallax mirrors. Best, Nicolas
TCK30844 | User | Service Request | INFO RACCORD | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 16:26 | 2024-03-19 09:50 | 1 | 0 |  | Salut Véronique, la longueur d'un 10265059 ARIN-2306 est de 55mm. pour le raccord que tu as trouvé, il ne devrait pas exister, je vais demander la clôture, c'est une erreur. Best, Nicolas
TCK30841 | User | Service Request | mtbf y 91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 15:38 | 2024-03-18 15:58 | 0 | 0 |  | Here it is: Best, Nicolas
TCK30840 | User | Service Request | TR: Ersatztyp für Haenni ZED 514 30207719 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 15:37 | 2024-03-19 12:09 | 1 | 0 |  | Nathalie, je ne trouve rien ni sur internet ni dans les anciens catalogues ou datasheets que j'ai archivé. Le plus proche semble être le ZED503 de chez Haenni (Datasheet ci-jointe). pour le remplacer, avec affichage, points de consignes et sortie x...20mA, je suggère de proposer le TED et le PFMN mais beaucoup de détails restent à revalider avec le client. Best, Nicolas
TCK30827 | User | Service Request | TR: RFQ - Manometer R134A - KD KASTURI | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 14:15 | 2024-03-18 15:52 | 0 | 0 |  | Bonjour, Pour moi, a part le remplissage qui est en BH1 (le BH9 n'existe plus, a voir ce qui peut aller pour eux mais a priori ils ont déjà commandés ce produit dans 4 sales unit donc le BH1 devrait passer). 11252165 MEX5-AX1.X81/9712_9704_D815 Best, Nicolas
TCK30811 | User | Service Request | VALIDATION REFERENCE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 11:24 | 2024-03-18 15:44 | 0 | 0 |  | Salut Yannick, juste à vérifier qu'ils ne veulent pas le réglage à l'horizontale car le seul code que je trouve est avec cette option. 11220178 MX7-C30.N2KA/0217 ce code sera à réouvrir et le MX7-C30.N2KA sera à créer Best, Nicolas
TCK30760 | User | Service Request | Raccors Y92 pour sames | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-18 08:03 | 2024-03-18 09:47 | 0 | 0 |  | Salut Jacques, maintenant qu'il y a une claim je ne pourrais plus faire grand chose mais je voulais t'apporter quelques éléments intéressants pour simplifier les discussions par la suite. la connexion ancienne et la nouvelle sont les mêmes en terme de dimensionnel, seule lépaisseur de la membrane (plus fine) et le type d'usinage (monobloc) ont changés, pas les dimensions. étant donné qu'il doit y avoir assez de place pour faire passer le joint avant la disparition de la soudure a l'avant ne peut pas avoir d'impact et le G1/2 est toujours du G1/2. j'ai bien peur que si il y a un problème ce soit de leur coté, par exemple qu'ils n'aient pas respecté le standard et ne mettent pas de joint a l'avant avec un montage très ajusté dans ce cas la, le problème vient de leur non respect du standard. ci-dessous les gammes d'un ancien et d'un nouveau, tu verra que le code du raccord est le même et le plan n'a pas changé (fiche de modif du plan indique un seul changement depuis 2010 au niveau de la membrane sans influence sur le montage). Old Y92x3Gxxxx (11009802 based) New Y92x3Gxxxx (11209565 based) 11167912 RACCORD G1/2+ MEMBRANE BP RACCORD G1/2+ MEMBRANE 25 m Best, Nicolas
TCK30722 | User | Service Request | TR: Kindly confirm IC price | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-15 16:03 | 2024-03-19 13:58 | 1 | 0 |  | Anne, pour moi ce que veut Anil comme bride c'est celle-ci: 11012774 XBR:FCBM12A8T600 Best, Nicolas
TCK30717 | User | Service Request | TR: Kindly confirm the price , availability & delivery period of food grade oil. Tri--Therm 300- food grade fill fluid | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-15 15:47 | 2024-03-15 16:23 | 0 | 0 |  | Anne, The FDA certificate is applicable only for the filling fluid in case there is a breach in the diaphragm and the filling leaks into the process. and it's stated on the MSDS you sent. the 3A is about the design and apply to the complete mounting as mentionned on the datasheet, please find our 3A certificate enclosed. Best, Nicolas
TCK30567 | User | Service Request | option 9750 - MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-14 13:20 | 2024-03-14 13:56 | 0 | 0 |  | Hi Estera, for the logo, the standard format is JPEG and usually it needs to be more than 4Mo to look good at printing. Best, Nicolas
TCK30529 | User | Service Request | logitrade -0230294893. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-14 10:16 | 2024-03-14 11:29 | 0 | 0 |  | Bonjour Yannick, je vois dans les commentaires de ce code qu'on l'a fermé pour le recréer avec l'option 0074 qui semble t il était manquante. tu peux faire la demande de création pour MEX2-F20.B38/0683_0753_0074_9711_D686 en remplacement du code 10263590. Best, Nicolas
TCK30517 | User | Service Request | TR: CTX O-Ring | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-14 09:27 | 2024-03-14 11:26 | 0 | 0 |  | Nathalie, les joints de raccords ne sont pas fournis avec les CTX. le joint de cellule en Néoprène est bien créé dans le code 11264899. pour fournir le joint de raccord, il va falloir le créer et trouver un fournisseur, je pense que Mario peut regarder si chez Baumer il y a des joints néoprène pour G1/4 DIN, mais chez nous on en a pas a priori. Best, Nicolas
TCK30466 | User | Service Request | Y92 non conformes | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-13 16:22 | 2024-03-14 11:02 | 1 | 0 |  | Bonjour Cyrille, oui le design de la connexion a légèrement changé pour une meilleure soudabilité des membranes. cet épaulement n'assurait pas de fonction pour le client et n'aura pas d'incidence lors du montage. les joints vont dans la gorge a l'avant et sur le dessous de l'HEX en haut. Best, Nicolas
TCK30406 | User | Service Request | Pressure switch enquiry 12/03/24 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fechtig, Kerstin [fke] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-13 10:48 | 2024-03-14 17:37 | 1 | 0 |  | Hello Anja, the first code has the range 224 which is not compatible with the contact, I suggest to offer the 221 which covers it. the second code was old and can be replaced as below: RDY-8C3.221 RDN-8W3.234/SETP I added a setp on the second one as there seems to be instructions for the setting in the description (even if incomplete). Best, Nicolas
TCK30385 | User | Service Request | Enquiry or Barflex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-13 08:50 | 2024-03-14 10:54 | 1 | 0 |  | Hi Bee Tin, the code is not a reference from us, please see the attached picture and send us a picture of either the external or internal label. [Image] Best, Nicolas
TCK30326 | User | Service Request | RE: EQUIVALENCE STROMAG - ROEDERER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-12 15:18 | 2024-03-12 17:36 | 1 | 0 |  | Salut Yannick, c'est un peu surdimensionné par rapport a ce qu'ils ont mais le plus proche va être un RP2N353, avec pour différence la connexion en G1/2M (G1/4F serait spécifique, on pourrait voir pour 100 pièces par an) et létendue du réglage décart (plus large sur le XML) par contre au niveau robustesse on est a un autre niveau. les puissances de coupures sont à vérifier si ils bossent plutôt dans la partie basse de ce que propose le XML. Best, Nicolas
TCK30290 | User | Service Request | TR: BAIT - Precision - Accessory | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-12 11:23 | 2024-03-12 14:05 | 1 | 0 |  | Bonjour Anne, voici le code: N 05017.0005 MUFFE DIN 16283-G1/2-E.ST N05017.0005 10329915 Best, Nicolas
TCK30278 | User | Service Request | EQUIVALENCE STROMAG - ROEDERER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-12 09:58 | 2024-03-12 17:25 | 1 | 0 |  | Answer on 30326
TCK30277 | User | Service Request | TR: Kindly confirm the price , availability & delivery period of food grade oil. Tri--Therm 300- food grade fill fluid | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-12 09:58 | 2024-03-12 14:08 | 1 | 0 |  | Bonjour Anne, la voici. Best, Nicolas
TCK30230 | User | Service Request | Baumer India AQ Differential Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-11 17:38 | 2024-03-12 09:07 | 0 | 0 |  | Hi Daniel, the replacement will probably be an MFT. Best, Nicolas
TCK30228 | User | Service Request | MEX5-B31.B20 CAD File | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-11 17:25 | 2024-03-11 17:29 | 0 | 0 |  | Hi Andrew, sure, here you go. Best, Nicolas
TCK30207 | User | Service Request | TR: MTTF - Werte (Alfalytik - GIG Karasek) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-11 15:13 | 2024-03-11 15:28 | 0 | 0 |  | Nathalie, ci-joint le MTTF du PBSN. Best, Nicolas
TCK30102 | User | Service Request | Séparateur haute température | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-08 16:33 | 2024-03-08 16:43 | 0 | 0 |  | Salut Cyrille, nous n'avons pas d'huile de remplissage qui résistera au dessus de 400C, ils faudra une solution externe, un col d'extension du coté client pour refroidir un peu avant le séparateur par exemple. Best, Nicolas
TCK30097 | User | Service Request | TR: Kindly confirm the price , availability & delivery period of food grade oil. Tri--Therm 300- food grade fill fluid | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-08 16:04 | 2024-03-08 16:41 | 0 | 0 |  | Anne, Nous n'avons pas de tri-term 300, nous ne somme pas fabricants d'huiles. je leurs proposerais ces 2 la avec les FDS en PJ: BH1 BIDON 10L BIDON DE 10 LITRES DE BH1 11006815 LRS1 BIDON 10L BIDON DE 10 LITRES DE LRS1 11099106 Best, Nicolas
TCK30096 | User | Service Request | request for recognition MG5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-08 15:53 | 2024-03-14 10:40 | 1 | 0 |  | Hi Estera, i can't find out the complete code based on those pictures but I can help a bit, the gauge looks like a MG5-1D30.B19 but the seal is not something standard from us, are you sure it's not just a adaptation piece from the customer ? anyway this is what I can do to help, please check and validate with your customer. Best, Nicolas
TCK30075 | User | Service Request | TR: CRM: Customer request: productInfo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-08 11:30 | 2024-03-08 14:48 | 1 | 0 |  | Nathalie, voila la SP de l'EF6. Best, Nicolas
TCK30038 | User | Service Request | MEX3-D20.B18 + DO40.02 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-07 18:56 | 2024-03-08 09:31 | 0 | 0 |  | Hi Daniel, this is old system codes from before SAP, here is what I can clarify: DO40.02 = D040-2H03 = DO42.2H03.01 DO40.02 = D040-2H03 = DO42.2H03.01 D301-13 = DT3 So what you found was correct, and I can just add that on the last one it's a DT3 but I don't have the details: MEX3-D20.B18 + DO42.2H03.01 MEX3-D20.B19 + DO42.2H03.01 MEX3-D20.B74 + DT3 Best, Nicolas
TCK30030 | User | Service Request | TBI with Triclamp | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-07 16:52 | 2024-03-07 17:12 | 0 | 0 |  | Hi Paul, no problem it's the option 9600 for a TBI with DN38 tri clamp, here is an example of SAP code. then for the length below, you select it in the code as for the other connections. 11156312 Best, Nicolas
TCK30010 | User | Service Request | Sample cert. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-07 13:55 | 2024-03-07 15:05 | 0 | 0 |  | Hi Paul, here is an example but I didn't find one with slicon free option, just O2 cleaning but it'll look the same except the option will be 0794 with silicon free cleaning written. [Image] Best, Nicolas
TCK29976 | User | Service Request | Kindly advise for paint application | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-07 11:47 | 2024-03-07 14:03 | 1 | 0 |  | Hi Anil, you should choose LRS7 for the paint application. I enclosed also the list of filling for reference. Best, Nicolas
TCK29870 | User | Service Request | No PTFE requirement | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-06 11:24 | 2024-03-06 16:12 | 1 | 0 |  | Hi Paul, perfect, I knew we would have a solution if we knew more. here is the best suiting device I can offer (nuclear cleanliness and material certificates for the parts in contact with the fluid). MEX3F20B19/0838/Q003 then I would suggest safety options for durability like the restrictor screw (0769 or 0770) And I'll make the list of the materials of the gauge to share it with your customer and ensure no PTFE is used at all. Best, Nicolas
TCK29856 | User | Service Request | Probe Seal XDSE:32445-09 - Oleon (21106342) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-06 10:06 | 2024-03-07 17:09 | 1 | 0 |  | Hi Rudi, We would prefer to offer below option: 150x (1"Gas) please check with your customer if it's ok but doing the specific seal would be extremely expensive for him and probably also long. if the customer is OK for the 150x, then please request the creation of the code and we'll create the drawing for the top piece of the seal in 1" and give a price. Best, Nicolas
TCK29803 | User | Service Request | Assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 17:43 | 2024-03-06 16:48 | 1 | 0 |  | Closed via 29870
TCK29798 | User | Service Request | TR: BOBST Purchase Order ZE5 471980593 SUPPLIER 53707 MAIL orders.ch@baumer.com----------Avis-obsolescence----- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 16:48 | 2024-03-06 09:56 | 1 | 0 |  | Hi Nathalie, I'm not sure I understand the request but I'll try an answer. 1. The type Bis made for a panel mounting from the front , and the collar is wider because it got 3 holes to screw the gauge to the panel. type F is designed for panel mounting from the back with the clamp system. the clamp in the picture is not ours, so better order it with the gauges to be sure you have something that fits together. 2. the values in Newton are a pressure, but we could offer something as cole as possible considering the values we know. a 10bar gauge for example (X22) could cover 1000N (the exact calculation if linear would be 10bar = 1042N) Best, Nicolas
TCK29764 | User | Service Request | FW: PTFE use | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 13:36 | 2024-03-06 10:00 | 1 | 0 |  | Closed, see Ticket 29803.
TCK29749 | User | Service Request | 9911 question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 12:03 | 2024-03-05 12:19 | 0 | 0 |  | Hi Paul, It's possible. here is the markings I'll offer that should fit their request: manufacturers name: Baumer Bourdon Haenni model of the gauge: MEX5D30B20/9911 (exemple) year and lot of production I would prefer order n and position as it's a unique tracking easily accessible from any sales unit: exemple: 2128654/20 customer code: 211EP21704B02 Best, Nicolas
TCK29717 | User | Service Request | [BHFR Material Data] TR: BAIT ¦ SPINAGROUP 71529 - RDO ITEM 1 e 5 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Deschamps, Josette [jau] | 2024-03-05 09:52 | 2024-03-05 15:51 | 1 | 0 |  | Anne, voici les codes: MIX7461B22_0751_0678_0078 MIX7461B15_0751_0678_0078 Best, Nicolas
TCK29716 | User | Service Request | TR: BAIT ¦ INTECH 276738, richiesta di offerta BOURDON (BAUMER) 20/02 0706 20537 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 09:49 | 2024-03-05 14:23 | 1 | 0 |  | Anne, on peut réduire le choix à ceux ci-dessous, a valider avec le client, mais pense à regarder les "codes anciens" dans SAP, ils y étaient à l'identique (le plus probable et le seul ouvert est en vert). RPPNEA6209PIL2 _RPPN-7E6.209/1990_9941_SETP_Q1061 EN 11105076 RPPNEA6209PIL2 _RPPN-7E6.209/1990_9941_SETP_Q1061 EN 11105077 RPPNEA6209PIL2 _RPPN-7E6.209/1990_9941_SETP_Q1061 EN 11105527 RPPNEA6209PIL2 _RPPN-7E6.209/SETP/1990/9941/Q1061/PILLA EN 10297721 RPPNEA6209PIL2 RPPN-7E6.209/9941_ED02_1990_Q011_SETP EN 11105073 Best, Nicolas
TCK29715 | User | Service Request | TR: BAIT ¦ INTECH Rif. 278783, richiesta di offerta BOURDON (BAUMER) 27/02 1246 9583 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 09:49 | 2024-03-05 14:45 | 1 | 0 |  | Bonjour Anne, les items 1, 2 et 3 sont des sondes de températures qui pourront surement être remplacées par du Baumer type TE2 ou TCR6. pour les 4, 5 et 6 mes suggestions en orange: 4. RPPN 7C3 201 "BOURDON" MEDIUM RANGE PRESSURE SWITCHES TYPE ZP, ALL FLUIDS COMPATIBLE WITH THE MEASURING ELEMENT -50 TO +80 DEG.C (STD) -50 TO +200 DEG.C (ZPN-ZPX-ZPC) OPERATING AMBIENT TEMP. -10 TO +55 DEG.C, STORAGE TEMP -50 TO +70 DEG.C, REPRODUCIBILITY +/- (% OF FULL SCALE, READING ACCURACY +/- 5% OF FULL SCALE REF. BOURDON SEDEME FRANCE MODEL SCALE ADJ. DEADBAND RANGE AT TOP AT BOTTOM ZPN201SH -1 -2.5 BAR PartNo :200 -1200 150-1200 QTY:2 5. RPPN 7B3 205 Connexion a clarifier j'ai un 6 dan le code (1/2NPT) et 1/2gas dans le texte (3). PartNo :RPPNBA6205 DISCH. PRESS. SWITCH 3.5/25B, PMS-30B,1/2" GAZ, MALE QTY:1 6. RPPN-7V3.209 Check overpressure ZPX209GSHH -1- 3.5 BAR QTY:2 Best, Nicolas
TCK29712 | User | Service Request | solyro -- DDP 2403JB007 --- francillard | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 09:06 | 2024-03-05 14:09 | 1 | 0 |  | Bonjour Yannick, comme tu le sais on manque un peu de normes en terme d'hydrogene. pour l'ATEX pas de soucis, ci-joint notre certificat. après nous conseillons les options suivantes sur recommandation de Air Liquide pour les applications H2 mais le client est le garant final de son choix. sauf dashpot et ATEX en standard sur le MG5 [Image] Best, Nicolas
TCK29693 | User | Service Request | FW: Baumer differential pressure switch quote./ RDPNAA8256 / Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-05 06:35 | 2024-03-05 15:36 | 1 | 0 |  | Hi Bee Tin, I was not able to find the exact part number because the chrono code is incomplete on the product (A3 is for January of 2003, 2013 and 2023) and I checked all the entries without finding the product. but we have the code and the setting typpically look like a standard setting so I suggest to offer a RDN-8A8.256. Best, Nicolas
TCK29677 | User | Service Request | TR: REI_P02P4720_Bitumen Project_ RFQ for Gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 17:01 | 2024-03-06 15:43 | 0 | 0 |  | Nathalie, le 0...2bar c'est possible, OK avec Nico et ce sera du B17 et pas de surcoût pour la création du cadran. le 0...12bar est spécifique mais possible avec des frais de création de cadran. Best, Nicolas
TCK29655 | User | Service Request | BAIT ¦ Marval ¦ Y91+flange | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 15:16 | 2024-03-04 17:24 | 0 | 0 |  | Hi Erald, a D820 will be cheaper and perfectly fine, the D920 are really designed for smart process transmitters like Yoko or Honeywell etc... but this mounting is not possible below 1bar, I can only suggest the following: Y91-3.3.3.N15.R + D820-10L.42CH.A2001 or: PBMN...N11.... + D435-2L20.4CH.A2001 or PBMN...N11.... + D631-2L20.4CH.A2001 Best, Nicolas
TCK29651 | User | Service Request | PBSN sur séparateur D820 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 14:33 | 2024-03-04 17:11 | 1 | 0 |  | Bonjour Cyrille, nous ce qu'on peut proposer de mieux c'est un Y913 en 0...2.5bar monté sur ton D820 (avec les options habituel, raccord montage séparateur et precision x2). après pour avoir un 2 bar il faudra passer sur du PBMN réglé (avec les mêmes options). Best, Nicolas
TCK29642 | User | Service Request | [BHFR Material Data] Cryogenic Press.Gauge f/Vapor Manifold 0-2.5 bar VS: Forespørgsel: 536483 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | DKAAB - Århus V | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-03-04 13:20 | 2024-03-08 11:59 | 1 | 0 |  | Hello Jens, some are Bourdon gauges and some are from Nuova Fima in Italy. those are simple gauges and can be replaced by the MEX5. I would offer MEX5-A33.K26/0678 MEX5-A33.K24/0678 both are dual scale bar/psi with micrometric adjustable pointer. they ask for silicon filled but the gauges seems dry. also chech the connections, it could be 1/2NPT on the nuova fima (I see PTFE tape). Best, Nicolas
TCK29637 | User | Service Request | MPG6-D30.B19/9711_0816_0192 Fitted to ABR6232Q09 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 12:52 | 2024-03-04 17:06 | 1 | 0 |  | Hi Andrew, the only thing to check is if your connection male (gauge) goes with the female on the ABR to mount it together. in your customer request, the MPG6DX0B19 /9711 /0816 /0192 /0751+ ABR6232Q09 is fine it's the code below: 11270073 MPG6-AX0.B19.ABR6232Q09/0693_9711_0816++ it includes also the TAG plate he seems to need and the option 0751 is removed because standard on the MPG6. otherwise if you want to change the connection to 1/2G, you can but change it also on the ABR. Best, Nicolas
TCK29631 | User | Service Request | Alternative to Parker's flanged manifold - TotalEnergies Petrochem. Feluy (21107894) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 12:05 | 2024-03-04 15:38 | 1 | 0 |  | Hello Rudi, for this I think it's good if you check with Jerome, he might have faced this problematic before and he offers a lot more than me specific manifolds monoflange. can you send him this request with myself in copy please. Best, Nicolas
TCK29606 | User | Service Request | TR: D002E.00549_AVILES CPL__WG: Baumer PT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 10:30 | 2024-03-04 12:05 | 0 | 0 |  | Salut Nathalie, c'est tout bon pour moi, le marquage colle et il demande bien la même protection de membrane PTFE que sur sa dernière commande. il est possible selon les fournisseurs que le revêtement ai une couleur différente mais ça n'impacte pas la fonctionnalité. Best, Nicolas
TCK29598 | User | Service Request | PFSX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 09:57 | 2024-03-04 11:58 | 0 | 0 |  | Bien sur, tu la trouveras en pièce-jointe, attention, le produit n'est plus du tout disponible. Best, Nicolas
TCK29591 | User | Service Request | etude John deeere | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-04 09:32 | 2024-03-05 16:39 | 1 | 0 |  | Salut Yannick, pour ce projet, si tu fais une SPR avec revue complète j'ai peur que ce soit très long et pour les qtés on ne fera pas un mano spécifique en terme de dimensions. par contre on a des choses proches a proposer en MEX3, MMD3 ou même le futur design du MEX3 DIN, donc je proposerais d'abord des dimensionnels avec les datasheets pour validation chez eux et si ça leurs va il ne reste qu'a faire des cadrans spéciaux et ça ce sera possible. Best, Nicolas
TCK29539 | User | Service Request | TR: Kindly confirm IC price | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-01 15:24 | 2024-03-04 09:44 | 1 | 0 |  | Anne, il est faisable, pas de soucis tu peux faire ta demande de prix. Best, Nicolas
TCK29505 | User | Service Request | BAIT ¦ FITAVA ¦ Cooling curve | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-03-01 12:28 | 2024-03-01 14:09 | 0 | 0 |  | Hi Erald, I can share with you the conclusion of the report on the test of the ASIPD2L3 with oil as a fluid. [Image] Best, Nicolas
TCK29406 | User | Service Request | Drawing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-29 15:23 | 2024-02-29 16:20 | 0 | 0 |  | Hi Paul, here you go: [Image] Best, Nicolas
TCK29390 | User | Service Request | Switch questions | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-29 13:33 | 2024-02-29 16:35 | 1 | 0 |  | Hi Paul, for the first 2 questions it's difficult to answer, I'll ask for the datasheet of the supplier of the microswitch to check if a 2kV is mentionned, but is it linked to a norm or something ? for the second question, I don't understand, what do they want us to comply to, about corrosing the materials in contact are mentionned and the IP to state nothing will go inside but I don't recall any norm with C class except the DPE but not applicable. enclosed is the manual where you have the mounting possibilities with a drawings (bottom right) for the cable gland here are the options we have for a 16mm cable: [Image] Best, Nicolas
TCK29380 | User | Service Request | PTFE use | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-29 12:38 | 2024-03-06 11:15 | 0 | 0 |  | see ticket 29803
TCK29373 | User | Service Request | RTNA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-29 11:38 | 2024-02-29 13:56 | 1 | 0 |  | Salut Veronique, oui j'avais vu ce code mais il ne correspond pas non plus au marquage (même si c'est le plus proche). pour une longueur spécifique il faudra passer par l'approbation du PM (raison pour laquelle je propose d'abord du standard). pour le setp sur la commande axima cétait 2C A LA BAISSE ET + 2C A LA HAUSSE Best, Nicolas
TCK29339 | User | Service Request | RoHS and REACH Declaration | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-29 09:04 | 2024-02-29 09:59 | 0 | 0 |  | Hi Sascha, the way to deal with european conformity is changing, before we had one declaration by type of norm with many products. now it is more the other way around, we issue a conformity declaration for one type of product and list all the european norms and regulations it complys to. in attachement the example of the Non Electric Pressure gauges ATEX and non ATEX. Best, Nicolas
TCK29306 | User | Service Request | orano project -- via ROBIVAL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-28 16:51 | 2024-02-28 17:11 | 0 | 0 |  | Salut Yannick, ça va être directement ce code la, il est en RL et tu dois pouvoir demander le prix au commerce. 11179843 Best, Nicolas
TCK29264 | User | Service Request | RTNAG301E04E0-1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-28 11:48 | 2024-02-28 14:56 | 1 | 0 |  | Salut Véronique, On est mal tombé avec celui la car il y a manifestement des erreurs dans le code mais la date n'a pas été noté sur la plaque. il faut proposer un: Order code: RT-NAG.301.E00EJ/SETP RT old Code / SAP code: ensuite, pour les potentielles différences avec ton code, le 4 était peut être une longueur spécifique, à vérifier. le 0 en fin de code est une erreur car on voit clairement le filetage sur la photo et le 0 est utilisé dans les codes pour signifier l'absence. pour le 1 après le code, il y a énormément de chance que ça ait été l'option SETP mais je ne peut pas être catégorique. donc a bien vérifier avec le client et SETPs à définir. Best, Nicolas
TCK29217 | User | Service Request | stp.files | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-27 18:09 | 2024-02-28 09:29 | 0 | 0 |  | Hi Sascha, there is more questions here, are the DR staying in the catalogue (some will, some won't and some will chane name to integrate the Bourdon range (Mxxx)) to create a 3D is some work even if it's not days, with the numbers of requests for 3D and the numbers of codes we can create we can make it a full time job. but, I'm personally in favor of more 3Ds as the request is becoming more and more common and this is a point that is standing out after 2 years of ticketing and I'll bring this point to Jorg and R&D manager here. and at the end nothing is blocked, you can order the creation of a 3D with the order of the products. Best, Nicolas
TCK29213 | User | Service Request | AGFR Thermowell dimension requirements | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-27 17:51 | 2024-03-05 17:03 | 1 | 0 |  | Here are the definitions, be careful on the qty, one of each will be really expensive but if they are refurbishing a project and consolidate it will make the price drop as soon as you have like 2 or 3 of each. obviously as everything was not cristlal clear from client's definition, the definitions are up for approval. Order code: AGFR-X2E0L.9R100.P AGF old Code / SAP code:(50mm) Order code: AGFR-X2X0L.9R100.P AGF old Code / SAP code:(50mm)(3/4BSPT) Best, Nicolas
TCK29212 | User | Service Request | article code glasses 160mm gauges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-27 17:51 | 2024-02-28 09:14 | 0 | 0 |  | Hi Sascha, on a MEX8, here are the glass references. please check with sales for price and conditions. 10308634 WINDOW GLASS 160MM 10308635 VOYANT 160 TRIPLEX Best, Nicolas
TCK29210 | User | Service Request | PREN value | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-27 17:48 | 2024-02-29 10:24 | 0 | 0 |  | Hi Sascha, the Super Duplex is not available on any of our standard so we don't have a PREN value. if it's a special request validated for a project, then we can ask for it to the material supplier but I need the project ID. Best, Nicolas
TCK29087 | User | Service Request | TR: Offer for BSA1158002000 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-26 16:00 | 2024-02-26 16:41 | 0 | 0 |  | Hi Nathalie, Giorgio, Giorgio, my answer below stated that the article needed to be defined, it's still the case. Please come back to the customer to define a complete code and Nathalie will be able to get you a price then. Best, Nicolas my previous answer: Hi Giorgio, I checked internally but this document allows several variants and mountings. we can confirm it's a DRx080 (datasheet in the ticket) but you'll have to define the exact part number with the customer. Best, Nicolas
TCK29084 | User | Service Request | Contactblok - Dow Silicones | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-26 15:53 | 2024-02-28 09:20 | 1 | 0 |  | Hi Rudi, I don't find more information but with what we had, I would guess the 10357001 is the right fit. the names matches, it was sold to DOW also, and I don't find any other usage in SAP (as a component for a gauge for an example) I would say the risk is minimum to request a price for it and ask for a drawing or description at the same time and then we validate before the order. Best, Nicolas
TCK29056 | User | Service Request | RE: Our inquiry / prijsaanvraag 22379784 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-26 12:44 | 2024-02-26 15:26 | 1 | 0 |  | Hi Kees, Rudi, Thanks for the compliment that I answer in the hour even if it's not so true ;) as you said in the email, the pictures will probably not be anough to define everything especially as the marking on the E7 est erased and we don't see the marking on the flange. but we can make a first draft of a replacement codification and with the range confirmed and the marking on the flange we could complete it. I suggest to start with: Y91333HxxR + D821xxL42xxB2001 the 6000 rating would have been psi usually but then it's not possible with the mounting or consistant with a level application (maybe 6000kPa). Best, Nicolas
TCK28966 | User | Service Request | Part no. assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-23 16:13 | 2024-02-26 09:35 | 1 | 0 |  | Hi Paul, I defined what I can from the old codes, but there is still missing information, so I offered the standard on the blanks but the customer needs to check this out. On the ZP, please check the process connection on the ZT 403, please check the capillary, bulb and connection and be sure to get the setp requested to create the code. on the ZT 415, please check the capillary, bulb and connection (and the range is 20...80C instead of 20...90C). 1 pc. ZPN-204-SI RPPN-7A3.204/SETP 1 pc. ZTN-403-SAM RT-NAG.403.230E3/SETP (price on request with setp from customer) 1 pc. ZTN-405-SI RT-NAA.415.230E3/SETP Best, Nicolas
TCK28943 | User | Service Request | TBI with option 0227 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-23 12:06 | 2024-02-23 14:07 | 0 | 0 |  | Hi Estera, it's not impossible but normally this type of mounting is one of the reason why we have the every angle type of mounting that allows orientation of the reading. [Image] if really want the options because a reason prevent the use of the every angle type we will need the qties to define a price and a bit of time to investigate how to implement it in the system and in production. Best, Nicolas
TCK28856 | User | Service Request | TR: 395643000 ∙ SENS.TEMP.MBT3270KTY81-210 084Z7292 Richiesta disponibilità e prezzo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-22 14:06 | 2024-02-22 14:09 | 0 | 0 |  | Voici le code: [Image] Best, Nicolas
TCK28845 | User | Service Request | TR: BAIT ¦ intech Rif. 276306, richiesta di offerta BOURDON (BAUMER) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-22 11:51 | 2024-02-22 14:06 | 1 | 0 |  | Anne, les 2 codes demandés ci-dessous sont OK mais il manque le remplissage et la seule différence que je vois entre les 2 c'est le RAS1 noté dans le premier mais il est décris ensuite comme classe 1, je pense que les 2 appareils sont les mêmes: MEP5-460.B19 + D621-2N6.4213.G64xx MEP5-460.B19 + D621-2N64.213.G64xx Best, Nicolas
TCK28812 | User | Service Request | TR: SAP code | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-22 09:08 | 2024-02-29 11:24 | 0 | 1 |  | Nathalie, pour moi c'est ce code qu'ils veulent, il faut demander la réouverture, je ne vois pas de soucis pour le faire pour l'instant: 11106227 TED-6.3.3.B22.R.D820-01 53230-02/LRS1++ Best, Nicolas
TCK28730 | User | Service Request | [BHFR Material Data] TR: BAIT ¦ Request from Control-systems | Process Instrumentation > Material Data Management | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-02-21 11:34 | 2024-02-21 17:18 | 1 | 0 |  | Anne, pour le marquage d'un QR code ce n'est peut être pas impossible mais ça doit faire partie de la SPR, être validé et approuvé et intégré au process. j'ai joint le certificat 3A mais attention, un certificat 3A ne peut pas être seulement pour un séparateur car la norme demande le respect de certaines contraintes qui incluent l'instrument et le remplissage. Best, Nicolas
TCK28726 | User | Service Request | TSF L and Ls | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-21 11:12 | 2024-02-21 11:36 | 0 | 0 |  | Hi Zhang wei, you can make the requested L=150mm, 3.5m capillary, 0-120degree C with any stem diameter but the Ls will then be defined and will be maximum 120mm with 6mm diameter. but the customer should be glad that it's less than the full length of the stem as the Ls is the part that need to be immerged in the fluid so the smaller the easier. as long as the Ls is in the medium, the temperature gauge will be operational. Best, Nicolas
TCK28707 | User | Service Request | connector quantity of turns | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-21 09:58 | 2024-02-21 11:30 | 0 | 0 |  | Dear Zhang wei, I tested it with this connection on an accessory with a G1/2F and the standard gasket, there is a bit less than 6 turn from the moment you have the first thread just engaged and the moment I enter in contact with the gasket. then it's a question of tightening but I can say roughly about 6 turns. Best, Nicolas
TCK28699 | User | Service Request | TBHI100-264.140.27T/9003_0090_0751 / IPS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-21 08:46 | 2024-02-21 11:19 | 1 | 0 |  | Hi Bee Tin, sorry, the P2F cannot apply conditions on the manually typed length. the minimum is as stated on the datasheet 100mm. but as you have a sliding connection, a 100mm should be fine with the installation, you'll just have 1cm on top of the connection, I suggest to get a standard 100mm, it'll also be a bit less expensive. Best, Nicolas
TCK28696 | User | Service Request | Replacement For ED 517 / Hydraulic Systems | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-21 08:38 | 2024-02-21 10:57 | 1 | 0 |  | Hi Bee Tin, Here is the definition of those 2 products: 1)ED517/314.291/26D, 0-10,000 psi 2)ED517/314.291/27D, 0-15,000 psi Both are: Relative 0.7% accuracy DIN43650 4...20mA Female 1/4NPT With the pressure, I see only the PBMN HP to replace it (as far as I know the sensor products). but the connection process is not a standard. please contact your sensor sales guy to define a replacement. I enclosed te datasheet of the ED517 Best, Nicolas
TCK28668 | User | Service Request | [BHFR Material Data] TR: PFMN-55.BC3R.C1571.49202.1001+ DT3-2N62.0.660.01 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-02-20 15:14 | 2024-02-21 17:26 | 1 | 0 |  | Anne, J'ai changé la connexion entre le PFMN et le DT (33 et L) car c'est une connexion spécifique sans membrane sur le PFMN qui est la seule possible pour ce montage. le remplissage du PFMN passe a 0 car il est fait par le remplissage du montage. l'exactitude de 0.1 n'est pas possible, selon les règles de montage c'est 0.25 et x2 pour le montage séparateur (9010) une configuration est demandée, attention, elle ne doit pas régler une étendue de mesure plus basse que 0...1bar. PFMN-54.BC3R.C1571.33200.1001 + DT3-2L62.0.660.01/9010 Sinon c'est possible: [Image] Best, Nicolas
TCK28621 | User | Service Request | TR: BHFR - BTIN - Zoll - balm---Require MSDS of BIDON DE 5 LITRES DE LRS9. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-20 10:34 | 2024-02-20 10:39 | 0 | 0 |  | Anne, ci-joint au ticket la MSDS. Best, Nicolas
TCK28620 | User | Service Request | procédés de dégraissage O2 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-20 10:33 | 2024-02-20 10:37 | 0 | 0 |  | Salut Yannick, voici un document fait en 2018 par la R&D justement pour partager aux clients. Best, Nicolas
TCK28613 | User | Service Request | SDI 439 - Order code on the device - Lans Nivotherm (21107104) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-20 10:23 | 2024-02-21 11:40 | 0 | 0 |  | Hi Rudi, as replacement parts, we can still offer SDI and I checked with Jorg for the marking and it'll be possible on the dial. now you can define what they exactly need (a picture of the first dial removed and one of the dial that is just behind will help us find the right product code). then we ask the same code to be created with the marking and you can then make your offer. Please send the pictures to my personnal baumer email address and I'll come back with the exact reference. Best, Nicolas
TCK28503 | User | Service Request | TR: Alimentation TED Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-19 12:34 | 2024-02-19 15:06 | 1 | 0 |  | Nathalie, la doc sur le web que montre le client me parait correcte aussi, si vous voulez juste l'alimenter, c'ezst les pin 1 et 2 (lus la masse). Best, Nicolas
TCK28498 | User | Service Request | TR: MJX7-B61.B29 + ABR2233R0N | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-19 12:04 | 2024-02-19 16:02 | 1 | 0 |  | Anne, le montage bride ne pose pas de soucis, et il demande la soudure de la connexion (9030). le montage avec les trous alignés est le standard. MJX7-B61.B29 + ABR2233R0N / 9030 Best, Nicolas
TCK28481 | User | Service Request | 10304104 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-19 10:51 | 2024-02-19 15:55 | 1 | 0 |  | Salut Veronique, je pense que c'est un peu plus compliqué finalement que mâle femelle. A première vue, c'est sensiblement les mêmes connecteurs avec comme différence que les PIN (tout à gauche sur la photo sont en inox et laiton sur le 10304104 (en photo) et sont tout en laiton sur le 10304188. on est à court de 10304188 pour l'instant (les plus utilisés). Best, Nicolas
TCK28480 | User | Service Request | [BHFR Material Data] WG: Offer for BSA1158002000 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2024-02-19 10:38 | 2024-02-21 14:17 | 1 | 0 |  | Hi Giorgio, I checked internally but this document allows several variants and mountings. we can confirm it's a DRx080 (datasheet in the ticket) but you'll have to define the exact part number with the customer. Best, Nicolas
TCK28388 | User | Service Request | PBMN + D920 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-16 15:45 | 2024-02-19 15:24 | 1 | 0 |  | Hello Estera, there was a couple of adjustment to do but you can go ahead with the code below: PBMN-24J11RA11403201000 + D820-10L42CKB6001/9010/0592 Changes: PBMN in same G1/2 EN as the diaphragm seal No need for a D920 which is for process transmitters such as Yokogawa, honeiwell etc... code back in order (the code is read in the datasheet from left to right even if it goes down and up). "L" for corresponding G1/2 "01" LRS1, filling of the mounting to be confirmed (look at the temperature to be sure it fits and confirm with the customer for chemical compatibility in case of a breach 9010 the class of the mounting is 2x the class of the instrument (0.5%) 0592 option for reduce volume to mount transmitters with diaphragm seals. Best, Nicolas
TCK28373 | User | Service Request | ADAPTATION PRODUIT NAVAL GROUP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-16 13:58 | 2024-02-19 15:27 | 1 | 0 |  | Salut Véronique, le montage est limité a 400C a la zone de contact avec le fluide process (membrane du séparateur). il faudra qu'il reduise a 400C en ammont (par un col de refroidissement par exemple ou un simple tube d'extension). Best, Nicolas
TCK28361 | User | Service Request | TR: enquiry pressure sensor TED - 30201389 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-16 12:07 | 2024-02-19 15:36 | 1 | 0 |  | Nathalie, Ce code D82010L12CEB21 n'existe pas et n'existera plus, le code 1 pour la norme était pour l'ISO et n'est plus utilisée. maintenant, le bon code selon ce que je vois dans sa demande sera le D82010L42CEB21. Best, Nicolas
TCK28359 | User | Service Request | Information concerning BARFLEX® - Koninklijke Luchtmacht Vliegbasis Woensdrecht | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-16 12:05 | 2024-02-19 16:31 | 1 | 0 |  | Hi Rudi, here is the code of the device: 10270713 EA5-AP0.B28R if there was an NSN, it would usually appear on the orders as customer part number but I don't find any, and the sales were done to sales units, or usual customer, I don't see anything army related (I don't have access to all sales though). Best, Nicolas
TCK28303 | User | Service Request | TR: BAIT ¦ Intech Rif. 271339, richiesta di offerta BOURDON (BAUMER) 29/01 1527 21303 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-15 16:10 | 2024-02-15 16:30 | 0 | 0 |  | Anne, bien sur mais c'est juste un ancien code: RDPNEA6131-0201 RDN-4E6.131/SETP 10296548 le setp sera à définir ou à retrouver dans une ancienne commande client si tu as accès et qu'il connait le client. Best, Nicolas
TCK28272 | User | Service Request | TR: DOC | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pétré, Sandra [sape] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-15 12:25 | 2024-02-15 14:17 | 0 | 0 |  | Hi Sandra, the certificates I see for that mouting would be a Q003 material certificate for the parts in contact (bottom part of the DT2 and diaphragm), or Q001 certificates of conformity or Q002 hydrostatic test, both related to the complete assembly. Best, Nicolas
TCK28253 | User | Service Request | Option Code 0003 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-15 10:26 | 2024-02-15 10:53 | 0 | 0 |  | Hi Daniel, you seem to be right and I was not aware about that, hereafter what I find in the system and it seems to be correct, the old option 0003 was for the current tag plate option 9941: MEX5D61B12-0003 MEX5-D61.B12/9941 EN 10285147 MEX5D61B15-0003 MEX5-D61.B15/9941 EN 10285154 MEX5D61B16-0003 MEX5-D61.B16/9941 EN 10285168 MEX5D61B18-0003 MEX5-D61.B18/9941 EN 10285179 MEX5D61B19-0003 MEX5-D61.B19/9941 EN 10285199 MEX5D61B20-0003 MEX5-D61.B20/9941 EN 10285232 MEX5D61B22-0003 _MEX5-D61.B22/9941 EN 10285266 MEX5D61B24-0003 MEX5-D61.B24/9941 EN 10285317 MEX5D61B26-0003 MEX5-D61.B26/9941 EN 10285360 MEX5D61B27-0003 MEX5-D61.B27/9941 EN 10285388 MEX5D61B29-0003 MEX5-D61.B29/9941 EN 10285411 MEX5D61B31-0003 MEX5-D61.B31/9941 EN 10285427 MEX5D61B33-0003 MEX5-D61.B33/9941 EN 10285441 MEX5D61B35-0003 MEX5-D61.B35/9941 EN 10285453 MEX5D61B38-0003 MEX5-D61.B38/9941 EN 10285465 MEX5D61B74-0003 MEX5-D61.B74/9941 EN 10285499 MEX5D61B76-0003 MEX5-D61.B76/9941 EN 10285513 MEX5D61B77-0003 MEX5-D61.B77/9941 EN 10285524 MEX5D61B79-0003 MEX5-D61.B79/9941 EN 10285532 I didn't found any type F though but the principle should be the same. Best, Nicolas
TCK28243 | User | Service Request | RP2-EL.644/SETP / GUE/Hyperwave | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-15 09:38 | 2024-02-15 10:14 | 0 | 0 |  | Hi Bee Tin, I don't see any issue, you can go ahead. please keep in mind that the final validation on material compatibility is customer's because only he knows all caracteristics of the fluid he works on. Best, Nicolas
TCK28197 | User | Service Request | à propos du E930 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-14 15:14 | 2024-02-14 17:49 | 1 | 0 |  | Salut Cyrille, On ne fait plus la connexion R2 (DR2 + écrou) mais je ne sais pas si il trouvera ça ailleurs non plus. je n'arrive pas a trouver plus d'informations, il faudra probablement qu'il change sa connexion pour un standard plus récent, je suis désolé. Best, Nicolas
TCK28168 | User | Service Request | 9/16" Autoclave connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-14 12:24 | 2024-02-19 17:20 | 1 | 0 |  | Closed, in discussion with Jörg. Best, Nicolas
TCK28151 | User | Service Request | Serial numbers versus SAP codes - 22902100097377 - 22902100065112 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-14 10:17 | 2024-02-14 15:58 | 1 | 0 |  | Hi Rudi, here is what I found: 22902100097377 - Found nothing 22902100065112 = 11243890 MEX5-D32.B20.DANC01.2L2A.01/9941_ED02 Best, Nicolas
TCK28139 | User | Service Request | VALIDATION CONFIGURATION MA7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-14 09:10 | 2024-02-14 09:26 | 0 | 0 |  | Véronique, pas de problème pour moi, tu peux y aller. Best, Nicolas
TCK28097 | User | Service Request | 6685145209379 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-13 16:43 | 2024-02-20 14:35 | 1 | 0 |  | Bonjour Veronique, je suis désolé mais ni Christophe ni Jörg ni moi ne trouvons quoi que ce soit sur cette référence. SE40 ne veut rien dire ni dans nos références actuelles ni dans les anciennes et dans aucun système SAP. J'aurais voulu pouvoir plus t'aider mais il va falloir redéfinir le besoin. Best, Nicolas
TCK28093 | User | Service Request | MA7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-13 15:58 | 2024-02-13 16:26 | 0 | 0 |  | Salut Véronique, les MA7 sont toujours fabricable, j'ai des codes récents et ouvert, il faudrait vérifier ton code spécifique, peut être as tu une option ou une charactéristique qui ne se fait plus. pour le certificat ce sera un Q008, 90001319 Best, Nicolas
TCK27982 | User | Service Request | BAPL ¦ replacement ¦ 11126274 ¦ LA-A215A-D-1-D01 ¦ Zapytania ofertowe KM/564/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 16:50 | 2024-02-13 09:09 | 1 | 0 |  | Hi Anna, this is a level switch from BTIN (datasheet in the ticket), I don't know where you should order it, do you work directly with India, I let you check the rest with the sales. Best, Nicolas
TCK27981 | User | Service Request | TR: Customer request: productInfo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 16:39 | 2024-02-13 09:17 | 1 | 0 |  | Nathalie, la politique d'étalonnage est la responsabilité du client mais on peut donner des bonnes pratiques. pour un transmetteur électronique de pression, si c'est notre service métrologie, on commencerais par une vérification d'étalonnage tous les ans pendant 2 ou 3 ans et ensuite selon les résultats: si les résultats ne bougent pas, on espace les vérifications tous les 2ans si on voit déjà un frémissement (petit décalage dans le même sens tous les ans) on continue. si il est déjà plus bon la première/deuxième année, réétalonage et vérification des conditions d'usage (n'y a t'il pas beaucoup de vibrations, pulsations, changements de température etc...) Best, Nicolas
TCK27976 | User | Service Request | 3D drawing for DRO 80/313.111/0...6 BAR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 16:18 | 2024-02-13 09:20 | 1 | 0 |  | Hi Kerstin, There is drawing existing for DR as you guessed correctly, I let you contact sales so that they can give you a price to make a 3D. Best, Nicolas
TCK27973 | User | Service Request | Calcul de stress / CEFF | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 16:02 | 2024-02-12 18:06 | 0 | 0 |  | Salut Greg, vois directement avec JM stp. Best, Nicolas
TCK27966 | User | Service Request | ORANO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 15:31 | 2024-02-12 18:13 | 0 | 0 |  | Veronique, il faudrait 100bar pour ce montage ou au moins du 1" et on ne l'accepterait pas en 0.1%, ce sera un 0.25 pour une precision du montage complet a 0.5%. Best, Nicolas
TCK27957 | User | Service Request | Transmission infrarouge Barflex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 14:50 | 2024-02-12 18:09 | 1 | 0 |  | Bonjour Jacques, Nous ne proposons plus cette option car les adaptateurs USB IR e se font presque plus et notre logiciel actuel sous les windows récents ne le supporte plus. la transmission IR etait une des première techno de transmission de données mais elle est complètement obsolète depuis le bluetooth. Best, Nicolas
TCK27917 | User | Service Request | RAS | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 11:35 | 2024-02-12 18:05 | 1 | 0 |  | Salut Veronique, RAS 16 note qu'il y a une protection et que la pression max admissible est 16bar. pour savoir si c'est un RAS1 ou 2 il te faut le tableau, ici c'est un RAS 1 [Image] Best, Nicolas
TCK27911 | User | Service Request | TR: 11129820 / EF6-BOMB.3.0.B22.R.A1.0/C01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 11:16 | 2024-02-13 10:19 | 1 | 0 |  | Hello Nathalie, if we look at the basic functionnalities, on both you have an M12 electrical connection, G1/4 on process side, your output is 4...20mA and your range 2.5bar. so it seems like a good fit for replacing the EF6. ut then you have all the small differences I cannot list here, the dimensions of the PP10R are slightly different han the EF6, and the datasheets can give you all differences with a detailled review (both datasheet and SP enclosed). What I would like to highlight is: on the EF6, the range is 10bar set to 2.5 so the overpressure resistance is higher. the internal O'ring is NBR on your PP10R and was FVMQ on the EF6 Best, Nicolas
TCK27889 | User | Service Request | new Product Code? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-12 09:35 | 2024-02-12 09:53 | 0 | 0 |  | Hi Sascha, this is a CN1 pressure switch from India (datasheet enclosed). you can offer an RP2N for the replacement. Best, Nicolas
TCK27805 | User | Service Request | TR: TED5 / 10300307 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-09 09:23 | 2024-02-09 10:12 | 0 | 0 |  | Bonjour, je ne trouve rien de contradictoire dans nos docs, la datasheet, la notice et nos docs internes donnent bien 1 et 2 pour l'alim (et masse en 8) et les schémas de pins sont les mêmes. mêmes les couleurs standard de nos câbles sont bonnes et vérifiées en prod aussi (blanc brun et tresse). [Image] [Image]
TCK27772 | User | Service Request | TR: RFQ 70224.01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-08 17:57 | 2024-02-09 09:44 | 0 | 0 |  | Hi Paul, the mounting below is possible: MIX7-D60.F20/0678 + AORP0382/SETP + ARIN2606 + D621.2N24.214.G24.01 limiting factor in term of pressure resistance will be the ANSI flange class 150 1". Best, Nicolas
TCK27728 | User | Service Request | EQUIVALENCE Y913GZ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-08 10:57 | 2024-02-08 15:21 | 1 | 0 |  | Salut Véronique, on peut remercier Nathalie et ses sauvegardes des anciennes bases de commandes mais voici le détail de ce produit, la connexion était du G1/4: [Image] Best, Nicolas
TCK27672 | User | Service Request | DOUBT ABOUT TRICLAMP CONECTION | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-07 23:16 | 2024-02-15 10:44 | 0 | 1 |  | Hi Angel, you're right, it will fit together and will not leak, but you loose all the sanytary/hygienic features by creating a retention zone (in red below). this is because even though your drawing is right, it doesn't account for the diaphragm and surface of our diaphragm seal, here it is revised: [Image] Best, Nicolas
TCK27639 | User | Service Request | Technically possible ? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-07 14:43 | 2024-02-07 15:19 | 0 | 0 |  | Hi Rudi, the configuration is OK, just as usual, the setting of the pointer is not really accessible as they would have to empty the gauge, clean and refill for setting the pointer but it's OK. Best, Nicolas
TCK27534 | User | Service Request | 1 13/16" API-6A RTJ Flange Seals | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 17:17 | 2024-02-06 17:39 | 0 | 0 |  | Hi Daniel, the replacement for a CPG1500 from wika would be a PFMN (Baumer sensor) and we can mount it on diaphragm seals. then you need to define your diaphragm seal according to ANI RTJ and specify an X for the norm (it's not ANSI it'll be API). ask me for validation of the fitting in terms of size of the diaphragm to work with the PFMN. once validate, you make the request to sales with the qty. Best, Nicolas
TCK27528 | User | Service Request | Burst Pressure - TotalEnergies spec REQ-RC-INS-406 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 16:39 | 2024-02-06 17:21 | 0 | 0 |  | Hi Rudi, we comply to the REQ so yes we comply with those burst pressure. We do not usually give out the burst pressure of our tubes (also because it's pretty dependant on the conditions). Best, Nicolas
TCK27521 | User | Service Request | RAS1 & RAS2 - Operating principle | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 16:03 | 2024-02-06 16:30 | 0 | 0 |  | Hi Rudi, here is what I can give you: RAS1: we add an internal stop pin over the tube to prevent over-extension. RAS2: we change the tube for the next one (example, we use a 10bar tube on a 6bar gauge) which automatically raise the max pressure allowable but we have a oss in accuracy. the table below is a sum up based on a standard MEX5: [Image] and here is the table of the max pressure for RAS1 and 2 for different gauges : [Image] Best, Nicolas
TCK27517 | User | Service Request | Demande de certificat cofrac | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 15:46 | 2024-02-06 15:56 | 0 | 0 |  | Merci pour le ticket, répondu via teams entre temps. Best, Nicolas
TCK27502 | User | Service Request | Switch part no. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 14:38 | 2024-02-06 16:23 | 0 | 0 |  | Hi Paul, here you have the old and new designation and the code: RDPNGA6211 RDN-8G6.211 10296556 Best, Nicolas
TCK27424 | User | Service Request | TR: MEX3 hi-vis dial | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-06 08:37 | 2024-02-13 09:04 | 1 | 1 |  | Closure f the request confirmed by tereasa, opportunity lost. Best, Nicolas
TCK27396 | User | Service Request | Y913GZ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-05 16:22 | 2024-02-05 17:28 | 0 | 0 |  | Salut Véronique, je n'ai jamais vraiment compris l'utilité de la création des YGZ mais tu as parfaitement raison, ils devaient tous être avec propreté O2 donc Y913xxxxx/0765 Best, Nicolas
TCK27341 | User | Service Request | RP2-EL.644/ Hyperwave | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-05 09:59 | 2024-02-05 14:47 | 1 | 0 |  | Hi Bee Tin, I found it in the records but no setp was required so it was set at the middle of the scale falling with the hysteresys at the minimum (if adjustable deadband). Best, Nicolas
TCK27252 | User | Service Request | MEP5-460.F76/0693_0678_0191_0720_0763_9941 + D820-A9N.2227.G20.09_9030 / Servo | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-02 11:19 | 2024-02-02 16:19 | 1 | 0 |  | Hi Bee Tin, your mounting is possible. MEP5-460.F76/0693_0678_0191_0720_0763_9941 + D820-A9N.2227.G20.09_9030 Concerning the thicker capillary I guess your customer wants an outside diameter of 8mm (if it's about the internal diameter then it's not a capillary anymore it's a tube and the volume would be so big it could not work with a diaphragm seal). types A, and C respectively are 4.5, 6.3 and 6.5mm external diameter [Image] Best, Nicolas
TCK27247 | User | Service Request | Y91 sur séparateur avec refroidisseur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-02 10:56 | 2024-02-02 14:17 | 1 | 0 |  | Hi, I would suggest the mounting below, I also like the ZPX but it's status and mounting limits are not yet fixed in our product range. Y91-3.3.3.B20.R/2164_ASIP-D1L3_D820-10L.42CH.B60_9010/0592 Also your D820 is a DN50, here is the code if you need a DN25: Y91-3.3.3.B20.R/2164_ASIP-D1L3_D820-10L.42CE.B60_9010/0592 Best Nicolas
TCK27243 | User | Service Request | MEX5-D31.B20 pressure gauge mounted to 1 1/2" tri-clamp seal DANC01.2L2B.01 via AKPL1F5305 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-02 10:23 | 2024-02-02 14:09 | 1 | 0 |  | Hi Andrew, we have it but without the capillary (enclosed in the ticket). Best, Nicolas
TCK27233 | User | Service Request | TBI160-264.140.12T/9003_1135 / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-02 09:46 | 2024-02-02 14:06 | 1 | 0 |  | Hi Bee Tin, I'm sorry, 1m is the max length (code 9) Best, Nicolas
TCK27182 | User | Service Request | FW: Pressure gauge with Low/High switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-01 16:08 | 2024-02-12 18:00 | 0 | 1 |  | Hi Daniel, we can offer this solution, we think 6bar is too low and might be out of specs on the end of scales, but we validated with R&D we can produce it in 10bar: MR57D20B22/1547/LRS1 Best, Nicolas
TCK27179 | User | Service Request | MEX5 S3G | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-01 15:46 | 2024-02-01 17:42 | 0 | 0 |  | Salut Véronique, non, désolé, le seul mano avec rétroéclairage est le MFE. Best, Nicolas
TCK27158 | User | Service Request | TED sur séparateur à bride | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-02-01 12:37 | 2024-02-01 15:46 | 1 | 0 |  | Bonjour Cyrille, Voici ce que tu peux proposer: attention, précision du montage = 2x précision de l'instrument. TED-6.3.3.B76.R_D820-10L.22CH.B60.09_9010/0592 +Q005 [Image] Best, Nicolas
TCK27101 | User | Service Request | MA7- CALIBRATED 180º | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 21:17 | 2024-02-01 15:42 | 1 | 0 |  | Hi Angel, what I'm worried about right now is the SF6 marking on the gauge as SF6 is a very dangerous gas and needs special instruments. Wika is big in this and we never tried to developp products for this. Sorry, I'm afraid we can't offer a replacement for this one. Also I would try to warn the customer as I doubt the gauge is protecting anything anymore. Best, Nicolas
TCK27068 | User | Service Request | technical inquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 15:07 | 2024-01-31 16:33 | 0 | 0 |  | Hello Xu Qiang, I changed the filling as you can't have the cleaning and "use no oil" if you use oil in the filling, then you must use an oil free filling (LRS4) and 150C process is the limit but it's OK (list of fillings enclosed) 1) MEX5-430.E24+ D425-2L20.213.G2004/0765 2) MEX5-430.E20+ D425-2L20.213.G2004/0765 for the other questions, BHFR can provide: Calibration certificate: Q008 (list of certificates enclosed) Certificate of inspection of verification standards: Q001 I guess (certificate of compliance) Diagram showing systems of traceability: our ISO9001 certificate assess of the reliability of our quality system (enclosed) Can BHFR provide CSA or cUL for MEX5?: No, but our MEX are CRN registered (I cen send the certificate if you tell me for which provincein Canada I attached the one for Ontario) Whats the diaphragm diameter of D425 and D435?: we usually don't give it as we ensure compatibility and functions of the system. Best, Nicolas
TCK27064 | User | Service Request | CPX with ATEX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 14:17 | 2024-01-31 16:19 | 0 | 0 |  | Hi Estera, sorry, we don't have ATEX on the CPX. Best, Nicolas
TCK27011 | User | Service Request | D820 + Tantalum | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 10:00 | 2024-02-01 17:45 | 0 | 0 |  | Hi Paul, we never did a tantalum diaphragm on a D820, so it would be a special execution. We would first suggest to find another suitable material or coating that is more standard. if the tantalum is absolutely necessary, then we need an SPR with qty and business information to be able to gget you a validation and a price if possible. Best, Nicolas
TCK27005 | User | Service Request | BAPL ¦ product selection pressure gauge with separator | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 09:38 | 2024-02-06 09:07 | 0 | 0 |  | I don't need help anymore
TCK26983 | User | Service Request | Zero Needle at the zero point of MX/MZ diff gauges. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 06:18 | 2024-01-31 11:25 | 1 | 0 |  | Hello Zhang Wei, we don't have pointer stopper at zero to ensure the reading is really zero and the pointer is not just resting on the zero stopper. but for overpressure, there is an internal stop in the movement preventing the pointer to go over the end of the scale (it's actually stopping the pointer a bit after the 270 of the scale. Best, Nicolas
TCK26979 | User | Service Request | option 2153 of Y91 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-31 04:09 | 2024-01-31 11:22 | 1 | 0 |  | Hi Xu-Qiang, it means that we will lock the access to the setting. Best, Nicolas
TCK26971 | User | Service Request | MEX3 for autoclave | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-30 19:55 | 2024-01-31 11:20 | 1 | 0 |  | Hi Paul, With DANC (or some other sanitary diaphragm seals), MEX3 + DANCxxx01/1985 are compatible with SIP and CIP (Sterilization/Cleaning in place) so OK for autoclaves otherwise a dry MEX3 with the right autoclave connection (to define, usually some kind of UNF) could also work but I don't know of any norm compliance then. Best, Nicolas
TCK26833 | User | Service Request | Technically possible? - MEP5-461.K19.D425-2N70.214.G2309/0816_0771 - Ineos Aromatics (21108018) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-29 11:50 | 2024-01-29 14:02 | 1 | 0 |  | Hi Rudi, yes, the mounting is OK. [Image] Best, Nicolas
TCK26821 | User | Service Request | TR: Urgent STP / PBMN-24B17RA15403201000_ASIPD1L3_ DANC02-2L2A_LRS10 // MRI pour vendredi 26/01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-29 10:36 | 2024-01-29 13:58 | 1 | 0 |  | Nathalie, c'est possible mais pourquoi ne pas prendre un PBMH-24B17RA15484222010 qui est diectement en Clamp DN51 et avec un refroidisseur ? Best, Nicolas
TCK26592 | User | Service Request | Current PED Cert. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-25 09:44 | 2024-01-25 14:00 | 0 | 0 |  | Hi Paul, Sure, for non electrical pressure gauges, it's on the global CE certificate, here it is: Best, Nicolas
TCK26557 | User | Service Request | MFT5 40 Bar | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 17:50 | 2024-01-25 14:32 | 0 | 0 |  | Hi Daniel, it is possible but it's not an MFT, it's a dual bourdon tube differential. we purchase it from fantinelli but we can't give the documentation to the customer so I checked with Jorg how to do it. For the offer, we will give the customer this and I'll ask the sales to send you an offer: DOUBLE "C" TUBE DIFFERENTIAL PRESSURE GAUGE Characteristics Full code: XPG: (Drawing number to be defined) Dial nominal size NS: 100mm Process connection: 2 X 1/4"NPTF in line interax 27mm Range: 0/40 bar Accuracy: Class 1,6 Static: 2 x Full scale value Mounting: Surface mounting DP gauge Execution: Silicon oil filled case IP67 IP: IP67 Tubes material: AISI316L Operating T: -20 to 120C ATEX: II 2 G/D Option: Laminated Safety glass Option: Case and ring of AISI316 Dimensional drawing: [Image] For the purchase, you have to order a drawing with those DP gauges and a specific drawing with dimensions and characteristic will be done and linked to the SAP code. Best, Nicolas
TCK26530 | User | Service Request | TR: MPG6 + ABR Flange | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 14:02 | 2024-01-24 14:13 | 0 | 0 |  | Anne, Jérome est dessus en direct avec Satorp, ils peuvent laisser tomber la demande c'est une pièce de rechange. Best, Nicolas
TCK26514 | User | Service Request | "USE NO OIL" option for MEX+D821 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 11:59 | 2024-01-25 15:14 | 1 | 0 |  | Hi Xu Qiang, First of all, you can keep in mind that the cleaning steps on all cleaning options are pretty much the same. undependently of the customer's goal or norm, our role is to get the material clean with no traces of anything on the metal. The signs and markings will differ. to get the sign "Use no oil", please use the O2 cleaning options the EN ask for this mention to be mandatory on all gauges used on O2 applications. so options 0764 or 0765 depending on which one fits your configuration will get you this marking. Best, Nicolas
TCK26511 | User | Service Request | TR: BAES SAP Code Creation // Client : Hydac Technology, S.L. // Cde Clt : 4506949136 // TED-6.B.9.B35.R/0604 // Position 10 | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 11:34 | 2024-01-24 14:01 | 1 | 0 |  | Bonjour Marie Claire, voici un screen de la réponse du PM: [Image] Best, Nicolas
TCK26480 | User | Service Request | SETP Check | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 08:27 | 2024-01-24 09:46 | 0 | 0 |  | Hi Zhang Wei, everything is ok except for one where I need to precise: RPPE3-3A8 101 /SETP -500Pa(-5mbar) on this one, the SETP is right on the limit, which means it's possible rising pressure (as the reset point will then be also in the 10-90% range (-10mbar for example) but it's not possible if the setp you want is with falling pressure because the min reset point is then -2mbar which is out of the 10-90% of the range. Best, Nicolas
TCK26445 | User | Service Request | TBI100-262.140.22T/0751_9003_0225 with dial scale 2 degree as minimum value per scale | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-24 04:25 | 2024-01-24 09:38 | 1 | 0 |  | Hi Bee Tin, Our standard for the TBI is a 5C step which is a good call in regards to the norms, the accuracy of the system and the range. changing the step will not change the accuracy of the device. in case the customer nsist, it would be possible, you need to ask for a special code and there will be a price to create the dial (on time fee) and a cost for printing it special (for each thermometer). you can ask for an offer to the sales if you're sure it is necessary. Best, Nicolas
TCK26404 | User | Service Request | Repeated problems with the RP2 - Berg Maritieme Meetsystemen (50113750) | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-23 13:00 | 2024-01-23 15:56 | 1 | 0 |  | Hi Rudi, I checked with production, R&D and quality quite a lot of elements and I think a good first step would be to have their test procedure (I suspect they are putting pressure too quick to have a right correspondance between their reference and the switch point for example if the reference is a gauge). I checked the former claim results and found that even if there was some drift. In one case the devices were correct with very accurate setp and were found wrong because of a strong deadband (no issue for the customer). For the second the setp was slightly off (36 instead of 40) but it was a second resent so it had been tested and transported a lot. A strong comment I can do is that our return rate on RP2 is extremely low and we don't experience those issues anywhere else so I really think there is something in the testing or application going wrong on customer side but I'm sure we can find a solution for that. also if they feel the need I'm available for a live expertise via teams before we see if other actions are really required. Best, Nicolas
TCK26349 | User | Service Request | Pressure Switch ZPN204SAM / Moore Ind | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-23 07:08 | 2024-01-23 09:34 | 1 | 0 |  | Hi Bee Tin, this is what I would offer, we just don't have information about the process connection but I found one with a matching serial number to your request. 10297743 RPPN-7G3.204 Best, Nicolas
TCK26335 | User | Service Request | Assistance | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-22 16:28 | 2024-01-22 17:31 | 0 | 0 |  | Hi Paul, we miss a couple information like the type of capillary but I would start by offering this: Order code: RT-NAA.403.224C3/SETP RT old Code / SAP code: Type of switch: RTNA standard NA Microswitch: 1 standard changeover switch A Temperature rangeC: Temperature rangeC RTNA mounting with capillary 40 ... 120C 403 Type of design: TD2 stem transmission with stainless steel capillary and stainless steel protection 2 Type of transmission and length: Stainless steel sliding male connection TD2 TD2:stainless steel sliding male connection, length 2 m 2 Stem length: Length 250 mm 4 Bulb diameter: 9.5 mm C Connection: G1/2" 3 OPTIONS Accuracy / Setting: Set point adjustment SETP Best, Nicolas
TCK26318 | User | Service Request | D055-2L14 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-22 14:51 | 2024-01-22 17:28 | 1 | 0 |  | Hi Andrew, we usually don't fit a CPX on diaphragm seal and also, with a comparable volume transmitter, the minimum range would be 100bar. also, if I think we could reach the 0.4 on the diaphragm, it still doesn't make the connection hygienic, baterias will developp on the bottom threads. I would suggest to go for example with the TED flush or hygienic version. or to make an SPR for a CPX with hygienic diaphragm seal like clamp. Best, Nicolas
TCK26258 | User | Service Request | ALF2-4NM Baumer India? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-22 09:50 | 2024-01-22 17:22 | 1 | 0 |  | Hi Daniel, yes it's Indian, and your reference is correct, I would just add a comment on the color of the second scale: MEX5-D60.522 (dual scale black) MEX5-D60.522/9704 (second scale red) Best, Nicolas
TCK26252 | User | Service Request | 3D drawing of MEX5-D91.E20 and TBI100-252.142.12T | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-22 09:15 | 2024-01-22 17:18 | 1 | 0 |  | Hi Xu Qiang, I don't have the exact drawing you asked for but I attached to the ticket the closest I can find. MEX5 type D but in 1/2NPT and G1/2 TBI with aa different stem length. Best, Nicolas
TCK26197 | User | Service Request | BAPL ¦ pressure gauges for hydrogen applications | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-19 16:21 | 2024-01-19 16:41 | 0 | 0 |  | Hi Estera, Hydrogen is a focus and a slightly complicated subject, can you please set a quick 30min call with your colleagues that needs information and I'll give you some training and intels about this ? Best, Nicolas
TCK26195 | User | Service Request | BAPL ¦ Q1229 and Q1070 certificates for MJX7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-19 16:15 | 2024-01-19 16:38 | 0 | 0 |  | Hi Estera, yes, certificates are not treated as option as we don't create the codes with the certificates in the order (it's a separate line). so in many configurator you won't find it (for now) it is available except the names have changed, please refer to the attached documents for offering certificates. Q1229 = Q003 Q1070 = Q009 (standard 5 point calibration certificate Q008 is less expensive and often also accepted. Best, Nicolas
TCK26185 | User | Service Request | Hydrogen / Duratherm | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-19 15:04 | 2024-01-19 15:56 | 0 | 0 |  | Hi Paul, good question, but you'll get a bad answer, we have absolutely no data on the behavior of durathem with hydrogen and I don't find much information. Why do they want duratherm ? don't hesitate to call me so we can find a good solution to offer. Best, Nicolas
TCK26160 | User | Service Request | ZDPN--234SGM/9923 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-19 11:51 | 2024-01-19 15:50 | 1 | 0 |  | Salut Véronique, Oui ce n'est pas super facile vu que les infos sur la datasheet ne sont pas complètement juste. par contre c'est simple, ce ne sont que des soufflets et tous en Inox 316L. Best, Nicolas
TCK26126 | User | Service Request | TAX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-19 08:53 | 2024-01-19 09:11 | 0 | 0 |  | Salut Véronique, ci-joint la documentation du TAX pour faire la conversion. Best, Nicolas
TCK26093 | User | Service Request | Barflex 4 type YA4A - STIB-MIVB | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 16:15 | 2024-01-18 16:20 | 0 | 0 |  | Bonjour Rudi, on a la pièce détachée, mais je doute qu'on autorise le client à la changer vu qu'ils y a beaucoup de soudures à enlever puis à refaire pour changer l'écran. il vaut mieux faire une claim pour voir ce que préfère faire la qualité, mais à mon avis ils proposeront d'en racheter un neuf. Best, Nicolas
TCK26067 | User | Service Request | Z811U030 / 11181574 / MFT7-A00.B18R.62.000/CM01_9923 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 14:05 | 2024-01-18 14:34 | 0 | 0 |  | Salut Véronique, tu vois une différence entre les 2 ? ;) c'est une température max qui vient des tests et calculs pour respecter tous les autres paramètres selon les normes. ce qui veut dire qu'au dessus de cette limite, l'appareil ne va pas automatiquement se détruire, mais ça aura une influence, il résistera un peu moins bien à la Pmax statique car le joint deviendra trop ductile, ou la précision commencera à dépasser le seuil limite etc... donc une surtempérature ponctuelle a 125C ne mettra pas en péril l'appareil si il est utilisé en milieu d'échelle et que tous les autres paramètres sont standard mais si ça se reproduit, ça pourra influencer la durée de vie de l'appareil. Best, Nicolas
TCK26037 | User | Service Request | Y91 For /CON-LASH | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 10:43 | 2024-01-18 14:55 | 0 | 0 |  | Hi Bee Tin, the numbers are not codes or order numbers, I don't find more information thant the fact it's a Y913. you'll have to redefine it with the customer except if he can find other codes or SAP numbers that are really from us (or a picture of the order number on the product. Best, Nicolas
TCK26028 | User | Service Request | RE: Livraison manomètres sans n° de série | Service Desk > General | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 10:25 | 2024-01-18 11:38 | 0 | 0 |  | Closed via phone call and email
TCK26021 | User | Service Request | RE: Livraison manomètres sans n° de série | Service Desk > General | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 10:02 | 2024-01-18 11:38 | 0 | 0 |  | Closed via phone call and email
TCK26013 | User | Service Request | RE: Livraison manomètres sans n° de série | Service Desk > General | Low |  |  | Closed | Pletikos, Maurane [plem] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 09:34 | 2024-01-18 11:38 | 1 | 0 |  | Closed via phone call and email
TCK26010 | User | Service Request | RE: Livraison manomètres sans n° de série | Service Desk > General | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-18 09:24 | 2024-01-18 11:38 | 1 | 0 |  | Closed via phone call and email
TCK26008 | User | Service Request | RE: Livraison manomètres sans n° de série | Service Desk > General | Low |  |  | Closed | Faure, Arnaud [faa] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Cruz Torres, Marcelo [mcr] | 2024-01-18 09:20 | 2024-01-25 15:18 | 0 | 1 |  | Closed, Technical support only in CC. Best, Nicolas
TCK25946 | User | Service Request | TR: PRIX & DELAI - MFE / EXAMESURE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-17 15:05 | 2024-01-17 17:19 | 1 | 0 |  | Nathalie, pour la graduation c'est techniquement possible, il faudra payer une sérigraphie. pour les photos il y en a sur la datasheet, et au besoin je peux en faire d'autres si je sais ce qu'ils veulent voir (pour ca un teams me suffira et je prendrais des photos des appareils les plus proches que je pourrais trouver Attention a ne pas négliger les actions qui seront a faire pour avoir modifier un MFE (nouvelle spec produit, update des bases de codes pour le retrouver, nouveau numéro interne chez le client, approbations etc...) je vous laisse déterminer un cout au besoin. Best, Nicolas
TCK25929 | User | Service Request | MEX3 with D055 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-17 12:50 | 2024-01-17 17:11 | 1 | 0 |  | Hi Sasha, I'm sorry but the minimum range will be 10bar (B22) for the mounting you request. also please make the connexions instrument/seal match: MEX3-D21.B22_D055-2H1601 you can still make an SPR, and I'll quickly see with R&D and make tests to see if we can go lower but we can do it only via an accepted SPR that's worth it (important customer, big quantities, a big order depending on it...). Best, Nicolas
TCK25907 | User | Service Request | FW: REQUEST FOR QUOTATION -TRANSMITTER- SMS/23/1530 - 56039659ML1 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-17 10:47 | 2024-01-17 11:16 | 0 | 0 |  | Hi Bee Tin, I'm missing a lot of data but I would suggest to start with offering this and define the range with the customer Y91333BxxR + D820A4L42DHB2001 (range 01bar minimum) Best, Nicolas
TCK25585 | User | Service Request | TR: BAIT ¦ Barflex - can we still supply? | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-12 15:56 | 2024-01-15 11:47 | 1 | 0 |  | Anne, c'est Ok pour Céline pour produire: 10270545 EA4-AP3.B98R Best, Nicolas
TCK25547 | User | Service Request | RPPN-7A6.208/SETP / Chemicoat | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-12 10:55 | 2024-01-12 11:32 | 0 | 0 |  | Hi Bee Tin, if the customer really have constantly 200bar and want a switch at 7.5bar falling, we have nothing. if he's working at a lower steady pressure and can have accidentaly once in a while 200bar, we have solutions. so we need more information about the application. Also I always reccomend to start by offering an RP2 if possible. Best, Nicolas
TCK25446 | User | Service Request | TR: Is this a product you recognize ? VS: INQUIRY MAE-C-T-10-01-24, URGENT.. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-11 11:04 | 2024-01-11 15:38 | 1 | 0 |  | Anne, ce n'est pas de chez nous, c'est un potentiometre de marque Bours, ils peuvent le trouver avec la ref: 53RAA-R25-A15L [Image] Best, Nicolas
TCK25408 | User | Service Request | RPPN-7A3.601/SETP / Moore Ind | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-11 05:02 | 2024-01-11 09:39 | 1 | 0 |  | Hi Bee Tin, your scale is 30...350bar your usable zone (10...90%) is 62...318bar bottom limit is 62 falling because the reset point will be higher so within the limits. high limit is the reset (higher) so 318 and the min histeresys (difference between set and reset) is 30. so max 288 falling. a falling setp can be set between 62 and 288bar. Best, Nicolas
TCK25391 | User | Service Request | FDA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-10 16:36 | 2024-01-11 14:58 | 0 | 0 |  | Hi Sascha, here is the MSDS of LRS1 with the Norm compatibility. Best, Nicolas
TCK25331 | User | Service Request | FW: Missing plate - RPPN7 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-10 10:59 | 2024-01-10 11:45 | 0 | 0 |  | Hi Paul, those plates are for the wall mounting or with an accessory for the 2" pipe mounting. if the instruments are for direct mounting, it is not necessary to do anything as there is no impact on functionnality (I'll just tell it on the line). if they need to do a wall mounting, then we can resend a plate, you need to open a claim for it. Best, Nicolas
TCK25313 | User | Service Request | ATEX + Additional red pointer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-10 09:12 | 2024-01-10 10:55 | 0 | 0 |  | Hi Paul, you're right, it's not possible. the possible options to offer would be markings on the dial (like a red mark/sector) or if you think it differently, an ATEX contact gauge if they want this overpressure to be relayed electricaly. Best, Nicolas
TCK25215 | User | Service Request | need details of PG from sr. no. of instrument | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-09 10:06 | 2024-01-09 11:25 | 0 | 0 |  | Old order found at BTIN. Best, Nicolas
TCK25202 | User | Service Request | P417 - BAUMER THERMOSTAT K3ad SND | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-09 08:59 | 2024-01-09 11:08 | 1 | 0 |  | Salut Véronique, oui le bulbe doit être complètement dans le puits (même complètement en contact avec le fluide si on veut être précis). la connexion que tu choisis en dernier digit est une connexion coulissante sur capillaire de type TD1. tu est bien limitée à 310mm avec du 10m (210 avec du 8m si on pouvait réduire un peu la longueur de capillaire). donc ça ne rentera pas dans le puits Georgin, si ils passent sur du Baumer il faudra un nouveau puits, désolé. Best, Nicolas
TCK25180 | User | Service Request | BAIT ¦ INTECH : Rif. 265281, richiesta di offerta BOURDON (BAUMER) 06/01 1311 8160 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-08 17:38 | 2024-01-08 18:08 | 0 | 0 |  | Sure, here you go, I filled in the blank with standard (like the contact, as there is no details I selected a single max contact). Order code: M31-2A3.B18G M3 old Code / SAP code: 10271728 Nominal size: 150 mm 1 Functions of order: Maxi (4001) 2 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection, 3 back lugs fixing A Process connection: G1/2" 3 Pressure unit: Bar (EN standard) B Pressure range: 0 ... 2.5 bar 18 Static pressure: 6 bar G Order code: AKPLBF5302 AKPL old Code / SAP code: 10264889 Material: Used for gauges MX,MZ,ME,MA,RD,RPP with seals Stainless steel with stainless steel protection B Connector on instrument side: G1/2" (female) swivel nut F5 Connector on process side: Intermediate Connector G1/2" (male) 3 Capillary lenght: 2 m 02 Order code D621-2L60-212-G2001 D621 Flange material Stainless steel (1.4404) 2 Instrument connection G 1/2 Female L Diaphragm material Hastelloy C 276 (HC) 6 Diaphragm Coating Without 0 Standard Flange ANSI B 16-5 2 PN/ DN Class 150 / NPS 1/2" 12 Flange face Raised face RF 1,6 (Type B1) EN1759-1/ANSI B16-5 G Bottom part material Stainless steel 316L (SS) 2 Bottom housing Coating Without 0 Oil Filling LRS1 (-15+150C) Vaseline 01 Best, Nicolas
TCK25164 | User | Service Request | BH10 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-08 15:54 | 2024-01-08 16:20 | 0 | 0 |  | Anne, si le BH1 ne convient vraiment pas à son application le BH10 est disponible. On doit juste en vendre très rarement (surtout en bidon) et il a été bloqué pour ça. Best, Nicolas
TCK25107 | User | Service Request | AROE --- cap INSTRUM --- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-08 11:00 | 2024-01-08 14:38 | 1 | 0 |  | Bonjour Yannick, tu peux proposer un SAM mais je pense qu'on va être hors sujet avec notre produit (et surtout sur le prix) notre offre va être très surdimensionnée par rapport a leur besoin. on a pas cette typologie de thermostats basiques pour les usages en intérieur. Best, Nicolas
TCK24943 | User | Service Request | EDF UTO Z811UWB0 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-04 15:29 | 2024-01-10 15:00 | 0 | 0 |  | Véronique, la consultation fournisseur est faite et une offre est en train de t'être faite par l'ADV. les options étant très spécifiques, elles sont encore sous XXXX et seront définies avec la database si on a la commande. Best, Nicolas
TCK24865 | User | Service Request | RPPN-7A6.601/SETP/ Moore | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-04 07:16 | 2024-01-04 09:22 | 0 | 0 |  | Hi Bee Tin, the placement of the SETP is OK, 137 or 138 bar is fine, but the accuracy of the SETP is 1% of the ful scale so: 1%*(350-30) = 1%*320 = 3.2 so the accuracy of the setp will be +/-3.2bar. This means that asking for 134.7bar makes no sense as even if we base our setting on 134.7 the real setpoint will be between 131.5 and 137.9bar. please ask for a setp at 135 and 138. Best, Nicolas
TCK24860 | User | Service Request | RDN-8A3.234/SETP | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-03 18:01 | 2024-01-04 09:14 | 0 | 0 |  | Hi Daniel, Sure there will be solutions, ou can take a 224 instead of 234 and the SETP will be OK but DPmax = 1bar and StatPmax = 20bar the is also the solution of the RDN6 162 or 163 but the connection is only G1/4F and the pressure limits will still be slightly lower than a 234 (also it's more expensive). Best, Nicolas
TCK24823 | User | Service Request | accuracy class 1 in model MEX3 with case filling | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-03 13:32 | 2024-01-03 14:28 | 0 | 0 |  | Answered via teams and call. Best, Nicolas
TCK24812 | User | Service Request | ANCIEN PRODUIT MC3 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-03 12:05 | 2024-01-03 14:36 | 1 | 0 |  | Bonjour Véronique, Meilleurs voeux à toi aussi, c'est bien un ancêtre du MCX5, il avait un organe moteur bronze ou laiton, mais l'Inox sera plutôt un bonus, voici ce que tu peux proposer (les cadrans sont peut être à créer) MCX5D30X63 (-160...250mmCE) ou MCX5D30X62 (-100...160mmCE) Best, Nicolas
TCK24745 | User | Service Request | FW: Product selection assist - Pressure | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-02 13:00 | 2024-01-03 11:39 | 1 | 0 |  | Hi Paul, you're right about the fact it might be complicated to replace with Duplex, I know we have trouble finding sources for it. but if it is for numerous parts, we can do a special request. Anyway, for replacing a few parts, I would still quote an alternative either with HC276 full wetted parts (D824), or simply with a D820 full stainless (the 316L is still quite comparable to the duplex in terms of properties) but a better understanding of the process would help to offer the best solution. Duplex is a stainless steel alloy made or corrosion properties similar to 316L, it performs better in stress corrosion with chloride ions, acetic acid and formic acid but 316L is more versatile and performs good stress corrosion with a wide variety of chemicals and with Chlorine for example, Monel ould be a valuable option. In conclusion, we should have a good alternative to offer if they are not bound to Duplex that will give them good performance but if they have their mind set on duplex (or if it's specifically really the best solution on their application), then we might not win the case or we have to make a special request but it can take time. Best, Nicolas
TCK24733 | User | Service Request | Total Warranty | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-02 11:28 | 2024-01-03 15:18 | 1 | 0 |  | Hi Daniel, when are quite used to warranty extension, but it'll have a cost and a line needs to be added to the order. the process is you ask Anne, we need the value of the order and the products details (if there is slight changes, it can be adjusted after). she'll validate a price with Stéphane (usually a few % of the total value) and confirm it to you. Best, Nicolas
TCK24728 | User | Service Request | BAPL ¦ Order: 102844989 ¦ Zapytanie ofertowe Baumer - UP/3/2024 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-02 10:17 | 2024-01-03 11:57 | 1 | 0 |  | Hi Anna, it's a 11239423 TED-5.3.5.B77.R/SETP and the settings are as follow: [Image] Best, Nicolas
TCK24720 | User | Service Request | FW: PE1754 MJWilson Cardiff Office | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2024-01-02 09:42 | 2024-01-04 09:38 | 1 | 0 |  | Hi Daniel, I received the replacement part number from Stiko but Stockach is handling the prices for TS so I sent it to Jorg with the details of this serial number. Best, Nicolas
TCK24648 | User | Service Request | function of screw of top | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-25 09:50 | 2024-01-03 15:55 | 1 | 1 |  | Hi Zhang Wei, yes it is a ventilation system, and it seems blind but it's not, the part in the middle is a solid filter that only let pass the air and not the water to avoid condensation inside. this feature was pdated with the Nuclear authorities when they used our products on cold systems and the temperature shocks created problems of condensation inside because humide air was entering. Best, Nicolas
TCK24621 | User | Service Request | TR: BAPL ¦ replacement for ED701/Z20.B13.055.11 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 14:03 | 2023-12-22 16:50 | 1 | 0 |  | Nathalie, voici le détail: Transmetteur de pression raccord G1/2 affleurant (détails ci-dessous) 4...20mA CE Raccord presse étoupe IP67 (proposer un autre raccord). précision standard 0...4bar R INOX remplissage LRS1. [Image] un Y9233LB19R avec une connexion éléctrique spécifique serait mon choix. la compatibilité de la connexion nouvelle est a faire vérifier au client: [Image] Best, Nicolas
TCK24612 | User | Service Request | TR: REQUEST FOR QUOTATION- 262229 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 12:24 | 2023-12-22 13:56 | 0 | 0 |  | Nathalie, we have close to no information, but this fit the request as a first offer: Order code: RP2-EL.664 RP2 old Code / SAP code: 10296678 Type of switch: RP2E E Type of microswitch: Microswitches RP2E 1 standard changeover switch L Hydraulic connection: 1/2" NPT male 6 Pressure range (bar): Static pressure up to 100 bar, peak 200 bar 0 ... 40 bar 64 Best, Nicolas
TCK24598 | User | Service Request | BAPL ¦ Order: 547407/120 ¦ Kolejne zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 10:03 | 2023-12-22 11:02 | 0 | 0 |  | Hi Anna, here is the code you look for, status is phase out but you should still be able to unlock it when asking sales. 11017663 TED-5.6.3.B22.R/0609_2037 Best, Nicolas
TCK24591 | User | Service Request | Demande équivalence / | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 09:17 | 2023-12-22 16:21 | 0 | 0 |  | Bonjour Greg, ci-dessous le remplacement que je propose: MG5-7D30.B77/4710 + 1041 + LRS1 (Mb HC276 DT2 + Corps 316L 10341536) Best, Nicolas
TCK24586 | User | Service Request | BAIT - Ulix Innovation - combination | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 08:44 | 2023-12-22 09:06 | 0 | 0 |  | Hi Francesco, no, you should use a connection G1/2 or G1/4 EN for our mounting with diaphragm seals. but it should be available on everything we mount (Y9, PBMN, PFMN). Best, Nicolas
TCK24585 | User | Service Request | DPO 160/213.153/315/0776 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-22 08:41 | 2023-12-22 15:33 | 0 | 0 |  | I can offer replacement below: DPO 100/816.153/315/0776 Differences: Case diam. 100mm Flange diam. 150mm Silicon oil potentially slightly different, to confirm Best, Nicolas
TCK24545 | User | Service Request | TR: PRIX & DELAI - MFE / EXAMESURE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-21 14:58 | 2024-01-03 15:56 | 0 | 1 |  | Nathalie, voici les éléments de réponse de Philippe, attention ce sont des CCR: Nicolas, Nous ne faisons pas de MFE51. Voir ma réponse en rouge MFE520003JB3 : quantité =1 CCR2 = 260 + sérigraphie à créer (nexistez pas en non éclairé) MFE520003JB1 : quantité =1 CCR2 = 250 + sérigraphie à créer (nexistez pas en non éclairé) MFE520002JB1 : quantité =3 CCR2 = 250 + sérigraphie à créer (nexistez pas en non éclairé) MFE520003JB1 : quantité : 1 idem MFE520002JB1 : quantité : 2 idem La graduation du cadran des manomètres sera bien de 0.1 en 0.1 bar Non le 16 bar, la graduation est tous les 0.2 bar Le 6 bar, entre 3 et 6 bar tous les 0.1 bar et entre 0 et 3 bar tous les 0.2 bar. Best, Nicolas
TCK24521 | User | Service Request | Question MR5 + D055 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-21 11:38 | 2023-12-21 13:57 | 1 | 0 |  | Hello Kerstin, I'm sorry but we don't have a version with reduced tube for the contact gauges. the solutions are as follow: take the bigger size of D055 (at least 1") change for a transmitter with electronic contacts like PFMN if the "flush" feature is not crucial, offer a D04x or a DT. I'm sorry I don't have better options to offer you. Best, Nicolas
TCK24513 | User | Service Request | ABR Flanges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-21 10:38 | 2023-12-21 11:29 | 0 | 0 |  | Hi Daniel, yes, the datasheet is just old (a new one is coming). you can use this one for now to make the definition. Best, Nicolas
TCK24458 | User | Service Request | FW: Pressure Transmitter E713 / Seaquest Marine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-21 04:19 | 2023-12-21 16:30 | 1 | 0 |  | Hi Bee Tin, I suggest to replace by a : Y91323B38R. the difference is it's a 1/4"Gas connection and a 0...400bar but it's standard so the price will be kept low for a similar / newer product. if necessary for the connection to be 1/8, there is the Y91333B38R + ARIN210L (adaptor G1/2F / G1/8M) if there is qty, you can make a request to create the special connection G1/8 on the Y91, you send me the qty you need and I'll validate with the product manager. Best, Nicolas
TCK24457 | User | Service Request | Pressure Transmitter E713 / Seaquest Marine | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-21 04:17 | 2023-12-21 10:11 | 1 | 0 |  | Double 24458
TCK24397 | User | Service Request | Application For Naphtha / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-20 10:56 | 2023-12-20 11:13 | 0 | 0 |  | Hi Bee Tin, your mounting fit's and will resist the temperature, I would just take a graphite packing for the AMFD. the compatibility of 316 with the media is up to the customer, I don't have compatibility table with Naphta. MEX5-D60.B27/0693_0751_0760_9941 + ASIP20N604 + AMFD02N/1803 Best Nicolas
TCK24366 | User | Service Request | MEX5-D60.B20/0693_0751_0760_9941 + AMFD02N +ASIP206604 / Theerasak | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-20 07:25 | 2023-12-20 09:17 | 0 | 0 |  | Hi Bee Tin, yes it is ok it'll be a bit of weight above the ASIP, but for no vibration and if it's a customer request, no problem. Best, Nicolas
TCK24347 | User | Service Request | BarFlex | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-19 16:48 | 2023-12-19 17:55 | 0 | 0 |  | Salut Cyrille, il y a un problème dans la définition, tu demande des mbar (N) et des psi (1C) avec une étendue de mesure cohérente et l'accord de Franck ça devrait pouvoir se faire. Best, Nicolas
TCK24326 | User | Service Request | section câble PBSN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-19 13:50 | 2023-12-21 15:17 | 0 | 0 |  | Bonjour Jacques, j'attendais une réponse de notre fournisseur pour les dimensions sur plan mais je n'ai pas de réponse. donc j'ai fait les mesures au pieds a coulisse (moins précis car se sont des gaines plastiques déformables) Diamètre gaine externe grise: 5.9 mm Diamètre gaine interne bleu/blanche: 2mm Diamètre embout de tresse: environ 1.4mm photo ci-jointe pour référence. Best, Nicolas
TCK24302 | User | Service Request | FW: Pressure Gauge for Project 4T-2701 Tank Re-Construction,AMP//Pisanu | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-19 11:26 | 2023-12-19 12:05 | 0 | 0 |  | Hi Bee Tin, seems compliant with customer's request and as we don't know the concentration, if it's ok for him it's ok for us and you're right, order the graphite packing /1803 for the AMFD. Best, Nicolas
TCK24283 | User | Service Request | BAIT - Flenco - Pressure switch request | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-19 09:58 | 2023-12-19 11:49 | 0 | 0 |  | Hi Francesco, I would go with a standard RP2Y but it'll be european standard not US (ATEX instead of NEC and CE norms instead of UL). enclosed the certificates for the norms it complies to. Best, Nicolas
TCK24135 | User | Service Request | MPE6-D60.B31 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-18 12:18 | 2023-12-18 17:05 | 1 | 0 |  | Hi Daniel, I don't find much more than you, if it is something, it's a 2" pipe mounting but we wouldn't use this option anymore now, I suggest you offer both options: MPG6-D60.C31 + ARA10D60N + AMFD02N MPG6-D60.C31 + S006.06.215 (11219813) + ARA10D60N + AMFD02N + 11011546 (2" pipe mounting). [Image] Best, Nicolas
TCK24124 | User | Service Request | TR: Zapytanie ofertowe | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-18 11:31 | 2023-12-18 15:21 | 1 | 0 |  | Hi Nathalie, it's weird this question is rising from everywhere. The connection G1/2 with flush diaphragm of the TED (3G): [Image] And the one of the PFMN (G51 / 41) are the same: [Image] Best, Nicolas
TCK24079 | User | Service Request | Price and Code DRC 100/813.133/525/0677/Tri-Clamp DN20/1310/1359/SPEZ | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-18 09:50 | 2023-12-18 15:14 | 1 | 0 |  | Hi Sascha, we will not be able to o this low with gauges, but the exact same PFMH could be an option to replace it. otherwise, it would be a MEX5F30B77/0751 + DANC07-2L2B.01/1985 (+ specific marking option if still required). connection matching to be checked. + certificates as required. Best, Nicolas
TCK24077 | User | Service Request | Price and Productcode : ED701/82042554519/Tri-Clamp..... | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-18 09:40 | 2023-12-18 15:06 | 1 | 0 |  | Hi Sascha, we don't do any of those products anymore but he needs a 3A clamp mounted transmitter with a display so I would go with the PFMH (connection to be validated by the customer). PFMH-54.BC3R.A1571.51202.2001/9059 + documentations as requested. Best, Nicolas
TCK24035 | User | Service Request | Scale / range request / connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-15 16:35 | 2023-12-21 16:33 | 1 | 0 |  | Hi Paul, it is possible, there will be a fee to create the dial but you can request the price for this code: MCX5-EJ0.X06 (0...40g/cm) Best, Nicolas
TCK24027 | User | Service Request | Spare Parts | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-15 15:40 | 2023-12-18 14:47 | 1 | 0 |  | Hi Daniel, a list of tools is still available for SUs to be able to do minimum opertions on gauges (enclosed) and you can ask for updated prices to sales administration. Concerning the tools, I suggest to take only the pointer kit and supports (we can D print it) and then order online a belt wrench (example below): [Image] Concerning the spare parts, you can always ask for it but there might be minimum quantities. and it's difficult to be sure what you'll need. a few triplex lenses in 100 and 150mm and some top plugs are a must have but after that it's complicated. for your MX, you the standard window SAP n is 10308758. Best, Nicolas
TCK24020 | User | Service Request | BAPL ¦ N CDE/N: 1192383/10 ¦ odp ¦ Customer request: product | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-15 14:15 | 2023-12-15 14:26 | 0 | 0 |  | Hi Anna, it's probably the first order number and here is the content but you have to contact sales now for the lead time. [Image] Best, Nicolas
TCK24011 | User | Service Request | Check feasebility of configuration - Engie Electrabel (50148724) - RFQ 6000606653 Pos. 10 - M31-4A3.B12J/9712_9704_D000 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-15 12:03 | 2023-12-20 15:34 | 0 | 0 |  | Answered via email
TCK23940 | User | Service Request | Q008, Q009 and Q010 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | SGSIN - Singapore | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-15 05:58 | 2023-12-15 09:32 | 0 | 0 |  | Hi Bee Tin, those certificates are not for the same accuracy, you'll need a Q010, please find the details on the list enclosed. Best, Nicolas
TCK23916 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5492 Pos.20 - MCX5-D60.N09.AORPB501.ARIN2606.D435-2N20.212.G2001/0751_00789941_SETP_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-14 16:16 | 2023-12-14 17:18 | 0 | 0 |  | Ok for me: MCX5-D60.N09.AORPB501.ARIN2606.D435-2N20.212.G2001/0751_00789941_SETP_9030 Best, Nicolas
TCK23915 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5492 Pos.120 - MEX5-D63.B24.D621-2N20.214.G2001/0751_0078_9941_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-14 16:16 | 2023-12-14 17:17 | 0 | 0 |  | Ok for me: MEX5-D63.B24.D621-2N20.214.G2001/0751_0078_9941_9030 Best, Nicolas
TCK23914 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5492 Pos.170 - MEX5-D60.B79.D621-2N20.212.G2001/0751_0078_9941_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-14 16:16 | 2023-12-14 17:16 | 0 | 0 |  | Ok for me: MEX5-D60.B79.D621-2N20.212.G2001/0751_0078_9941_9030 Best, Nicolas
TCK23899 | User | Service Request | Special scale M31 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-14 13:47 | 2023-12-14 17:14 | 1 | 0 |  | Salut Jurgen, Frankignoul c'est les pieux franki mais c'est tout ce que je connais de lui, je ne trouve pas plus d'info sur le produit ou le fabricant. le M31 me parait adapté. Best, Nicolas
TCK23854 | User | Service Request | MIX7A65B22 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-14 08:57 | 2023-12-14 10:17 | 0 | 0 |  | Salut Veronique, ce ne sont pas des informations que nous donnons. le tube est le coeur de notre fabrication, nous ne pouvons pas donner les plans. Best, Nicolas
TCK23823 | User | Service Request | GE HYDRO PROJECT BRAZIL -- | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-13 16:20 | 2023-12-14 10:16 | 1 | 0 |  | Oui ça me parait bien, après je ne vois pas l'étendue de mesure mais pour le reste pas de soucis. Best, Nicolas
TCK23799 | User | Service Request | R: BAIT ¦ Promeco ¦ High temperature flanges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-13 12:02 | 2023-12-13 12:08 | 0 | 0 |  | Hi Erald, here is the answer I sent, tell me if you need more details: if we're talking diaphragm seals I don't have solutions for you sorry, the filling fluid will be the limitation, over 400C, you'll loose accuracy and it might reduce the longevity of the mounting. but if we talk about a simple flange and you choose the right gasket for the mounting (stainless or graphite) it should be OK. Best, Nicolas
TCK23772 | User | Service Request | Y91 sortie câble | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-13 09:40 | 2023-12-13 11:52 | 1 | 0 |  | Bonjour Cyrille, oui on a arrêté de faire des longueurs de câble spécifiques et le standard est 1.5m. il n'est pas impossible de faire du 3m mais il faut créer un article spécial donc assurer un certain nombre de pièces et demander la validation a Franck. Best, Nicolas
TCK23739 | User | Service Request | PBMN sur séparateur | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-12 16:40 | 2023-12-12 18:08 | 0 | 0 |  | Salut Cyrille, ce montage est possible avec comme bémol que la précision sera 0.5% au lieu de 0.25% (une option est ajouté pour le préciser). Best, Nicolas
TCK23728 | User | Service Request | TR: PO NO. 4231110156 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-12 15:36 | 2023-12-12 18:00 | 0 | 0 |  | Dear Mr Dhumak, I'm the one that sent through Anne the question of the protection, I just wanted to warn you that you order a flange class 2500 (very high pressure rating) with a gauge that can't withstand more than 10bar overpressure. it might come from the customer request because the piping is overrated, or the customer has a monoflange valve below and then everything is ok, but it could also be a mistake because an overpressure protection was missing, just please check on your end it reply to the customer request. Best, Nicolas
TCK23701 | User | Service Request | LEAD 30192092 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-12 11:40 | 2023-12-12 14:19 | 1 | 0 |  | Salut Véronique, j'ai extrait tout les MEX5....X14.... mais je ne trouve pas cette échelle spécifique il faudra la recréer. en même temps c'est un mano qui a une bonne 50aine d'années. Best, Nicolas
TCK23629 | User | Service Request | YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-11 16:16 | 2023-12-14 10:37 | 1 | 0 |  | Bonjour Véronique, je ne vois pas de problème pour la fabrication et le raccord est noté dans la désignation, après c'est une question de conformité DRS et REQ Total que tu pourras vérifier (avec Jérome au besoin). Best, Nicolas
TCK23624 | User | Service Request | TR: Question Thermometer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-11 15:55 | 2023-12-12 15:38 | 1 | 0 |  | Nathalie, on fera un TBL100-232.150.27T/9003/0100 pour remplacer. Best, Nicolas
TCK23618 | User | Service Request | TR: Purchase Order | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Lestrange, Priscilla [dpri] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-11 15:26 | 2023-12-12 16:48 | 0 | 0 |  | Bonjour Priscilla, le 10307775 est pour le MPG/J (ancien) le 10307774 est pour le MPE/F (phased out) le 11251778 pour le MPG6 redisigned actuel (attention, elle vient de l'inde avec les sous ensemble donc on a pas beaucoup de stock ici. Best, Nicolas
TCK23516 | User | Service Request | [BHFR Material Data] BAPL ¦ SAP No for Y91 from 1192383/10 | Process Instrumentation > Material Data Management | Low |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2023-12-08 12:33 | 2023-12-20 09:50 | 1 | 1 |  | Already answered.
TCK23514 | User | Service Request | CTX for Oxygen | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-08 11:52 | 2023-12-08 14:51 | 1 | 0 |  | Hi Andrew, we created a couple but it's blocked, unless it's really important business, I would suggest the Y91 which can be O2 cleaned and bonus point is ATEX. Best, Nicolas
TCK23484 | User | Service Request | Kindly confirm for TQ comments in drawing of pressure switch model RPPN | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-08 08:36 | 2023-12-08 10:31 | 0 | 0 |  | Hi Anil, Why take a 2" pipe mounting if it is to wall mount it. a standard plate is allways delivered for wall mounting, you fix the plate on the support and then clip the pressure switch on the plate. (make sure you have room below for the sensing element. [Image] Otherwise, we have an option for a stainless wall mounting as below without the 2"Pipe (10302534). [Image] Best, Nicolas
TCK23483 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5365 Pos.90 - MCX5-D60.N67.AORPB01.ARIN2606.D435-2N20.212.G2001/0078_0751_SETP_9941_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-08 08:31 | 2023-12-08 09:25 | 0 | 0 |  | Hi Rudi, it's ok. Best, Nicolas
TCK23482 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5365 Pos.110 - MCX5-D60.N66.D435-2N20.212.G2001/0078_0751_SETP_9941_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-08 08:31 | 2023-12-08 09:28 | 0 | 0 |  | Hi Rudi, OK without the SETP: MCX5-D60.N66.D435-2N20.212.G2001/0078_0751_9941_9030 Best, Nicolas
TCK23481 | User | Service Request | Check feasibility of configuration - Kaneka (21107133) - RFQ RFE 5365 Pos.170 - MEX5-D60.B24.D631-2N60.212.G2001/0751_0078_9941_9030 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-08 08:31 | 2023-12-08 09:07 | 0 | 0 |  | Hi Rudi, configuration OK. Best, Nicolas
TCK23446 | User | Service Request | Thermostats TSSE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-07 14:35 | 2023-12-14 09:50 | 1 | 0 |  | Salut Cyrille, voici la capacité des switchs standards des TSSE: [Image] Best, Nicolas
TCK23441 | User | Service Request | 11239986 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Purchase, BHDE | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-07 14:22 | 2023-12-08 14:45 | 1 | 0 |  | Hi Kerstin, the drawing you sent is the code 11239978 from the project Long Son Vietnam. this is the code that was provided, without the changes they are asking for in blue. this project will be handled as usual via the malaysian distributor and BASG, if I were you I wouldn't bother with it, there is no chance it'll be handled in EU. Best, Nicolas
TCK23423 | User | Service Request | Demande de faisabilité | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-07 12:25 | 2023-12-07 14:10 | 0 | 0 |  | Bonjour Cyrille, ça me parait bien, et le minimum de réglage est -1...0bar, par contre tu ne peux pas aller en dessous. Best, Nicolas
TCK23339 | User | Service Request | pfmn et TFRx monté sur bride Aseptique | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-06 18:00 | 2023-12-08 10:40 | 1 | 0 |  | Répondu via email groupé avec Yannick.
TCK23314 | User | Service Request | BAIT ¦ Promeco ¦ High temperature flanges | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-06 14:42 | 2023-12-06 16:57 | 1 | 0 |  | Hi Erald, if we're talking diaphragm seals I don't have solutions for you sorry, the filling fluid will be the limitation, over 400C, you'll loose accuracy and it might reduce the longevity of the mounting. but if we talk about a simple flange and you choose the right gasket for the mounting (stainless or graphite) it should be OK. Best, Nicolas
TCK23209 | User | Service Request | **SPAM** Information about data processing at AS-Schneider in terms of GDPR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-06 01:07 | 2023-12-06 10:55 | 1 | 0 |  | SPAM
TCK23208 | User | Service Request | **SPAM** Information about data processing at AS-Schneider in terms of GDPR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-06 01:07 | 2023-12-06 10:55 | 1 | 0 |  | SPAM
TCK23190 | User | Service Request | BAPL ¦ Nrder 1020022/10 ¦ Zapytanie | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-05 15:01 | 2023-12-05 15:50 | 0 | 0 |  | Perfect, thanks Anna, then this is the product, you can replace it with the same: 11107097 Y91-3.9.3.B24.R/2152 Best, Nicolas
TCK23058 | User | Service Request | CEA-COMBALUZIER-04.12.2023 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-04 14:47 | 2023-12-04 15:28 | 0 | 0 |  | Salut Yannick, ici il n'y a que la température qui pose problème et il ne faut jamais commencer par un RPP. je proposerais un RP2 avec un capillaire ou un ASIPD. comme c'est de l'huile thermique à haute température le système s'en accomodera très bien. il faudra faire attention a ne pas surdimensionner le capillaire, pour ne pas baisser trop bas en température (2m je pense mais à revoir avec le client). Best, Nicolas
TCK23052 | User | Service Request | Differential Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-04 14:23 | 2023-12-04 15:18 | 0 | 0 |  | Hi Daniel, your thinking is right and I would also start with offering this. the MT7 is not S3 as S3 requires a baffle wall, a safety window AND a blow out back so it's just S1. Best, Nicolas
TCK23042 | User | Service Request | Need General arrangement drawing against tender enquiry of GAIL | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-04 13:33 | 2023-12-04 17:31 | 0 | 0 |  | Hi Anil, I'm sorry but the 3D or dimensionnal drawing doesn't exist except for the one on the datasheet. if you need to have it created, you can see with the sales (only 2 drawings would be necessary for your 4 references as the filling or range don't change anything). [Image] Best, Nicolas
TCK22962 | User | Service Request | Question for contact gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Sung, JinKyung [sji] | Birrien, Nicolas [bin] | KRSEO - Seoul | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-04 05:29 | 2023-12-04 15:12 | 1 | 0 |  | Hello JinKyung, you're perfectly right, both the selected switches will ensure the right break and make you need for your application. the 5021 is a 3 wires which means the powering of each contact is made through the same wire and the 5065 is a 4 wires so the contacts are completely independants. Best, Nicolas
TCK22920 | User | Service Request | TR: Enquiry | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-01 17:32 | 2023-12-04 17:38 | 1 | 0 |  | Hi Kent, here is the reference to offer as a replacement: MEX5-A30.F18/0151_0670_0840_0752_0412_A009 please check the connection with them as it can be either 1/2G or M20x150 code X + C006 according to the standards. Best, Nicolas
TCK22892 | User | Service Request | Information sur capteur ZDPWN-M +10+400 MBAR CHM | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-01 13:40 | 2023-12-04 15:04 | 1 | 0 |  | Salut Véronique, si c'est du débit, il y a de l'eau dans les 2 chambres mais pas d'effet de colonne d'eau. ça ne posera pas de vrai problème, par contre on peut revoir l'étendue de mesure pour coller mieux à son process, il nous faudra la Pmax de la tuyauterie et ces pression min/max d'utilisation en amont et aval du filtre. Best, Nicolas
TCK22876 | User | Service Request | possible.... 100.....125 kPa / 100...500 kPa with MEX8 NW160 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-01 11:40 | 2023-12-04 17:52 | 1 | 0 |  | Hi Sascha, for the first one you can offer: MCX7-B30.X10/9709 (100...125 kPaA) we don't have a capsule 160mm. For the second one: MEX8-B30.X19/0840/0670/9709/9711 (100...125kPaA; technical inscription of the graduation step). On both the dial will need to be created and I don't see the connection so to be verified. the dials might be slightly different, I suggest to ask for customer validation upon creation. Best, Nicolas
TCK22855 | User | Service Request | MAR3M78 57557-1/MAR3BX3A15MN09 - NOI : 144825294 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-12-01 09:43 | 2023-12-04 14:26 | 1 | 0 |  | Salut Véronique, la connexion utilisée est celle ci-dessous et selon la quantité il sera soit plus interessant de faire un raccord intermédiaire spécifique (G1/4 - M20*1.5 n19) soit de faire des raccords spéciaux pour le MEX3, a moins de 10 pieces je penche pour le raccord intermédiaire et pour plus le raccord mano (très grossière estimation). [Image] Best, Nicolas
TCK22785 | User | Service Request | TR: question / manomètre pneumatique / douille pour ampoule. CFF, Ateliers d'Yverdon-les-Bains, site d'Yvonand | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-30 14:25 | 2023-11-30 15:43 | 0 | 0 |  | Nathalie, c'est un DR96 de chez KHW, on en fera plus de nouveaux mais on peut encore peut être réparer quelques pièces en claim. Best, Nicolas
TCK22784 | User | Service Request | EQUIVALENCES DRO | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-30 14:11 | 2023-12-01 09:12 | 0 | 0 |  | Bonjour Véronique, je proposerais un remplacement par des MEX3, ci-dessous les équivalents: MEX3-D51.V19 MEX3-D51.V27 la seule vrai différence va être le tube et connexion en 316L au lieu de laiton. Best, Nicolas
TCK22758 | User | Service Request | RP2E Pressure Switch | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-30 10:38 | 2023-11-30 11:35 | 0 | 0 |  | Hi Daniel, the best fit if they work with AC is the L, otherwise the R set with a minimum deadband. Best, Nicolas
TCK22678 | User | Service Request | EU declaration of conformity | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-29 12:19 | 2023-11-29 15:29 | 1 | 0 |  | Hi Paul, sure, here it is: Best, Nicolas
TCK22636 | User | Service Request | ARPX sur H2 -- AIR LIQUIDE DI lyon - Mr Decke. | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-29 09:30 | 2023-12-12 17:40 | 1 | 1 |  | Bonjour Yannick, Voici mes réponses à tes questions: As-tu des information complémentaire de la part de notre fournisseur sur dautre norme applicable que celle mentionné dans la documentation DIN 16270 ? les informations sur les materiaux utilisés prennent en compte les databases SANDIA, l'ISO/TR 15916 (édition 2015-12-15), la SAE J2579 (édition 06/2018), l'IGC DOC 121/14 et les recommandations d'un laboratoire spécialisé. Peux -t-on créer la référence ARPX03L/1803/0765 ? Oui, cette référence évoluera surement avec l'arrivée d'une nouvelle option spécifique a l'utilisation sur H2. Peut-on réaliser un test Q001 ancienneté Q1201 test détanchéité ? les versions H2 passent des tests réguliers en production et un certificat de conformité mentionnant l'utilisation sur H2 sera possible sinon, un test hélium sera aussi faisable Concernant l'écrou de presse étoupe, y a t'il risque d'éjection (possibilité qu'il se desserre lors de l' ouverture maximale par action sur le volant) ? l'écrou de purge est serré et ne devrait pas bouger. Pour la poignée de fermeture du robinet, il y a un blocage d'ouverture maximale qui empêche la sortie du pas de vis et 4 joints d'étanchéité. J'éspère que ça réponds a toutes tes questions, et quand tu passera a la demande de prix, il y aura peut être l'option 0766 qui sera pour l'utilisation sur H2 (mais ce n'est pas encore fait ). j'ai aussi lancé une consultation pour: ARPX03L/1803/H2 ARPX03L/H2 on devrait avoir des prix à te donner quand tu les donnera. Best, Nicolas
TCK22601 | User | Service Request | Doubt about thermometer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | VECAR - Caracas | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-28 21:25 | 2023-11-29 16:42 | 1 | 0 |  | Hello Angel, the number before BIL is for us to know the exact sensing element that was used and allow us to make traceability. in the same prduct, we could use several different sensing element to allow a more flexible and precise calibration and the marking keeps track of this. Best, Nicolas
TCK22542 | User | Service Request | MEX5 + D435 - Medium H2SO4 approx 53% upto 83% | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-28 11:43 | 2023-11-28 14:16 | 1 | 0 |  | Hi Rudi, it's really a question of long term reliability vs cost and the answer is more on the customer side with the costs. I would offer an HC diaphragm and PTFE liner anyway because it's not too expensive and will give best protection. then the coating of the bottom part is really up to him as it's a long term risk. Best, Nicolas
TCK22514 | User | Service Request | MIX5-DAPO-B20 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-28 09:52 | 2023-11-28 14:20 | 1 | 0 |  | Hi Daniel, I didn't think it meant anything but I still checked the Bourdon Haenni old datasheet and the codification system was the same as what we do know (except it was MIX instead of MEX). go ahead with a MEX5D30B20 Best, Nicolas
TCK22465 | User | Service Request | TR: zapytanie na Y913 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-27 14:05 | 2023-12-05 09:31 | 0 | 0 |  | Closed externally with the customer. Best, Nicolas
TCK22440 | User | Service Request | California proposition 65 legislation - Atlas Copco | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-27 12:01 | 2023-11-27 15:24 | 1 | 0 |  | Hi Rudi, there is a document from the group for the customers on Cal Prop 65, you can find it enclosed: Best, Nicolas
TCK22389 | User | Service Request | TR: #INC284438 - Création BAFR // Client : TMH AMS // Cde Clt : à venir // MEX3-BX0.X38/0789_0762_9911_9712_9710_++ // Position X Your Support Ticket was received | Service Desk > General | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-27 09:17 | 2023-11-27 11:40 | 1 | 0 |  | Bonjour Marie claire, tu envois sur la mauvaise boite, il faut envoyer à : FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> Best, Nicolas
TCK22353 | User | Service Request | [BHFR Material Data] FW: Offerte omvormer type Y913 | Process Instrumentation > Material Data Management | Low |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Dolbeau, Nadia [ndo] | 2023-11-24 16:55 | 2023-12-04 14:19 | 1 | 1 |  | Email sent instead of closing sorry. Nicolas
TCK22351 | User | Service Request | FW: 202311241521 RE: inquiry_k 202311211312 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 16:31 | 2023-11-27 09:51 | 1 | 0 |  | Hi Paul, it's because in terms of deadband, the microswitch C will have the same values you can see when used with a range 234. so he minimum deadband would be comparable to the full scale, so not usable. this have to do with the travel of the sensing element that is similar on the 214, 224 and 234 and makes t usable only on the 234. Best, Nicolas
TCK22320 | User | Service Request | THERMOSTATS ZT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 12:18 | 2023-11-28 16:34 | 1 | 0 |  | Bonjour Véronique, je suis un peu étonné de la question, en générale pour ce type de mesure de niveau, la pression haute est prise sur le ciel gazeux en haut de cuve donc il n'y a pas de poids de colonne de liquide et sur la partie basse, ce poids nous donne la pression hydrostatique et donc la mesure. Si tu peux avoir plus d'informations ce serait super pour comprendre. Je clos le ticket pour l'instant car je connais les délais de réponse de EDF, mais n'hésite pas à me recontacter avec des détails, je serais ravis de donner d'autres infos. Best, Nicolas
TCK22289 | User | Service Request | TR: BP-20230255-R0 RE: RFQ for the Provision of BP-20230197-R0 SG-0000452-R00 OGDCL - Pressure Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 09:13 | 2023-11-24 09:29 | 0 | 0 |  | Donc ca devrait etre cette connexion, a confirmer: [Image] Best, Nicolas
TCK22288 | User | Service Request | TR: 20271-521 & 20271-522 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 09:12 | 2023-11-24 09:27 | 0 | 0 |  | Nathalie, voici le certificat IECEx: Best, Nicolas
TCK22285 | User | Service Request | MFT5+5way manifold | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BEBRU - Westerlo | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 09:03 | 2023-11-24 09:12 | 0 | 0 |  | Jurgen, il te faut 2 ARIN26F8 entre le MFT et le manifold. Best, Nicolas
TCK22269 | User | Service Request | Kindly confirm process connection of 9/16-18-UNF-2B (F) for 0/1600 bar range in Model MEX5 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | INMUM - Mumbai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-24 06:20 | 2023-11-24 09:25 | 0 | 0 |  | Hi Anil, it would be this connection code and drawing but there is pretty much no gauge created with it and for 1unit I'm afraid the price will be really high and 1/2NPT is suitable for 1600bar. C010 9/16 - 18 UNF (type 85) [Image] Best, Nicolas
TCK22219 | User | Service Request | TEMPERATURE AMBIANTE THERMOSTATS ZT | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-23 12:03 | 2023-11-23 15:58 | 1 | 0 |  | Salut Véronique, ça ne craint rien pour le fonctionnement de l'appareil, par contre ça pourrait avoir une légère influence sur la précision. Best, Nicolas
TCK22216 | User | Service Request | TSF100 P2F issues | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-23 11:55 | 2023-11-23 13:59 | 1 | 0 |  | We'll check this out together, do you have availability this afternoon for a quick teams ? feel free to set a meeting in my agenda, I should be free from 3pm. Best, Nicolas
TCK22197 | User | Service Request | BAPL ¦ create product ¦ Y91-3.3.3.B24.R/2151+DANC01-2L2A.01 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-23 10:14 | 2023-11-23 11:47 | 0 | 0 |  | Hi Anna, the mounting is OK (you'll have the options added for accuracy x2 and mounting with seal to add. but for the code creation request, you need to send it to FUNC MATERIAL DATA [Image] Best, Nicolas
TCK22184 | User | Service Request | CDE 5100-4210543853 POSTE 40 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-23 09:25 | 2023-11-23 11:42 | 1 | 0 |  | Je ne comprends pas, la référence qu'ils ont existe mais c'est un TSF: 0657.5297 TSF100-222.1B2.27T.70/738C_0777 10261490 et pour définir une valve il me manquerait des infos, il va falloir vérifier avec le client ou est l'erreur et si il veut vraiment une valve il faudra la définir. Best, Nicolas
TCK22144 | User | Service Request | TR: MEX5 with 1/2" tube stub | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 17:03 | 2023-11-23 09:10 | 1 | 0 |  | Nathalie, c'est possible et ce serait une connexion C038 (ou C036) et avec cette pression c'est bon. [Image] Best, Nicolas
TCK22142 | User | Service Request | RE: CAF - Fwd: RV: CR1 - Mallorca - transductor X.18X02212X01 ¦ OP CAF Construcciones y auxiliar 12062023 20522046 ¦ Ped 1037 PO 1410587 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 16:30 | 2023-11-23 09:12 | 1 | 0 |  | Bonjour Jacques, c'est une connexion ancienne qu'on a pas sur place donc c'est un peu compliqué de revenir aux plans fournisseur et même dans les données que j'ai pu trouver ce n'est pas défini. Mais je continue mes recherches. Best, Nicolas
TCK22138 | User | Service Request | NCDE/N Orders: 1331042/10 ¦ zapytanie ofertowe | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | PLWAR - Łódź | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 15:57 | 2023-11-23 09:06 | 1 | 0 |  | Hi Anna, here is a screen print of the order: [Image] best, Nicolas
TCK22136 | User | Service Request | DRO 80/311.111.440/8907-0004 process connection | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 15:52 | 2023-11-23 09:05 | 1 | 0 |  | Hello Zhang Wei, the connection behing the G1/2 is a G1/4 in the code: 11167373 DRO 80/311.111.440/8907_0004 Best, Nicolas
TCK22097 | User | Service Request | DROE | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 11:17 | 2023-11-22 14:21 | 1 | 0 |  | Hi Sascha, we can't mount a gauge on a D055 1/2" with a 25bar range, the minimum normally is 250bar. the mounting with a 3/4" is already below the standard limit. Best, Nicolas
TCK22059 | User | Service Request | TR: BP-20230255-R0 RE: RFQ for the Provision of BP-20230197-R0 SG-0000452-R00 OGDCL - Pressure Gauge | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-22 08:47 | 2023-11-22 09:26 | 0 | 0 |  | Nathalie, Quand on dit 9/16 HP, on parle je pense d'un 9/16"-18 UNF comme filetage (donc 18 filets par pouce). par contre il faudra valider si la forme est celle ci-dessous: [Image] Best, Nicolas
INC284463 | User | Incident | RE: #INC284438 - Création BAFR // Client : TMH AMS // Cde Clt : à venir // MEX3-BX0.X38/0789_0762_9911_9712_9710_++ // Position X Your Support Ticket was received | Archive > BHFR Technical Support | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-11-22 08:36 | 2023-11-24 09:58 | 1 | 0 |  | Bonjour Marie Claire, vous vous etes adréssée au mauvais service, il faut envoyer votre demande de création de code a FUNC MATERIAL DATA. Best, Nicolas
INC284462 | User | Incident | RE: #INC284435 - Création BAFR // Client : GUILMAN // Cde Clt : CAFR1423-065875 // MEX5-D20.N59 // Position 50 Your Support Ticket was received | Archive > BHFR Technical Support | Low |  |  | Closed | Labrosse, Marie Claire [labm] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-11-22 08:09 | 2023-11-24 09:59 | 1 | 0 |  | Idem, mauvais service de ticketing. Best, Nicolas
TCK21999 | User | Service Request | TR: Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-21 12:23 | 2023-11-21 16:36 | 1 | 0 |  | Nathalie, J'avais répondu à Giorgio le jour même mais tu n'étais pas en copie: Bonjour Giorgio Cest une de ces 2 options et vu lombre et la position de laiguille, je pencherais pour 0060 laiguille a maxima. [Image] Best, Nicolas
TCK21993 | User | Service Request | TR: BAIT - Water System - Price for combination | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-21 11:32 | 2023-11-21 16:35 | 1 | 0 |  | Anne, ça ne va pas marcher, le minimum de montage pour MEX5 + DANC01 est 6bar et avec un MX35 ce serait -1...1.5bar le plus proche possible. Best, Nicolas
TCK21977 | User | Service Request | TR: Question | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-21 10:13 | 2023-11-21 17:44 | 1 | 0 |  | Nathalie, tu peux lui proposer en remplacement: 10319947 / DRO 63/412.111 -1..24BAR R134A MMD3 FI2 B82/9712/Dxxx (Dxxx to be checked but existing) Best, Nicolas
TCK21921 | User | Service Request | TR: Prix et délai MEP3-120.B36/0816_0761_0769_0655_0792 MAGYAR | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-20 17:17 | 2023-11-21 17:10 | 0 | 0 |  | Nathalie, voici ce qu'on peut faire ou pas sur un MEP3: 0816: no 0655: no 0769: no but 0770 ok (0.4mm instead 0.3mm) Best, Nicolas
TCK21909 | User | Service Request | YOKOGAWA | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-20 16:07 | 2023-11-20 16:34 | 0 | 0 |  | Bonjour Véronique, j'ai préféré vérifier avec Jérome pour ça et on ne le propose pas. il est au courant que Total continue à le demander mais on a pas la capabilité de le faire en interne. Best, Nicolas
TCK21908 | User | Service Request | D912 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-20 16:05 | 2023-11-20 17:28 | 0 | 0 |  | Ok, je vois mieux, Le SIL concerne uniquement l'éléctronique et ne peut pas être applicable. Par contre, selon le THR, Total a accès à la SP2706 qui donne des données de fiabilités à utiliser pour les calculs de fiabilité de l'ensemble (THR 5.2). Best, Nicolas
TCK21869 | User | Service Request | TR: DEMANDE ATELIER | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-20 12:32 | 2024-01-10 15:27 | 1 | 1 |  | Clôture en attente de réponse, autre solution envisagée. Best, Nicolas
TCK21843 | User | Service Request | BAIT ¦ Fluid Processing ¦ AKPL tech information | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | ITASS - Assago | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-20 10:08 | 2023-11-20 10:57 | 0 | 0 |  | Hi Dema, yes, this connection is standard and necessary for the differential mounting. Once mounted, the connection will be tightenned and fix and the capillary will be allowed some rotation that will avoid putting torsion efforts on it. Best, Nicolas
TCK21792 | User | Service Request | TR: MEX5-D30.B73 / 0751 / 0078 / 9941 + D631 – 2 . L 6 . 0 4 C E. A 6 0 + LRS1 / 9030 de dietrich | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-17 16:40 | 2023-11-20 10:54 | 1 | 0 |  | Nathalie, c'est bon, par contre le 9030 ne doit pas etre possible sur du G1/2, il faut faire les 3 points de soudure (ou passer en NPT). Best, Nicolas
TCK21781 | User | Service Request | TR: urgent inquiry_k 202310251204 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-17 15:22 | 2023-11-17 16:06 | 0 | 0 |  | Le premier pas possible, le reste ça va. Best, Nicolas
TCK21780 | User | Service Request | TR: urgent inquiry_k 202310251204 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-17 15:18 | 2023-11-17 16:06 | 0 | 0 |  | Nathalie, tu as raison, pour le premier ca va etre trop juste, il faudrait au moins 2.5bar. Best, Nicolas
TCK21768 | User | Service Request | WG: Preisanfrage | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-17 13:14 | 2023-11-21 17:34 | 1 | 0 |  | Hi Felix, here is what whe can offer instead of: DROE 100/213.111/065/0794/9711/1020/5001/4726 Replaced by: MR5 2D3K B20/0794/9711/4710 + 1020/LRS7/0794 9711: "Silikon Frei" 4726 connector not available as standard if necessary, please make a special request. Best, Nicolas
TCK21607 | User | Service Request | TR: Tube stub pricing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-16 08:55 | 2023-11-16 09:04 | 0 | 0 |  | Nathalie, en raccord tubulaire (OD) on aura rien au dessus de 100bar désolé. Best, Nicolas
TCK21594 | User | Service Request | DN40 PN20 PN25 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-16 07:49 | 2023-11-16 09:06 | 0 | 0 |  | Bonjour Véronique, tu trouvera le dimensionnel sur les datasheets des séparateurs (exemple pour le D820): [Image] Best, Nicolas
TCK21517 | User | Service Request | UL certificate for CTX | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | CNSHA - Shanghai | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-15 10:47 | 2023-11-15 14:56 | 1 | 0 |  | Hi Xu Qiang, please find the UL certificate attached in the ticket. Best, Nicolas
TCK21513 | User | Service Request | MFT5-P03.X10X.N2/9704 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | GBSWI - Watchfield, Swindon | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-15 10:33 | 2023-11-15 15:25 | 1 | 0 |  | Hi Daniel, it would have been my assumption too but I don't see anything created. but you can ask directly to the sales for price like this: MFT5-P03.X10X.N2/9704 "https://servicedesk.baumernet.org/wm/" Special pressure units (mbar / kg/h) Pressure range: 0 ... 200 mbar and 0 11800kg/h there will probably be a fee for dial creation. Best, Nicolas
TCK21320 | User | Service Request | TR: Diaphragm seal pricing | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | FRVEN - Vendôme | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-13 14:03 | 2023-11-13 17:21 | 0 | 0 |  | Nathalie, j'aimerais qu'on travaille plutôt avec ce document qui se basait sur des quantités vendues, contient les accessoires, les MOQ et ou tous les codes SAP sont dispo, il faudra juste le vérifier si ça va à Tereasa. si c'est OK pour elle je fais la verif, sinon je suis dispo pour un call. Best, Nicolas
TCK21276 | User | Service Request | WG: Enquiry REF ML-10111808-Baumer | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | DESTO - Stockach | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-13 10:16 | 2023-11-13 16:00 | 1 | 0 |  | Hi Felix, sure you can, there will be a fee as usual for the dial creation and we'll have to confirm what you want is as below, but otherwise it's possible: 0 mbar = 0kg/h 250mbar = 14750kg/h Linear scales. Best, Nicolas
TCK21271 | User | Service Request | TR: CAF - Fwd: RV: CR1 - Mallorca - transductor X.18X02212X01 ¦ OP CAF Construcciones y auxiliar 12062023 20522046 ¦ Ped 1037 PO 1410587 | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | FRFIL - Fillinges | [Role] SA-Technical Support FRVEN (SA-TS) | Birrien, Nicolas [bin] | 2023-11-13 09:59 | 2023-11-27 17:46 | 1 | 1 |  | Bonjour Jacques, j'ai eu une première réponse du fournisseur (que je trouve un peu bizarre, je leurs ai redemandé si c'était bien un M16 ou M20 ?) je clos le ticket avec cette réponse et je te redirais si j'obtiens mieux. [Image] Best, Nicolas
TCK21201 | User | Service Request | Test Technical Support | Process Instrumentation > Sales Technical Support | Low |  |  | Closed | Zuberbühler, Lucca [zlu] | Zuberbühler, Lucca [zlu] | CHFRA - Frauenfeld | [Role] SA-Technical Support FRVEN (SA-TS) | Zuberbühler, Lucca [zlu] | 2023-11-10 17:00 | 2023-11-13 08:59 | 0 | 0 |  | -
INC284430 | Workgroup | Incident | 3D drawings | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-09 15:06 | 2023-11-10 13:24 | 0 | 0 |  | Hi Kerstin, we don't have 3D for those and I can't disclose the complete technical drawing, but here is a detailled dimensional for all 4 types: [Image] Best, Nicolas
INC284428 | Workgroup | Incident | demande de KNORR BREMSE | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-09 12:27 | 2023-11-13 14:45 | 0 | 0 |  | Cyrille, j'ai un peu plus de détails, il semblerait qu'on ai fait un TB233 dans le debut des années 90 pour ce type d'applications. Evidemment on ne le fait plus mais il etait basé sur la meme techno que le Y913 qui pourrait etre un remplacement. par contre je n'ai pas plus de detail sur la codif, on a pas retrouvé de fiche technique. l'EF6 ou le PPxxR pourraient aussi en cas de besoin de l'EN50155. Best, Nicolas
INC284405 | Workgroup | Incident | TR: OFFRE ACTEMIUM DDP URGENT 20539383 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-08 16:39 | 2023-11-09 09:15 | 0 | 0 |  | Nathalie, c'est OK pour moi avec ces PN/DN
INC284399 | Workgroup | Incident | Assistance | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-08 11:27 | 2023-11-09 12:13 | 1 | 0 |  | Hi Paul, we don't have much on this but mainly because we don't have much polluting processes or chemicals. you can find our statements on the page 4 of the quality manualthat is the main document we can provide. Best, Nicolas
INC284393 | Workgroup | Incident | BAPL ¦ replacement ¦ ED 510/314.241/075 | Archive > BHFR Technical Support | Middle |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-08 09:45 | 2023-11-08 10:47 | 0 | 0 |  | Hello Anna, you can offer a Y9233XB22R and the customer can choose one of the below connection for X (probably the G for replacement): [Image] The Y920 is not relaunched yet but it'll happen shortly, you can offer this one already. Enclosed are both datasheets: Best, Nicolas
INC284380 | Workgroup | Incident | TR: MEX5 rear mount with 3/8" tube stub | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-08 09:00 | 2023-11-08 09:27 | 0 | 0 |  | Nathalie, on a jamais fait de 3/8 tube lisse, ce serait une nouvelle demande et on aura le meme probleme que sur les OD disponnibles qui est qu'on a pas de montage pour aller au dessus de 100bar (1000psi/60bar echelle). Best, Nicolas
INC284367 | Workgroup | Incident | EN 1092-1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-08 07:54 | 2023-11-09 11:57 | 0 | 0 |  | Veronique, a part pour des considérations specifiques de matieres exotiques par exemple, le PN20 et 25 pour un DN40 sera le meme, tu peux le voir dans les tables de cotes des datasheets des séparateurs. Best, Nicolas
INC284349 | Workgroup | Incident | RFQ 71123.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-07 11:00 | 2023-11-07 11:31 | 0 | 0 |  | Hi Paul, here is your mounting code: MEX5-A30.B22.AKPL2F5302.D043-2L03.09 it is the standard mounting, Best, Nicolas
INC284342 | Workgroup | Incident | TR: MCX5-D60.R06/0751_0722/ Theerasak | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-07 09:57 | 2023-11-10 14:50 | 0 | 1 |  | Nathalie, nous n'avons pas de montage disponnible pour faire le poinconnage donc ouvrir cette option sur les MCX (0722) serait possible mais demande un petit travail de codification et de mise en place d'outillage. je preconniserais de l'ouvrir sous SPR. Best, Nicolas
INC284339 | Workgroup | Incident | TR: NTECH: Zeichnung zu Anzeige D180 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-07 09:09 | 2023-11-07 10:21 | 0 | 0 |  | Nathalie, c'est ce code SAP, mais ce n'est pas de chez nous, c'est lié a BAFR mais je ne sais pas trop comment: 11041943 D180P.03S4A01 Best, Nicolas
INC284330 | User | Incident | TR: ZENITH - RFQ - MANOMETRE | Archive > BHFR Technical Support | Low |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-11-06 13:48 | 2023-11-09 09:12 | 1 | 0 |  | treated via email
INC284329 | Workgroup | Incident | FW: W 201 589 13 21 00 Pressure tester not working. | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 11:30 | 2023-11-07 10:16 | 1 | 0 |  | Hi Anil, this is specific material we only supply to mercedes germany, they have to contact them, we cannot handle directly. they'll probably have to provide more details as it was calibrated and working fine in august this year. Best, Nicolas
INC284328 | Workgroup | Incident | ZPHN / OMEXOM | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 11:18 | 2023-11-07 11:21 | 1 | 0 |  | Véronique, je suppose que ce n'est pas pour du nucleaire, dans ce cas je proposerais la reference ci-dessous, c'est le remplacement direct pour le micro GS: Order code: RPPN-4E3.102/SETP RPPN old Code / SAP code: Sensing element and diaphragm: Type: PHN - medium pressure 0 ... 2500 mbar 4 Microswitch: 1 ultra sensitive changeover switch E Hydraulic connection: G1/2" male 3 Pressure range: Pressure range (mbar) -2 ... 10 mbar 102 OPTIONS Accuracy / Setting: Set point adjustment SETP Best, Nicolas
INC284327 | Workgroup | Incident | WG: Anfrage 794490 YK | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 10:57 | 2023-11-07 10:13 | 1 | 0 |  | Hi Felix, seems like a sliding connection with compression fitting to me, you can find it on the website: Selector ¦ Sliding screw connection ¦ Bourdon Instruments (bourdon-instruments.com) "https://www.bourdon-instruments.com/fr/en/product-overview/accessories-temperature/sliding-screw-connection/c/37882/products" Best, Nicolas
INC284318 | Workgroup | Incident | WG: Angebot STP-Modell für Manometer DR100/211.111.065 (C) | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 09:06 | 2023-11-07 09:58 | 1 | 0 |  | Hi Felix, sorry but we don't have the STEP file for those, you can find the dimensionnals 2D on the datasheet. Best, Nicolas
INC284317 | Workgroup | Incident | 2108243 / 11264746 / MX7-130.X15G/9941_9709_9704_DC50 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 09:03 | 2023-11-07 09:57 | 1 | 0 |  | Bonjour Véronique, 1bar differentiel meme à l'envers ne devrait pas endommager les souflets. par contre le cadran me parait bizarre, ils n'ont pas demandé leur équivalence à l'envers ? Parce que là, le mano est bien à 0bar differentiel. Best, Nicolas
INC284315 | Workgroup | Incident | PBSN-1.1.B15.R.A1.44.03.2.1 (11219815) | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-06 08:45 | 2023-11-07 10:57 | 1 | 0 |  | Hi Bee Tin, sorry, the G1/2 aluminum gasket is not provided with the PBSN. Best, Nicolas
INC284309 | Workgroup | Incident | TR: RFQ 21123.02 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-03 16:45 | 2023-11-07 11:27 | 1 | 0 |  | Nathalie, c'est à eux de nous donner les valeurs de diamètre en haut et en bas (sachant que 13mm est le diamètre mini et 29mm le maxi) par exemple: 25mm sous bride et 13mm au bout. Best, Nicolas
INC284296 | Workgroup | Incident | TR: Check for TED50 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-03 09:18 | 2023-11-03 15:54 | 0 | 0 |  | Nathalie, as discussed it's too old to find the detailled information, and it means I can't see if there is internal options. but from the marking and the picture, I would advise it's a TED-5.3.3.B20.R (seems like a G1/2 but to be confirmed). Best, Nicolas
INC284289 | Workgroup | Incident | TR: Enquiry_HAENNI | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-02 18:00 | 2023-11-03 09:20 | 0 | 0 |  | Nathalie, on a plus de transmetteur differentiel, ni chez nous ni chez Baumer, je ne vois pas de vrai remplacement. par contre il y a d'autres solutions possibles et qui peuvent etre moins cher comme par exemple utiliser 2 transmetteurs 4...20mA et faire la difference sur les signaux via un automate ou un afficheur. ou utiliser du mecanique (ce qu'on a). Best, Nicolas
INC284285 | Workgroup | Incident | RFQ 21123.03 | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-02 15:15 | 2023-11-02 17:11 | 0 | 0 |  | Hi Paul, here you go just have them check the microswitch, it's the same type but we don't exactly use the same ones sine 1995: [Image] Best, Nicolas
INC284274 | Workgroup | Incident | MEX5D62B22/0751+D1112N202N0/LRS1 /Unico | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-11-02 09:27 | 2023-11-02 10:56 | 0 | 0 |  | Hi Bee Tin, the D111 became the DT1, here is the new code: MEX5D62B22/0751+DT12N2002N001 Best, Nicolas
INC284256 | Workgroup | Incident | RFQ311023.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-31 16:21 | 2023-10-31 16:32 | 0 | 0 |  | Hi Paul, you might be working on the same offer as Jurgen De Mezel, he's defining the same product. you can find it in this table, for you it'll be 200mm [Image] Best, Nicolas
INC284253 | Workgroup | Incident | marine approval for Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-31 13:24 | 2023-10-31 14:29 | 0 | 0 |  | Dear Xu Qiang, please find the letter stating the phase out of the marine approval and the potential replacement. Best, Nicolas
INC284250 | Workgroup | Incident | BAPL ¦ LPGR1 ¦ replacement | Archive > BHFR Technical Support | Middle |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-31 11:30 | 2023-10-31 11:39 | 0 | 0 |  | Hi Anna, you're right, the LPGR1 is outdated and is replaced by the flexprogrammer: [Image] Best, Nicolas
INC284248 | Workgroup | Incident | Valve request | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-31 09:30 | 2023-10-31 10:15 | 0 | 0 |  | Hi Paul, I would offer either a 2VT from BTIN (2VT1 is more standard as the vent is over the valves but yours seems lika a 2VTF) 11176288 2VT.SCTP.44FM.N or Schneider: 11226376 XMFD:SAMFMA-N4N4-A Best, Nicolas
INC284241 | Workgroup | Incident | TR: RFQ 2023394 - Vacuum pressure gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-30 16:27 | 2023-10-30 17:08 | 0 | 0 |  | Nathalie, c'etait du specifique notament la connection, et on a pas eu de commande au moins depuis 2009 quand il a été bloqué. je doute du coup qu'on refasse du spécifique pour ce client, il faudra voir plutot avec un Y91-4.B.3.B59.R/2160/0765 qui serait plus adapté pour une utilisation sur Oxygène (degraissage O2 et ATEX) mais attention connection standard et pas d'inversion de signal. Si l'utilisation n'est pas vraiment pour de l'O2, alors un CTX-4.B.3.B59.2/4780 pourrait etre possible mais la aussi la connexion est standard. J'ajouterais que sur le long terme il est toujours mieux d'utiliser des connexions standard pour avoir un remplacement plus facile. Best, Nicolas
INC284240 | Workgroup | Incident | BAIT ¦ UNILEVER ¦ Pressure transmitter with APV-ISS connection | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-30 15:36 | 2023-10-30 16:46 | 0 | 0 |  | Hi Erald, I'm sorry but we don't have this norm for seals in our standard. but for 70-80units a year it might be worth an SPR, please send one to Jorg and myself for this opportunity and we will analyse it. Best, Nicolas
INC284237 | Workgroup | Incident | S/N request | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-30 14:51 | 2023-10-30 15:30 | 0 | 0 |  | Hi Jorg, sorry it's not in the traceability software (too old ?), nor in the production quality documents database. Best, Nicolas
INC284235 | Workgroup | Incident | I: BAIT ¦ Sterilgarda ¦ inline separator + PBMN | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-30 13:30 | 2023-10-31 14:24 | 0 | 0 |  | Hi Erald, as far as I can judge from the picture, I would thing it's a 1627, he picture is hereafter and the DN of the clamp on this one is 25mm. [Image] it can be fit with a PBMN at -1...5bar just the accuracy will be 0.5% instead of 0.25% and you need to choose the G1/2 EN connection to fit with it. Best, Nicolas
INC284232 | Workgroup | Incident | FW: MEX3 diaphragm seal, PG Flowteknik | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-30 11:36 | 2023-10-30 15:49 | 0 | 0 |  | Hi Paul, here are what we could offer in terms of tube connections and we are limited to 70bar. [Image] On the MEX3we can only do the small diameters (OD6 and OD1/4"). Best, Nicolas
INC284220 | Workgroup | Incident | TR: BAIT ¦ INTECH BOURDON - 251002125 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-27 15:54 | 2023-10-27 16:23 | 0 | 0 |  | Anne, c'etait ce code, à l'époque le SETP était précisé à la commande, maintenant il faudra recréer le code avec le SETP intégré au code SAP. 10297692 RPPN-3E6.104/SETP Best, Nicolas
INC284215 | Workgroup | Incident | TR: BAIT ¦ Sterilgarda ¦ inline separator + PBMN | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-27 11:57 | 2023-10-30 15:32 | 0 | 0 |  | Doublon demande directe.
INC284212 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-27 09:57 | 2023-10-27 14:59 | 1 | 0 |  | Nathalie, On a pas le joint exact qu'ils demandent, ça semble etre une fourniture de frauenfeld donc je leurs demanderais si j'etais eux. par contre en dépannage on a le joint du G1/2 qui est très proche et pourrait pour moi assurer la fonction: [Image] Best, Nicolas
INC284211 | Workgroup | Incident | TR: WEP5 quote | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-27 09:30 | 2023-10-27 14:45 | 1 | 0 |  | Nathalie, je vais appeler Jim pour clarifier. Best, Nicolas
INC284188 | Workgroup | Incident | PLAN 52298-03 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-26 14:33 | 2023-10-26 16:30 | 0 | 0 |  | Bonjour Véronique, envoyer le plan que tu demandes n'est pas possible car c'est un plan détaillé de l'appareil complet. par contre, tu trouveras ci-dessous le détails de la connexion: [Image] Best, Nicolas
INC284187 | Workgroup | Incident | WG: Anfrage Manometer Befüllen und Nullen | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-26 12:27 | 2023-10-26 13:57 | 0 | 0 |  | Hi Jorg, here is the code and designation: 11187776 MEX5-B20.B33/0751 Best, Nicolas
INC284182 | Workgroup | Incident | PRESSOSTAT AMBIANCE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-26 11:00 | 2023-10-26 11:48 | 0 | 0 |  | Bonjour Veronique, de memoire je ne pense pas, et nos ZPN ne seraient pas adaptés car l'interieur et l'exterieur de l'element sensible seraient a la Pa si il est simplement installé dans une piece. par contre ca pourrait marcher si il depressurisent (ou surpressurisent) légèrement une pièce pour des raisons de propreté ou autres. mais il faudrait que la connection se fassent comme pour un process (la pression de la piece reliée a la connection de l'instrument qui sera placé a l'exterieur de la piece. Best, Nicolas
INC284175 | Workgroup | Incident | CASE POLYPROPILENE UV RADIATION | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-25 23:57 | 2023-10-30 12:11 | 0 | 0 |  | Hi Angel, I discussed a bit within R&D and we tested the case with High UV constraints (similar to full sun desert location) in order to test the gaskets and the MPG passed the test with no apparent damage and functionnality ok after the test. but the test was not done in the purpose of testing the case itself and I'm not aware of a specific test or homologation for that. another point is that we sell those gauges since at least 30years without any return that was linked to this cause. And lastly, our new case for the MPG6 Redesign is in PBT and though will not be subject to this potential issue. I hope this helps you answer the customer. Best, Nicolas
INC284172 | Workgroup | Incident | BAPL ¦ request for sensor data from order 213402/10 | Archive > BHFR Material Data | Middle |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Dolbeau, Nadia [ndo] | 2023-10-25 16:18 | 2023-10-26 09:42 | 0 | 0 |  | Hello Estera, it was an E91333B18R2151 which is replaced now by the 11106495 Y91-3.3.3.B18.R/2151 Best, Nicolas
INC284163 | Workgroup | Incident | FW: Our inquiry / prijsaanvraag 22369966 | Archive > BHFR Material Data | Middle |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Dolbeau, Nadia [ndo] | 2023-10-25 10:57 | 2023-10-26 09:49 | 0 | 0 |  | Hi Kees, here is the datasheet, I guess a PBCN would have been good for this but now I don't know if there is a new capacitive sensor at Baumer. otherwise, our Y91 is robust and can probably do then job. Best, Nicolas
INC284160 | Workgroup | Incident | MFE Railcar PG's ? - Infrabel (50157534) - RFQ 20231020 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-25 10:33 | 2023-10-25 17:08 | 1 | 0 |  | Hi Rudi, you can present them the MFE that could be a potential replacement for those. please find the datasheets enclosed: Best, Nicolas
INC284152 | Workgroup | Incident | For MFE523323JF2-01 / 6346 34 010-01 - SAP 11092558 - Bombardier (50110507) | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-25 09:57 | 2023-10-26 10:17 | 1 | 0 |  | Hello Rudi, please find enclosed the conformity certificate for the MFE: Best, Nicolas
INC284141 | Workgroup | Incident | TR: VIS THERMOSTATS PRESSOSTATS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-24 15:03 | 2023-10-26 10:03 | 1 | 0 |  | Salut Véronique, voici le détail des 2 vis utilisées pour fixer la plaque de montage mural sur les cotés de l'appareil: 10304045 VIS TH M4x6 INOX nous ne précisons pas de couple de serrage pour ces vis. Best, Nicolas
INC284140 | Workgroup | Incident | TR: BAIT - Jaes - MEX5D31B18/0751+D2012L20230+ARPX03L | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-24 15:00 | 2023-10-25 11:56 | 1 | 0 |  | Anne, avec un DT2, un remplissage et si l'ARPX est en dessous du montage c'est OK: MEX5D31B18/0751+DT22L20023001+ARPX03L Best, Nicolas
INC284136 | Workgroup | Incident | COMPATIBILITE MATIERE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-24 12:03 | 2023-10-24 14:43 | 0 | 0 |  | Salut Véronique, Tout d'abord le fait de connaître la compatibilité d'un produit avec les matières en contact incombe au client qui gère le produit. Après j'ai pu regarder dans nos tables de compatibilité et l'HC276 pour la membrane est noté A avec l'acide sulfurique concentré et le 316L est noté B ce qui ne me parait pas problématique pour les parties massives. Le problème pourrait être le joint graphite qui ne semble pas vraiment compatible. Attention aussi, lors du processus de desulfurisation (traitement du H2SO4), des dégagements dhydrogènes peuvent aussi êtres à prévoir et demander une protection Or (je ne sais pas a quel niveau du process ils sont). Par experience, Total utilise nos séparateurs sur ces applications depuis longtemps dans problème avec le joint (vu avec Jérôme) Best, Nicolas
INC284133 | Workgroup | Incident | VIS THERMOSTATS PRESSOSTATS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-24 10:12 | 2023-10-24 14:58 | 1 | 0 |  | Bonjour Veronique, les 3 vis ne sont pas fournies, le client doit enlever la plaque, la fixer au mur avec les vis et taraudages (le cas échéant) de son choix puis replacer le ZPN sur la plaque. Best, Nicolas
INC284129 | Workgroup | Incident | TR: Differential quote | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-24 09:06 | 2023-10-24 09:18 | 0 | 0 |  | Nathalie, c'est le meme qu'on a redéfini il y a peu de temps, le cadran existe et ne sera pas à recréer: Order code: MFT7-A00.XxxN.62/9712/9704 MFT old Code / SAP code: Nominal size: 150 mm 7 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection, 3 back lugs fixing A Type of case: Case without partition of safety 0 Oil Filling: no oil filling 0 Pressure unit: Special pressure units X Pressure range: Pressure range for mBar Double scale 0/200 m3/hr S.R. - 0/0,05 Mpa xx Static pressure: 100 Bar N Process connection: AISI 316L (1.4404), stainless steel 1/2" NPT male 62 Best, Nicolas
INC284124 | Workgroup | Incident | TR: Quote request | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-23 17:21 | 2023-10-23 18:07 | 0 | 0 |  | Nathalie, voici le code que je proposerais pour sa demande (certificat matière à compter à part): si il y a un peu de quantité il sera opportun de proposer plutôt un D820 avec un anneau (a revoir au besoin). MEP8D60B20/0078/9941 + Order code D438-L100-N-20-226G0L-0A033-09 D438 Top housing material ►Hot rolled NACE compliant 1,4404 stainless steel EN10088-3 (table A4) L Stainless Steel Capillary type with Direct mounting 1 Outlet position ►Axial outlet 0 Capillary length No capillary 0 Connection on measuring instrument 1/2 NPT female for pressure gauge N Diaphragm material ►Stainless steel 1,4435 2 Diapraghm coating ►Without 0 Flange standard ANSI B16-5 2 PN / DN Class 300 / NPS 1"1/2 26 Flange facing ►Raised face 1,6 (RF) G Flange face finish ►Standard finish 0 Flange material Hot rolled NACE compliant 1.4404Stainless steel EN 10088-3 L Bottom housing coating Without 0 Bolting ►ASME (UNC) A Tracing circuit ►Without 0 Drain hole (bottom hole) 1/2NPT with plug 3 Vent hole (top hole) 1/2NPT with plug 3 Oil Filling LRS9 (-40 +400C) Syltherm 800 09 Best, Nicolas
INC284123 | Workgroup | Incident | BAPL ¦ N CDE/N Order: 651222/10 ¦ Oferta na czujnik ciśnienia | Archive > BHFR Technical Support | Middle |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-23 15:54 | 2023-10-23 16:58 | 0 | 0 |  | Hi Anna, here is the code you ordered on the mentioned order. please contact the sales dept for price and lead time: 11009796 Y91-3.3.3.B77.R/2151 Best, Nicolas
INC284111 | Workgroup | Incident | PERENITE MA7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-23 11:06 | 2023-10-23 15:25 | 0 | 0 |  | Bonjour Véronique, les MA7 sont d'actualité, c'est notre seul mano absolu et il est basé en grande partie sur le MX7 donc pas de problème de pérennité pour l'instant. Best, Nicolas
INC284110 | Workgroup | Incident | EU declaration of conformity | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-23 10:39 | 2023-10-23 15:29 | 1 | 0 |  | Hi Rudi, I have below CE conformity declaration but it's only for instruments: [Image] I can send it on request. if it's not on this list, then please specify to which EN/EU norm/approval you need the conformity. Best, Nicolas
INC284101 | Workgroup | Incident | Mounting | Archive > BHFR Technical Support | Middle |  |  | Closed | Stojanovic, Zoran [stoz] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-20 15:36 | 2023-10-23 09:14 | 0 | 0 |  | Hi Zoran, it can be mounted from a vertical position down to the almost horizontal position (0 vertical to 87 almost horizontal). but honestly, the last 3 caused the drift due to the position to be slightly over what is tolerated but nothing making it unusable. Best, Nicolas
INC284102 | Workgroup | Incident | INFOS RTN4 30182615 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-20 15:36 | 2023-10-20 17:11 | 0 | 0 |  | Salut Veronique, en cas de rupture hautement improbable du capillaire (ils sont résistants en pression a plus de 1600bar), il y aura une perte en charge qui dépendra de la fuite. la pression que recevra le switch sera donc inférieur à la pression réelle (plus ou moins selon la fuite). sur un système clos la pression va descendre jusquà 0 et si elle passe le point de contact, il switchera. sur un système ouvert avec une pression d'arrivée constante la fuite pourras créer une fausse pression plus basse que la réelle jusquà reparation de la fuite, si cette pression descend en dessous du setp, l'appareil switchera. Best, Nicolas
INC284097 | Workgroup | Incident | Connection piece | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-20 13:21 | 2023-10-20 14:07 | 0 | 0 |  | Hi Paul, it exists with NPT but not G on the other side. but what you need is an ARIN2X0L (M20x150 F) you can ask for a price to sales and precise the qty (the price will be based on it). Best, Nicolas
INC284096 | Workgroup | Incident | MMD3 Datasheet | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-20 12:36 | 2023-10-20 17:05 | 1 | 0 |  | Hi Daniel, please find enclosed the draft of the MMD3 datasheet, the contant is ok but the drawings don't look nice. Best, Nicolas
INC284095 | Workgroup | Incident | AGWR | Archive > BHFR Technical Support | Middle |  |  | Closed | Stojanovic, Zoran [stoz] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-20 11:42 | 2023-10-23 09:28 | 1 | 0 |  | Hi Zoran, we can make thermowells in exotic materials but usually we do it on the AGFR because it's easier to find the material and have it machined with our partner than to find specific thickness exotic metal sheets to build an AGWR. you can build your AGFR in the P2F and specify an X for the material (adding the note it's for duplex and send the price request to your sales contact). don't forget the qties, it ca dramatically change the price. Best, Nicolas
INC284085 | Workgroup | Incident | rfi MR5/MS5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-19 19:54 | 2023-10-20 10:59 | 0 | 0 |  | Hi Angel, our products are not tested according to those norms. we respect the RCCM/E and some of our products are K3 approved which is the european (ALSTOM/EF) Nuclear standard. as we were never consulted for US designed plants, we never went through those qualifications. Best, Nicolas
INC284084 | Workgroup | Incident | RFI Doubt about MIX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-19 17:57 | 2023-10-23 10:15 | 0 | 0 |  | Hi Angel, I think you confuse accuracy and resolution: Accuracy is the comparison between the value read on the gauge and the actual pressure value, but you can read between the increments (if the pointer is between 3.5 and 4bar increments, you can read 3.7 even if there is no mark for it). Resolution is the number of increments on a dial. they're of course related but there is also a compromise to do to get the best readability (reason why the resolution also depends on the size of the dial). last thing to add is that the studies done about this shows that there is a point where you have too much increment and even if it make the gauge look like it's really accurate, it leads to false reading or increase the time to define the real value because of a difficult reading. Our dial ensure to respect the accuracy class and a confortable reading according to the EN837. Best, Nicolas
INC284082 | Workgroup | Incident | TR: [Chung Chia Project] - RFQ for Pressure Gauges/ Ovivo SQ12301897 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-19 16:12 | 2023-10-19 17:14 | 0 | 0 |  | Nathalie, il faut faire une demande aux achats, normalement les fournisseurs de ce type de séparateurs peuvent le faire mais je ne sais pas quel sera notre prix. Best, Nicolas
INC284081 | Workgroup | Incident | MEX8-463.B74/0678 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-19 16:03 | 2023-10-19 17:25 | 0 | 0 |  | Hi Kerstin, the P2F is right, with the small pressure and the filling this pointer is not possible. but you can choose the 0679 which is the friction adjustable pointer (just another technology for the same use) Best, Nicolas
INC284053 | Workgroup | Incident | RPPN Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-18 11:39 | 2023-10-18 14:11 | 0 | 0 |  | Hi Paul, it's an old type of code some customers had there own "option code" containing a serie of option. here is the recent codes matching the customer code RPPEA6209PIL2: RPPN-7E6.209/1990_9941_SETP_Q1061 11105076 RPPN-7E6.209/1990_9941_SETP_Q1061 11105077 RPPN-7E6.209/1990_9941_SETP_Q1061 11105527 only the setpoints are different. we don't include the Q1061 (setp calibration certificate) anymore, it's now in the order. the new code would be: RPPN-7E6.209/1990_9941_SETP (setp to define). Best, Nicolas
INC284039 | Workgroup | Incident | RE: Zapytanie. | Archive > (I) Out of Material demand | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Dolbeau, Nadia [ndo] | 2023-10-17 15:21 | 2023-10-19 13:55 | 1 | 0 |  | Anne pr les demandes techniques c'est Nicolas Birrien : FUNC BHFR Bourdon Technical Support <bourdontechnicalsupport@baumer.com> merci de faire un ticket à cette adresse mail et modifier l'info svp . Nadia
INC284038 | Workgroup | Incident | TR: Zapytanie. | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-17 15:18 | 2023-10-17 16:51 | 0 | 0 |  | Anne, je peux te dire ce que c'etait mais je ne pense pas qu'on ai d'equivalent pour ce type de montage. ci-dessous quelques details et je conseillerais de contacter Baumer pour une alternative en TFRH. [Image] Best, Nicolas
INC284034 | Workgroup | Incident | BAPL ¦ TBX080 ¦ RFQ Baumer | Archive > BHFR Technical Support | Middle |  |  | Closed | Galusińska, Estera [gae] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-17 15:00 | 2023-10-17 17:51 | 0 | 0 |  | Hi Estera, the TBX with it's conical shaped immersion tube is really specific in term of length and sizes to be mounted with the thermowells designed for it. if the application requires specific length and connection, I would go more toward a TB 080 or a TBI 080 where you'll be able to choose different lengths and connections. Sélecteur ¦ Thermomètres bimétalliques ¦ Bourdon Instruments (bourdon-instruments.com) "https://www.bourdon-instruments.com/fr/fr/apercu-des-produits/thermometres/thermometres-bimetalliques/c/38229/products" Best, Nicolas
INC284025 | User | Incident | Ajout du R dans la désignation des MEP5R | Archive > BHFR Technical Support | Low |  |  | Closed | Pletikos, Maurane [plem] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-10-17 10:36 | 2023-10-19 16:39 | 1 | 0 |  | Mauranne, peux tu envoyer la demande a Material data, c'est arrivé chez moi et le transfer de ticket entre le ticketing de Nadia et le mien bug. Best, Nicolas
INC284019 | Workgroup | Incident | TR: quick pricing for D920 please | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-17 09:18 | 2023-10-17 09:54 | 0 | 0 |  | Handled via Nathalie
INC284012 | Workgroup | Incident | TR: Prix d'achat Y91-3.3.3.B24.R + 1622 + LRS1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-17 08:27 | 2023-10-17 09:56 | 0 | 0 |  | Nathalie, oui c'est OK: [Image] Best, Nicolas
INC284009 | Workgroup | Incident | REGLAGE MFT | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-16 17:15 | 2023-10-17 16:42 | 0 | 0 |  | Réponse Fantinelli + extrait de l'EN837 envoyé via email. Best, Nicolas
INC284007 | Workgroup | Incident | Nestle Gauge and Seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-16 16:18 | 2023-10-16 16:45 | 0 | 0 |  | Hi Daniel, it was closed because not sold since 2016, I don't see other reason so you can ask for the price as is and we will probably open the code as it is. Best, Nicolas
INC283993 | Workgroup | Incident | MG5 safety pattern? | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-13 17:51 | 2023-10-16 16:47 | 1 | 0 |  | Hi Paul, you can have the ATEX on an MG5 (inductive contacts) with the option 0751 and 0078 (triplex mandatory). but we don't have it with safety pattern. Best, Nicolas
INC283979 | Workgroup | Incident | WG: Customer request: product | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-13 07:51 | 2023-10-13 12:03 | 0 | 0 |  | Hi Felix, we don't have it and in fact you're the first one to ask for it. I'll ask Baumer china for their opinion on this certification to help us decide if it is a necessity for the future. Best, Nicolas
INC283978 | Workgroup | Incident | Switch enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-12 20:45 | 2023-10-13 09:21 | 0 | 0 |  | Hi Paul, yes, I defined the code with him as a replacement of the picture one. the G is existing but it's the special microswitch with 2 separates SETP and require our approval and setting as there is many complicated restrictions. that's what the customer had in place and we can still offer it but needs the details. Otherwise, you can check with him if a simpler switch with double setp at the same point could be used. Best, Nicolas
INC283972 | Workgroup | Incident | TSSE à valider | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-12 10:42 | 2023-10-12 14:06 | 0 | 0 |  | Hi Cyrille, je ne comprends pas, tu peux juste le faire au P2F, il n'y a rien a valider: verifie juste la connection, il manquait la lettre, en standard je t'ai mis du G1/2 Order code: TSSE100-211.166.12T.5001/738C TSSE old Code / SAP code: Nominal Size: DN 100 100 Case and bezel ring: AISI 304 (1.4301), stainless steel 2 Type of mounting: Direct monting, immersion tube bottom 11 Immersion tube diameter: 8 mm 1 Immersion tube material: AISI 321 (1.4541), stainless steel 6 Immersion length: 250 mm 6 Temperature range: Single scale inC 0 ... 100C 12T Contact function (with clockwise pointer motion): Mechanical contacts Contact makes (NO) 5001 OPTIONS Process connections: Adjustable connection sliding on stem G1/2" 738C Best, Nicolas
INC283969 | Workgroup | Incident | Complaint from Reliance - PG with Saddle Flange | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-12 09:30 | 2023-10-18 17:22 | 0 | 1 |  | Hi Anil, we finalized the check internally and we think the capacity of the seal in term of pressure might have been over accepted on this one. I cannot go further as the next steps will be to either block the product, or ask for a study to find an alternative solution and we might have to evaluate risks. Please create an official claim on this so that it can be pushed to the next decisions and actions. Best, Nicolas
INC283963 | Workgroup | Incident | NACE cert | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-11 22:48 | 2023-10-12 14:02 | 1 | 0 |  | Hi Paul, you're right but the parts are coming from india for those standard monoobloc seals. I checked the supply and the certificates, the good thing is you can order the 3.1 certificate and the bottom part and diaphragm materials are on it. the bad news is only the diaphragm is notified in the certificate as NACE. we could issue a conformity certificate 2.2 NACE for the diaphragm. If it is necessary to have it for all parts in contact, I think it might be possible but we have to go through an SPR to either change the material of the bottom part of talk to the supplier in india for the material o see if he can NACE certify it. Best, Nicolas
INC283947 | Workgroup | Incident | WG: Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-11 08:15 | 2023-10-11 17:39 | 0 | 0 |  | Hi Felix, I guess this one is old and was not a nuclear one, normally it should be replaced by this: RPPN7G3209 Microswitch type G is the only one with 2 separated setp, it is very special and we need the request from the customer in term of SETP and resetpoint for both switches in order to validate it is possible. Best, Nicolas
INC283942 | Workgroup | Incident | EQUIVALENCE DP370 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 15:39 | 2023-10-11 09:18 | 0 | 0 |  | Bonjour Veronique, voici léquivalent le cadran est existant donc pas à recréer: Order code: MFT7-A00.XxxN.62/9712/9704 MFT old Code / SAP code: Nominal size: 150 mm 7 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection, 3 back lugs fixing A Type of case: Case without partition of safety 0 Oil Filling: no oil filling 0 Pressure unit: Special pressure units X Pressure range: Pressure range for mBar Double scale 0/200 m3/hr S.R. - 0/0,05 Mpa xx Static pressure: 100 Bar N Process connection: AISI 316L (1.4404), stainless steel 1/2" NPT male 62 Best, Nicolas
INC283941 | Workgroup | Incident | # 11259900 FOURN.CLIENT + DANC01.2L2A.01/0400_9912 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 15:27 | 2023-10-10 16:54 | 0 | 0 |  | Hi Sascha, I checked with R&D and even though we mounted some of those danfoss in the past, the code should still not exist without details. Please send the complete code of what the customer will provide and I'll validate the mounting and ask for the creation of a new code and the closing of the old one. Best, Nicolas
INC283940 | Workgroup | Incident | ZT-Y405SCM210/SETP_2237 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 15:21 | 2023-10-10 15:43 | 0 | 0 |  | Bonjour Véronique, oui c'est bon, il est bien ouvert dans SAP, il avait été fermé en 2021 pour le changement du fluide mais ré-ouvert sur approbation de Mr Vrignaut en septembre 2022. c'est bon pour nous. Best, Nicolas
INC283935 | Workgroup | Incident | RE: MG 5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 12:33 | 2023-10-10 13:58 | 0 | 0 |  | Anne, la connection electrique d'un MG5 se fait par un cable qui vient se brancher sur un terminal a travers un presse etoupe M20x150 et cela en standard. Best, Nicolas
INC283933 | Workgroup | Incident | stp. files | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 12:27 | 2023-10-10 14:49 | 0 | 0 |  | Hi Sascha, Sorry but with those old German codes it's very difficult to find anything, I managed to get you the most possible information on each codes. Also if you can switch some to our standard products, it would be great (ASIP, ARP). SAP Code Designation Details Drawing Datasheet Replacement 10326559 S'HÜLSE NG6 LS=60 G1/2"INOX FOR BULB D=6 MM; G 1/2 STAINLESS STEEL; LS= 60 MM T 09346.0300 T6.210 OK 10296203 W.ROHR DIN16282-AG-EDELSTAHL N05013.0001 SIPHON TUBE G 1/2 MALE/TIGHTENING COUPLING; U-TYPE, STAINLESS STEEL N05013.0001 D5.205 ASIP 10296204 W.ROHR DIN 16282-BG-STAHL SIPHON TUBE G 1/2 WELD.ENDE/TIGHTENING COUPLING; U-TYPE, STEEL N 05013.0002 D5.205 ASIP 10326421 EINSCHRAUBHÜ. G1/2 INOX LS 48 / LT 55MM Screw in pocket; FOR BULB D=8 MM; G 1/2 STAINLESS STEEL; LENGTH TOTALLY LT = 55 MM; LS= 48 MM T 09143.0005 TBH100 OK 10325852 VENTIL DIN 16270-AG-EDELSTAHL Valve; G 1/2 MALE/ADJUSTING NUT; STAINLESS STEEL S004.16.200 AS-Schneider - Gauge Valves DIN 16270 Type A "https://www.as-schneider.com/gb/en/Products/Instrumentation-Valves-and-Manifolds/Gauge-Valves-and-Pressure-Gauge-Accessories/Gauge-Valves-Without-Test-Connection/Gauge-Valves-DIN-16270-Type-A.html" ARP-A03L "https://www.bourdon-instruments.com/fr/fr/apercu-des-produits/accessoires-de-mesure-de-pression/robinets-/arp-a03l/p/37845" Best, Nicolas
INC283932 | Workgroup | Incident | TR: Price | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 12:12 | 2023-10-10 14:02 | 0 | 0 |  | Oui, le montage est OK, Best, Nicolas
INC283930 | Workgroup | Incident | PFMN sur séparateur | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-10 10:57 | 2023-10-10 11:57 | 0 | 0 |  | Bonjour Cyrille, On a pas de fluide qui permet de résister à 420C, le maxi est de 400C. à part ca pour le montage, le ZPX ne devrait pas etre nécessaire, un grand ASIPD ou un capillaire seront suffisant. pour le montage murale il faudrait voir mais je pense qu'on peut se prendre sur le capillaire ou la connection du PFMN. Le probleme reste en tout cas la temperature de toute facon on ne pourra pas avancer avec plus de 400C. Best, Nicolas
INC283913 | Workgroup | Incident | TR: OFFRE ACTEMIUM DDP URGENT 20539383 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 15:12 | 2023-10-09 15:52 | 0 | 0 |  | Nathalie, ok pour tout sauf pour le remplissage silicone sur le mano degraissé O2, a remplacer par du BH5. Best, Nicolas
INC283911 | Workgroup | Incident | Neobee Fill Fluid | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 13:03 | 2023-10-09 15:59 | 0 | 0 |  | You're right, Neobee rings a bell, but it seems to simply be a glycerin oil for FDA applications and we can answer with either LRS1 or 10, the MSDS are enclosed. Best, Nicolas
INC283910 | Workgroup | Incident | MEX-D31.B20 + DAEL 02.0L7 (Tetrapak) | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 12:36 | 2023-10-10 17:33 | 0 | 0 |  | Hi Daniel, yes it doesn't help uncoding what they asked for but anyway the code they sent can't exist, I would offer them the most standard mounting according to what they asked and purchased in the past and ask them to check: MEX5D31B20 + DAEL11.2.L23.01 and send the datasheet for reference and confirmation. Best, Nicolas
INC283908 | Workgroup | Incident | Dial question | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 11:12 | 2023-10-09 16:18 | 0 | 0 |  | Hi Paul, hereafter the 2 dials for MEP5/0843 based on 11263875: [Image] Based on 11218455: [Image] Best, Nicolas
INC283907 | Workgroup | Incident | Demande technique / SIFEL / Marquage spé MFE | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 11:03 | 2023-10-09 16:44 | 0 | 0 |  | Bonjour Greg, oui c'est faisable, c'est un marquage sur le cache avant et c'est possible sur les rétroéclairés (qui ont un cache) le marquage de l'unité et du nom de la voie ou des 2 voies est habituel mais les 3 marquages demandés n'existent pas, il faudra donc créer ce marquage (frais probables idem creation de cadran, valider avec le commerce). il serait bon de preciser aussi les couleurs de chaque marquage. exemple: MFE422020JB2 marquage: bar (blanc) CG / RE (Blanc /jaune) Best, Nicolas
INC283901 | Workgroup | Incident | 3D drawing of TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 09:36 | 2023-10-09 09:55 | 0 | 0 |  | Hello Zhang Wei, no problem, please find enclosed the step file of a standard TED. Best, Nicolas
INC283898 | Workgroup | Incident | Replacement Model For E93318B22R / Gilang | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-09 08:51 | 2023-10-09 09:10 | 0 | 0 |  | Hi Bee Tin, here is the replacement you can offer: Y91333B22R + DANC012L2A01/9010/0592 (+/-1%) or a PBMH you can configure on the website Best, Nicolas
INC283882 | Workgroup | Incident | TR: Demande de prix IC | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-06 15:45 | 2023-10-06 16:16 | 0 | 0 |  | Nathalie, les montages sont ok. Best, Nicolas
INC283880 | Workgroup | Incident | DEFINITION MFT7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-06 14:18 | 2023-10-06 15:26 | 0 | 0 |  | Veronique, ma proposition serait ceci: MFT7-D00.X10Q.62 0mbar = 250mbar = contact à definir au besoin avec le detail du besoin client (alim, intensité ou pouvoir de coupure type de contact) si contact, un declassement ou une echelle reduite est à prévoir et la connection sera un PE standard. Best Nicolas
INC283879 | Workgroup | Incident | TR: BAIT ¦ IMA Fwd: Richiesta chiarimenti manometro con trasduttore di pressione integrato cod. IMA 6250115 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-06 14:18 | 2023-10-06 15:58 | 0 | 0 |  | Hi Anne, the code 11197645 will be replaced to phase out the Clamp diaphragm seal from KHW and replace it with ours. the functionnality and connection will be the same, just the design of the seal is slightly different. please offer a : DRCG101/211.133.535/3272 + DANC01L2B01 the body of the seal will be electropolished to 0.4microns and the complete seal will never exceed 0.8microns (conform 3A) Best, Nicolas
INC283877 | User | Incident | Création BAFR // Client : 50110954 // Cde Clt : 20539218 // ZT--403SHM220/9920_9922 // Position 50 | Archive > BHFR Technical Support | Low |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-10-06 13:54 | 2023-10-11 17:48 | 1 | 0 |  | Pour mat data
INC283875 | Workgroup | Incident | ZERO MX7 / ASG-ND Lot C : Seg803LP | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-06 11:15 | 2023-10-06 11:22 | 0 | 0 |  | Salut Véronique, Si c'est vraiment au moment ou il se connecte a son systeme, alors ca ne vient pas du mano mais de son systeme et il peut y avoir plusieurs explications: un montage non vertical peut entrainer un leger décalage selon les pressions. une pression residuelle dans ses canalisations le montage lui meme sur un faible volume statique peut aussi entrainer une legere augmentation de pression. mais cette pression lue est reelle (sauf dans le cas du montage non vertical) et c'est une lecture de l'etat de la pression dans ses canalisations donc il ne faut rien toucher, lors de la mise en service cette pression sera inclue dans la lecture. Best, Nicolas
INC283873 | Workgroup | Incident | WG: CAD-Daten | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-06 10:42 | 2023-10-06 11:24 | 0 | 0 |  | Sure, here is the step file: Best Nicolas
INC283866 | Workgroup | Incident | 11254108 FDA | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-05 15:39 | 2023-10-05 16:13 | 0 | 0 |  | Hello, please find enclosed the FDA approval for the LRS1 filling liquid. Best, Nicolas
INC283857 | Workgroup | Incident | Material Description (Textcode) ? - QuantiParts (27100130) - SAP 10264049 - AGF SVT PL 39741-1 L=83MM M27X200 ETIQIX | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-05 11:09 | 2023-10-05 11:32 | 0 | 0 |  | Hi Rudi, when we do a textcode with a drawing n it's because it can't be described by the datasheet but here is the drawing with the complete information. Best, Nicolas
INC283838 | Workgroup | Incident | TR: DT3 & DT2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-04 15:57 | 2023-10-05 09:25 | 0 | 0 |  | Hi Angel, the Tantalum was never available for machined part on any of our seals. the reason is it's difficult to find in bars and really hard to machine (which means really expensive IF we find someone to do it). I suggest to find another material for the massive parts, I'm sure you'll find something that resist and keep in mind that finding the perfect material on diaphragm is really important as it's thin and so subject to corrosion damages, the massive parts are much less. Best, Nicolas
INC283826 | Workgroup | Incident | DP370 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-04 10:00 | 2023-10-04 11:24 | 0 | 0 |  | Bonjour Véronique, tu as raison, il faut commander un DP370, mais à priori pas de problème comme c'est pour un client qui en a déjà un. le mieux est de demander la création du code (si vous avez la commande) sinon de demander le prix de ce code: DP370 BP V39.43.51.58, NS:150, 1/2NPTM, 0-200m3/h, 0-0.05Mpa en cas de creation de code, donne l'exemple du code 11257828 Best, Nicolas
INC283824 | Workgroup | Incident | DANC01-2L6A | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-04 09:39 | 2023-10-04 11:28 | 0 | 0 |  | Hi Andy, you're absolutely right, only the diaphragm will be HC276, the massive part is 316L. but we do that only on the diaphragm because ist's much thinner and though more subject to corrosion corroding importantly the massive part will be much longer. Best, Nicolas
INC283819 | Workgroup | Incident | WG: MEx 5 with D821 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-04 08:54 | 2023-10-04 11:39 | 0 | 0 |  | Hi Felix, I agree with the customer, the PTFE would be the best option if he have no overpressure or temperature (10bar / 100C max) instead of: D821-10L.42DE.B60 I would suggest: D425_2L2.24DE.B23 (11207436) Best, Nicolas
INC283815 | Workgroup | Incident | Y91 : perméation de l'hydrogène au travers du joint et de la cellule | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 16:51 | 2023-10-03 17:23 | 0 | 0 |  | Bonjour Cyrille, Je demande depuis longtemps la mise en place de choses pour l'hydrogène mais pour l'instant peu ont vu le jour. Pour l'instant aucun test avec de l'hydrogène n'a été fait (en meme temps c'est un peu compliqué en termes de moyens, il faut passer par un labo spécialisé). par contre, nous avons le Y92 (membrane flush) avec membrane dorée qui est fait specialement pour une utilisation sur H2. voici un exemple de code, toutes les demandes de creation doivent passer par Franck mais c'est assez rapide. 11253447 Y92-3.5G.B22.R/2164_7912 sinon, si c'est une demande globale pour air liquide, nous sommes ouverts à la discussion. Best, Nicolas
INC283813 | Workgroup | Incident | PFMN sur séparateur hygiénique | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 16:15 | 2023-10-05 09:28 | 0 | 0 |  | Bonjour Cyrille, ton montage est possible: PFMN-54.BC3R.C1581.33200.2001_DAVA20-2L6A01_9010_1996 Best, Nicolas
INC283804 | Workgroup | Incident | URGENT PROBLEME SOUFFLETS MX7 NAVAL GROUP | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 11:57 | 2023-10-03 15:19 | 0 | 0 |  | Bonjour Véronique, je pense comprendre que tu cherche une solution alternative ur un sujet en cours avec quelqu'un de la R&D, mais dans e cas que tu me propose, je sais que je n'ai pas toutes les informations et je vais de toute facon me retourner vers philippe ou nicolas pour faire le point car ils ont la vue d'ensemble sur ce sujet particulier. Je conseillerais plutot de faire remonter ton blocage hierarchiquement pour debloquer directement la situation. Best, Nicolas
INC283801 | Workgroup | Incident | TR: RFQ FIRM BUYING - 3000016206 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 11:45 | 2023-10-03 15:15 | 0 | 0 |  | Bonjour Anne, C'est ce code: 11246088 vendu à Saad al alyan [Image] Best, Nicolas
INC283797 | Workgroup | Incident | MEX5 low temp. | Archive > BHFR Technical Support | Middle |  |  | Closed | Skeldon, Paul [skp] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 09:30 | 2023-10-03 16:35 | 0 | 0 |  | Hi Paul, or now the standard use of MEX5 to -40C is still under test. but I can offer you the MEX5 + BH4 filling as an option to answer the request you have. [Image] Best, Nicolas
INC283796 | Workgroup | Incident | DRS EN 1092-1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 08:57 | 2023-10-03 09:22 | 0 | 0 |  | Salut Véronique, je vois que c'est pour du vide à priori et avec un séparateur vraiment hors standard. il nous faudra la plage de mesure de pression et aussi une SPR pour valider le montage stp. Best, Nicolas
INC283793 | Workgroup | Incident | FW: Price for Model Diaphragm seal D820-A6N.2K67.Q20-LRS9 | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-03 07:27 | 2023-10-03 09:24 | 0 | 0 |  | Hi Anil, sure, here are the details: [Image] Best, Nicolas
INC283790 | Workgroup | Incident | MEX3-BX2.A15/0151 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-02 15:00 | 2023-10-02 15:44 | 0 | 0 |  | Veronique, le MAR3 avait une conception legerement differente du MEX3 qui permettait ed faire de la connection en M20*150. le MAR3 est phase out et nous ne pouvons plus le faire, il est remplacé a 90% par le MEX3 avec les options appropriées (peinture noire etc...) mais nous ne pouvons plus le faire en M20*150. ll faudra donc soit proposer une autre connection ou un adaptateur, soit passer sur un MEX5 pour avoir un carré assez grand pour cette connexion. Best, Nicolas
INC283787 | Workgroup | Incident | TR: Chiffrage | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-10-02 11:48 | 2023-10-02 14:53 | 0 | 0 |  | Bonjour Jacques, ce que je suis sur de pouvoir te proposer pour répondre au client c'est: PBMN-24B10RA11403201000 + DT3-2L62.0.233.01 globalement les choix possibles avec l'acide fluorhydrique semblent etre l'HC276, le PTFE et le Viton. ne alternative est donc de prendre la bottom part en PTFE directement. sinon si tu as un grand nombre de pièces, je pense qu'il serait interessant d'étudier d'autres options (un D04x en HC276 avec un Y91) qui seraient moins cheres mais possibles seulement en grande quantité car il faudrait faire des essais. Dernier point, le client est responsable du choix des materiaux au final car c'est le meilleur expert de son process, nous ne pouvons que conseiller en nous basant sur des données limitées. Best, Nicolas
INC283751 | Workgroup | Incident | TR: RFQ/SAP CODE D825-10N-2217-GX0-D/1996 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-28 17:57 | 2023-09-29 09:25 | 0 | 0 |  | Nathalie, si ils ne prennent pas d'option pour la soudure, le capillaire sera seulement vissé. Best, Nicolas
INC283750 | Workgroup | Incident | TR: Demande de prix Sensotec Instruments ¦ AKPL4F5X05 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-28 17:39 | 2023-09-28 17:46 | 0 | 0 |  | Nathalie, il faut quand meme passer par la R&D mais normalement ca ne coutera pas plus cher qu'un M/F car on soude les raccords de la meme facon des 2 cotés donc ils devraient etre interchangeable. Best, Nicolas
INC283739 | User | Incident | 11253680 demande de blocage du code | Archive > BHFR Technical Support | Low |  |  | Closed | Pichon, Sylvain [psy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-09-28 11:00 | 2023-09-28 14:00 | 1 | 0 |  | Traité par Func Material data
INC283734 | Workgroup | Incident | TR: RFQ-MANOMETER WALL SUPPORT_Bourdon Haenni | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-28 08:57 | 2023-09-28 09:18 | 0 | 0 |  | Nathalie, c'est ce code SAP: 10296198 Best, Nicolas
INC283710 | Workgroup | Incident | WG: EIL - Anfrage-Nr.: 01/20230707 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-27 09:00 | 2023-09-27 09:11 | 0 | 0 |  | Hi Felix, I don't know this DIN norm for the connection but I would start by offering this and ask the customer to really check the connection: PBMH-24B79RA14453522000 Best, Nicolas
INC283706 | Workgroup | Incident | MMN5-460.D22 (10292388) / Christine | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-27 05:03 | 2023-09-27 17:04 | 0 | 0 |  | Hi Bee Tin, yes, it'll still be the old case design for the MMN5. Best, Nicolas
INC283701 | Workgroup | Incident | WG: 23 09 26 FontM - Anfrage CAD- 3D-Model - DE Wehrheim - Angebots-Nr.: 20520199 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-26 11:18 | 2023-09-26 12:02 | 0 | 0 |  | Hi Felix, you're lucky on this one, here it is: Best, Nicolas
INC283700 | Workgroup | Incident | TR: Demande de prix SPR / Achats en Nord | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-26 10:54 | 2023-09-29 09:32 | 0 | 1 |  | Nathalie, le MMD4 ne doit pas encore exister, pour l'instant il faut le creer en DR80 et il sera plus cher qu'un MEX3/5, je ne le conseille pas si cest pour une SPR, voici le code que je mettrais (pas sur que ce soit accepté). DRO80 411 133 065 9695 / 9696 / 9695 / 9750 / 0765 / 0751 (option remplissage O2 non existante). Best, Nicolas
INC283671 | Workgroup | Incident | WG: rvt. Zeichungsdatei Manometer | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-25 10:35 | 2023-09-25 10:58 | 0 | 0 |  | Hi Felix, we don't use revit for the design but we have standard .stp files they should be able to open (see enclosed). Best, Nicolas
INC283665 | Workgroup | Incident | TR: RFQ EA5Y-F0-H30R | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-25 10:35 | 2023-09-25 10:56 | 0 | 0 |  | Nathalie, on n'aura plus l'atex, mais la SU peut demander à Franck l'autorisation d'ouvrir ce code si ca va au client sans l'ATEX: 11100867 Best, Nicolas
INC283660 | Workgroup | Incident | stainless steel movement for MCX | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-24 16:30 | 2023-10-11 09:48 | 0 | 1 |  | Hi Xu Qiang, very sorry for the long delay, the answer is the datasheet is right, below 16mbar, we cannot have the stainless mouvement. so 0651 is not compatible with ranges 01, 02, 03, 48, 49 et 60 Best, Nicolas
INC283644 | Workgroup | Incident | MEP5...... ? - Global Enery - External TotalEnergies | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-21 15:21 | 2023-09-21 15:28 | 0 | 0 |  | Hi Rudi, not on this one, but I see an ATEX marking on the dial which means you have the order n and position in this marking. Best, Nicolas
INC283638 | Workgroup | Incident | BAIT ¦ Fluid processing CO2 pressure measurement | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-21 13:00 | 2023-09-22 09:15 | 0 | 0 |  | Ok, sorry, franck commented on the ticket, I thought you could see that. you can work with this with a transmitter if you add a capillary (at least 1m or a small ASIPD). it should reduce the temperature enough to work if the capillary or ASIPD is getting the transmitter to a warmer environment, otherwise, 2m min or a big asipd and ambiant temperature for the body of the transmitter needs to be respected. Best, Nicolas
INC283637 | Workgroup | Incident | TR: Demande de prix SPR / SNCF 37 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-21 12:36 | 2023-09-21 15:16 | 0 | 0 |  | Nathalie, Aurélien, Réponse un peu compliquée, je vais essayer dêtre clair. Aurélien, tu me rappel au besoin. Déjà ce code est bloqué pour un problème de composant donc la SPR s'impose de toute façon. Ensuite d'un point de vue purement technique un remplissage est sûrement possible (boîtier/connection idem MEX3). Par contre, faire un remplissage donnera un produit non conforme au plan d'origine (59190-01) et a la SP2576 SNCF/Baumer car ces documents ne mentionnent pas de remplissage et font état d'un poids de 130g ce qui ne sera plus le cas avec le remplissage. le client doit donc être conscient que cette demande provoquera chez lui et chez nous la création ou la revision de ces documents et un circuit d'approbation complet peut-être avec des tests à réaliser (procedure qui peut être lourde pour le client aussi et entraîner des coûts). Je conseil donc d'en parler au client avant de lancer cette SPR pour être sur que tout le monde soit sur la meme longueur d'onde. (et probablement de refaire la SPR sans le BH) Best, Nicolas
INC283631 | Workgroup | Incident | FW: RFQ- BOURDON PRESSURE SWITCH & CAPILLARY | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-21 10:51 | 2023-09-21 14:56 | 0 | 0 |  | Hi Anil, this was produced in January 2005 and is too old to be in the system. hereafter the replacement for the switch and the capillary, and I suggest a DT2 for the seal but it seems parts were in monel and you need to check this out with the customer to define the right bottom part and diaphragm material (or coating ?). RPPN-7K6.204/SETP (min 1.5bar, to be defined) + AKPLB0N605 + DT2-2Nx0.0.x6x.01 Best, Nicolas
INC283618 | Workgroup | Incident | TR: Question microprint | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-20 12:24 | 2023-09-25 16:50 | 1 | 0 |  | Nathalie, notre butée est bien placée de la meme façon que sur la photo sur nos MEX3 (sous le 0 en fin déchelle). Best, Nicolas
INC283613 | Workgroup | Incident | TR: TRES TRES URGENT Demande de prix AO 1693954 | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-20 10:57 | 2023-09-25 16:52 | 0 | 0 |  | Aurelien, voici la fiche d'obsolescence pour le mano demandé. Best, Nicolas
INC283609 | Workgroup | Incident | FW: RFQ: Steering Gear Room- Level Gauge_XAO HE202309145 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-20 06:39 | 2023-09-21 10:05 | 0 | 0 |  | Double, accepted on INC283608
INC283608 | Workgroup | Incident | FW: RFQ: Steering Gear Room- Level Gauge_XAO HE202309145 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-20 06:06 | 2023-10-05 09:30 | 0 | 1 |  | Hello Bee Tin, I close the ticket for now as we don't have answer from the customer. tell me if you get anything new. Best, Nicolas
INC283603 | Workgroup | Incident | TR: Demande de prix Kem One | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-19 14:12 | 2023-09-19 14:28 | 0 | 0 |  | Bonjour, un changement de joint n'est pas nécessaire sur ce montage (vu R&D) on garde le PTFE. Best, Nicolas
INC283593 | Workgroup | Incident | IP class and Explosion proof certificate | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-19 10:00 | 2023-09-19 11:23 | 0 | 0 |  | Dear Zhang Wei, for everything about ATEX, you find it on the CE certificate (I enclosed the one for Ex Proof products (RDE and ZDxE)) I also attached the IECEx cert. For IP I don't have a certificate, if necessary we need to do a check with quality
INC283587 | Workgroup | Incident | WG: 23506 Anfrage Manometer / Angebot 20532204 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-19 08:21 | 2023-09-19 15:08 | 1 | 0 |  | Hi Felix, here is for the MEP3 and for the MZ7. sorry for the MZ7 it's the only mounting type I have. Best, Nicolas
INC283584 | Workgroup | Incident | Demande de faisabilité / MR5+AKPL+D820 / BOSTIK | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-18 11:03 | 2023-09-19 15:15 | 1 | 0 |  | Salut Greg, la membrane du D820 va être trop petite pour faire marcher un différentiel a soufflet. sachant que le D821 ne te donnera pas de possibilité de matière exotique pour le flasque et que les D824/5 ne passent pas en taille de membrane, la meilleure option est de prendre un D435, avec la membrane HC276 et le flasque inférieur revêtu ou en matière exotique ou plastique etc... aussi le capillaire doit être de 2.5mm sur les montages diff. M31-4A3.B18H (2.5bar diff / 10bar static + contact simple) AKPLBF5301 (2.5mm for diff) D435-2L60.4FF.Bxx01 (PN25 / DN32 RF HC276 membrane) Best, Nicolas
INC283580 | Workgroup | Incident | SETP of ZT303 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-18 10:33 | 2023-09-19 14:45 | 1 | 0 |  | Dear Zhang wei, yes this is the standard label you can find inside your device for nuclear ones. it means it was set at 88C rising with a deadband of 8C (80C falling). Best, Nicolas
INC283574 | Workgroup | Incident | TR: enquiry Rolls-Royce - gaskets for Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-18 09:21 | 2023-09-19 11:13 | 1 | 0 |  | Hi Nathalie, here is the code and a picture of the EPDM gasket for G1/4 DIN3852-E 10303406 [Image] Best, Nicolas
INC283572 | Workgroup | Incident | DIAMETRE MEMBRANE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-18 08:30 | 2023-09-19 10:41 | 1 | 0 |  | Bonjour Veronique, J'ai cru voir une réponse de Christophe, en effet le diamètre de membrane pour tous les D912 est de 110mm. Best, Nicolas
INC283568 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-18 06:45 | 2023-09-19 10:39 | 1 | 0 |  | mailbox full
INC283563 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-16 12:36 | 2023-09-19 10:39 | 1 | 0 |  | mailbox full
INC283547 | Workgroup | Incident | TR: THERMOMETRE TGX | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-14 15:45 | 2023-09-15 09:59 | 0 | 0 |  | Aurelien, je ne trouve rien dans mes archives sur un TGX, c'est sûrement un TG (thermo gas) qui peut être remplacé par un TSS/TSF. les autres infos que je vois sont qu'il était rempli et que le plongeur faisait 150mm. C'est tout ce que je peux en tirer (c'est très vieux). Best, Nicolas
INC283541 | Workgroup | Incident | FW: Aanvraag | Archive > BHFR Technical Support | Middle |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-14 11:48 | 2023-09-15 10:02 | 1 | 0 |  | Hi Christof, Adel, please find enclosed the datasheet (available on the website) and you can use the P2F to configure it: Order code: MMN5-462.B76/0693_0720_0078 MMN old Code / SAP code: Nominal size: 100 mm 5 Type of mounting: Case and bezel ring AISI 316L (1.4404), stainless steel Bottom connection 4 Process connection: 1/2" NPT 6 Oil filling: BH2, glycerin, high viscosity 2 Pressure unit: Bar (EN standard) B Pressure range: -1 ... 3 bar 76 OPTIONS Dial: "G" added to the pressure unit 0693 Serial number on dial 0720 Approvals: ATEX EC II GD, must be ordered with option 0751 if not included 0078 Best, Nicolas
INC283525 | User | Incident | VS: Tryktransmitter Y913 GZ G1/4 FKM 0+250BAR R | Archive > BHFR Technical Support | Low |  |  | Closed | Soskic, Orhan [soo] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-09-14 09:54 | 2023-09-15 10:25 | 0 | 0 |  | Hi Orhan, I took one sample behind my desk, tested it and it worked, my screw is a bit more than 30mm (see picture). if you really have a shorter one and can't secure the connection, please confirm and I'll relay the information to sales dept so they can create a claim and send you a new connection. Best, Nicolas
INC283504 | Workgroup | Incident | WG: Anfrage Druckmessgerät Kapselfeder / Kunden-Nr. 55122423 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-13 09:18 | 2023-09-13 09:51 | 0 | 0 |  | Hi Felix, you can go ahead. Best, Nicolas
INC283503 | Workgroup | Incident | Y91-3.3.9.B74.R/2164_2151 CAD file | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-13 09:18 | 2023-09-13 09:49 | 0 | 0 |  | Hi Andy, here is the step file. Best, Nicolas
INC283502 | Workgroup | Incident | TR: Demande de Devis Thermostat 71364-01b symbole SNCF 77750362 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-13 09:06 | 2023-09-13 15:16 | 1 | 0 |  | Bonjour Jacques, on ne fait plus du tout cette référence et pas de remplacement, désolé. Best, Nicolas
INC283493 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-13 06:06 | 2023-09-13 09:41 | 0 | 0 |  | Mailbox full
INC283460 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-12 06:10 | 2023-09-13 09:40 | 1 | 0 |  | Mailbox full
INC283455 | Workgroup | Incident | BAIT ¦ Megawatt - Ohmica ¦ Alternative to Vegabar 83 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-11 17:39 | 2023-09-13 10:28 | 1 | 0 |  | Hi Erald, for the mounting and seals definition, for me it fits, but you won't have the slave/master function and I cannot tell you about replacing the VEGA with PFMN and this might need to be validated on the baumer side. the accuracy will be x2 with fitting with a seal (standard now for all electronic mounted on seals) and you can set the range below our min mounting range (0...1bar min for mounting with both D055 and D820). Flanged PFMN-54.BC2R.A1153.33200.1000 + D820-10L.42GK.C2001 2threaded PFMN-54.BC2R.A1153.33200.1000 + D055.2L1C.01 Best, Nicolas
INC283454 | Workgroup | Incident | MR5-4D33.B22 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-11 14:51 | 2023-09-13 09:58 | 1 | 0 |  | Hi Kerstin, here is the code for the glass, please check then with your sales contacts if there is min qty, pricing etc... 10308598 VOYANT TRIPLEX MG/MR/MS5 Best, Nicolas
INC283448 | Workgroup | Incident | TR: rfi MI7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-11 13:54 | 2023-09-13 10:00 | 1 | 0 |  | Anne, sûrement une erreur, un MIX7 plutôt. Best, Nicolas
INC283445 | Workgroup | Incident | TBI100 With Special Temperature Range | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-11 11:57 | 2023-09-13 14:54 | 1 | 0 |  | Hi Bee Tin, if you want to go outside the standard range, there will be an overcost of dial creation and here are the rules: mechanical limits of the TBI in temperature must be respected: -70 ... 600C we can chose a span and slide it up or down but the span must be respected (we can make a -50...350C with a 400C or a 20...420C, but not something that is not 400C large). so the ranges you request are not possible with TBI and must be chosen in TSF/TSS standard range. Best, Nicolas
INC283439 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-10 21:58 | 2023-09-11 11:56 | 0 | 0 |  | Mailbox full
INC283437 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-09 10:00 | 2023-09-11 11:55 | 0 | 0 |  | Mailbox full
INC283436 | Workgroup | Incident | Faisabilité SNCF Mano MFX3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-08 16:30 | 2023-09-19 10:31 | 0 | 1 |  | Aurelien, today we are still producing 4 codes all in type G and 0...6bar (very important as the scale is non-linear). for your request the code is no longer produced and I need you to do an SPR request in order to start over the production. Best, Nicolas
INC283435 | Workgroup | Incident | TR: 0905 RFQ BOURDON HAENNI | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-08 15:09 | 2023-09-13 10:15 | 1 | 0 |  | Bonjour Nathalie, 10262344 est toujours dispo via Gaelle, en stock en suisse, le prix sera à mettre à jour sûrement. Best, Nicolas
INC283427 | Workgroup | Incident | ZTD Switches | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-08 12:54 | 2023-09-08 14:24 | 0 | 0 |  | Hi Andy, it is from BTIN, please find enclosed the datasheet. we can probably offer cross reference or they can buy to india, but we need to check with Jorg how to process for that. Best, Nicolas
INC283419 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-08 10:00 | 2023-09-08 11:33 | 0 | 0 |  | Mailbox full
INC283418 | Workgroup | Incident | combination of Y91+ASIPD+D055 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-08 09:30 | 2023-09-19 14:48 | 0 | 1 |  | Hi Xu Qiang, yes, the mounting is OK. Best, Nicolas
INC283378 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-06 10:00 | 2023-09-08 11:29 | 1 | 0 |  | Mailbox full
INC283367 | Workgroup | Incident | TR: MEP5-463B20/0679/0816 + D621-2N6.4213.G64.02 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-05 15:39 | 2023-09-07 12:36 | 1 | 0 |  | Salut Anne, oui c'est OK. Best, Nicolas
INC283366 | Workgroup | Incident | TR: Capital Misr-RFQ | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-05 14:30 | 2023-09-13 09:36 | 1 | 1 |  | No News from customer, please come back to me with details if needed. Best, Nicolas
INC283354 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-05 09:57 | 2023-09-08 11:29 | 1 | 1 |  | Mailbox full
INC283344 | Workgroup | Incident | TR: Prix contact MFT / Robival | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-04 17:48 | 2023-09-08 11:31 | 1 | 1 |  | Nathalie, il faut faire la demande complete a Fantinelli avec le detail de la config plus les contacts, et ensuite valider la proposition de fantinelli avec Aurélien. Best, Nicolas
INC283308 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-03 21:57 | 2023-09-04 11:03 | 0 | 0 |  | mailbox full
INC283307 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-02 21:57 | 2023-09-04 11:03 | 0 | 0 |  | mailbox full
INC283306 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-01 21:57 | 2023-09-04 11:03 | 0 | 0 |  | mailbox full
INC283305 | Workgroup | Incident | Demande mano SDI | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-09-01 17:27 | 2023-09-19 14:47 | 1 | 1 |  | No feedback from customer, ticket closed until further notice. Best, Nicolas
INC283180 | Workgroup | Incident | DEFINITION MFT7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-29 11:00 | 2023-10-04 09:58 | 1 | 1 |  | Véronique, sans prendre en compte les contacts, le mini pour avoir une tenue à 250bar est 250mbar (avec échelle réduite). des contacts seraient possibles (sous-conditions) mais avec un presse etoupe standard M20*150. Si le client est flexible, les ventes pourront lancer une demande de prix. Best, Nicolas
INC283171 | Workgroup | Incident | OPTION CHANGEOVER CONTACT MR5/MS5 (5003) | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-28 23:09 | 2023-08-29 10:11 | 0 | 0 |  | Hi Angel, Hereafter are some diagrams of how the different contacts work (1: NO; 2: NC; 3: Changeover; 4: Tendency) you can see on NO or NC the circuit is either closed or open depending on where you stand from the setp. on the changeover, on or the other circuit is always closed, so you never have an open circuit signal, you switch from a closed circuit to another closed circuit on a different wire. concerning the tendency, it work the same way with 2 closed circuits, except it's not with a setp, it tells you if it's rising or falling. [Image] Best, Nicolas
INC283166 | Workgroup | Incident | Capillaire CNPE BLAYAIS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-28 14:18 | 2023-08-28 15:09 | 0 | 0 |  | Bonjour Véronique, Daprès la photo, je dirais qu'on est sur du G1/4M / G1/2F mais c'est juste du visuel et je ne suis pas sur de comprendre leurs mesure. ci-dessous les plans du G1/4M et du G1/2F avec les cotes pour verif: [Image] [Image] Best, Nicolas
INC283134 | Workgroup | Incident | TR: spare part of CLAMP, PRESSURE GAUGE 93462 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 16:57 | 2023-08-28 09:54 | 0 | 0 |  | Hi Xu Qiang, I'm sorry, but the description will really not be enough to advise anything. please define more with the customer. Best, Nicolas
INC283133 | Workgroup | Incident | TR: Hydrogen application | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 16:51 | 2023-08-28 09:58 | 0 | 0 |  | Hi Nathalie, both possibilities could be done, it still have to go through an SPR to Franck Ferarrese but you should get a pretty quick answer. please define further the product ou need with the connection type, output and electrical output and send an SPR to Franck. Best, Nicolas
INC283132 | Workgroup | Incident | TR: Demande de prix MEX3 - Dexis | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 16:45 | 2023-08-28 10:06 | 0 | 0 |  | Nathalie, le code du M20*150 standard est "9" mais c'est trop gros pour un MEX3 qui a un carré d'usinage de la connection de 14mm, il faudra passer sur un MEX5 et ce sera au config. Best, Nicolas
INC283120 | Workgroup | Incident | TR: inquiry Y9239GB24A | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 16:03 | 2023-08-28 10:10 | 0 | 0 |  | Nathalie, ca ne devrait pas poser de problème, tu peux demander la codif avec ce code en exemple: 11253583 Y92-2.3G.B29.A Best, Nicolas
INC283116 | Workgroup | Incident | TR: BAPL ¦ E913 ¦ Zapytanie ofertowe | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 12:15 | 2023-08-28 10:17 | 1 | 0 |  | Nathalie, on ne fait plus les E913 mais il est disponible en Y913: [Image] Best, Nicolas
INC283112 | Workgroup | Incident | TR: Y913-57-H31-RZIM9 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 09:36 | 2023-08-28 10:28 | 1 | 0 |  | Nathalie, c'est les anciens codes ZIMCO, ils sont bloqués car ils avaient tous l'option FM/CSA que l'on a pas renouvelé. ci-dessous le code sans FM/CSA et ci-joint la note pour léquivalence technique. 11208048 Y91-3.5.7.H31.R/2169_2158_2154_2151 Best, Nicolas
INC283108 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 06:23 | 2023-08-28 10:29 | 1 | 0 |  | Mailbox full alert
INC283106 | Workgroup | Incident | FW: Price and Availability - Baumer Gauges | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 05:33 | 2023-08-28 10:29 | 1 | 0 |  | Double
INC283105 | Workgroup | Incident | FW: Price and Availability - Baumer Gauges | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-25 05:00 | 2023-08-28 10:30 | 1 | 0 |  | Double
INC283099 | Workgroup | Incident | WG: RFQ 109311 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-24 13:42 | 2023-08-28 15:59 | 1 | 0 |  | Thanks, answered via JST
INC283093 | Workgroup | Incident | TR: Alternative for sensor Y93 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-23 17:21 | 2023-08-28 10:57 | 1 | 0 |  | Nathalie, le Y933 sera bientot remplacé et une nouvelle datasheet sera mise en ligne. en attendant, il est possible de les remplacer par des Y913 + DA mais ici la pression va etre trop basse voici le minimum de montage: [Image] Best, Nicolas
INC283083 | Workgroup | Incident | Besoin plan 3D ou 2D MEX5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-22 15:48 | 2023-08-28 11:05 | 1 | 1 |  | Bonjour Aurélien, ci-joint le plan 3D du MEX, par contre je n'ai pas le D821, tu peux demander la création d'un plan d'ensemble (3D ou 2D via les ventes) Best, Nicolas
INC283066 | Workgroup | Incident | Demande validation code MEP5 et MEX5 pour TFE | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-22 12:36 | 2023-08-28 11:33 | 1 | 1 |  | Aurélien, ce sont de vieux codes mais ce que tu propose me parait bien (avec les details manquants) pour le premier montage pas besoin de 0191, l'AORP va empêcher la surpression et un ARIN2606 est nécessaire car l'AORP est en F/F. a bien verifier la matière du séparateur sur le premier aussi si c'est pour du HCL. Best, Nicolas
INC283062 | Workgroup | Incident | TR: DUS - Maintenance capteurs | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-22 11:24 | 2023-08-28 11:41 | 1 | 1 |  | Aurelien, cette question fait entrer pas mal de paramètres en jeux, mais voici ce que je peux dire: une montée descente peux suffire pour assurer le fonctionnement mais plusieurs permettent de tester la répetabilité, en usine nous testons le setp mini de l'appareil avec Hysteresis min et max puis pareil pour le setp maxi et ensuite nous réglons le setp demandé a lhystérésis demandé ce qui fait au minimum 5 montés/descentes mais pour une verification sur site en faire 3 parait raisonnable. les calibrations en usines sont facilitées par nos moyens en terme de bains/fours mais ils peuvent peut etre limiter les monté/descentes en amplitude autour du point de consigne. Best, Nicolas
INC283058 | Workgroup | Incident | TR: documentation manomètre Baumer | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-22 10:33 | 2023-08-28 11:46 | 1 | 1 |  | Aurelien, c'est plus interne en fait, nos manos sont calibrés à l'air en dessous de 60 bar et à l'eau au dessus. Après, la norme EN837-2 donne des conseils d'utilisation des manomètre standard ou de sûreté selon s'ils sont utilisés sur du liquide ou du gaz, le remplissage et la pression d'utilisation: [Image] Best, Nicolas
INC283055 | Workgroup | Incident | Demande manomètre absolue | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-22 09:36 | 2023-08-28 11:52 | 1 | 1 |  | Aurelien, tu peux partir sur un MA7, je demanderais bien au client de s'assurer de la compatibilité du kérosene avec le faible diamètre de passage et la matière (je ne me rends pas vraiment compte de la viscosité surtout). et pour la résolution, il faudra lui faire valider le standard et la precision sera de +/-2%, il n'y a pas de sens a avoir une resolution de 0.001 si on est precis a +/-0.02. Best, Nicolas
INC283038 | Workgroup | Incident | Config Possible? | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-18 14:42 | 2023-08-28 11:59 | 1 | 1 |  | Hi Andrew, the mountings are possible. Best, Nicolas
INC283029 | Workgroup | Incident | Your mailbox is almost full. | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-18 06:41 | 2023-08-28 09:50 | 1 | 1 |  | Mailbox Full alert
INC283022 | Workgroup | Incident | WG: MEX3-421.N70/0751_0078_9704 special scale | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-16 15:36 | 2023-08-28 12:07 | 1 | 1 |  | Hi Felix, to go to the mbar range you need a low pressure sensing element (capsule) and if you need the 316L you'll only have the MCX, then the closest range is [Image] for the TBI, the scale can be: [Image] Best, Nicolas
INC282987 | Workgroup | Incident | Y91 0...250 bar burst pressure | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-10 10:15 | 2023-08-28 14:30 | 1 | 1 |  | Hi Andrew, the burst pressure for a 250bar range is 750bar, it's a mistake and will be corrected in the next version of the datasheet. Best, Nicolas
INC282969 | Workgroup | Incident | D040-01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-08 17:30 | 2023-08-28 13:53 | 1 | 1 |  | Answered via Jörg Strüfing.
INC282967 | Workgroup | Incident | WG: V23070 3D STP FILE | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-08 15:18 | 2023-08-28 14:21 | 1 | 1 |  | Hi Felix, please find enclosed the 3D drawings, the material and range doesn't change the dimensions. Best, Nicolas
INC282959 | Workgroup | Incident | bellow of ZPN | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-08 09:27 | 2023-08-28 14:15 | 1 | 1 |  | Dear Xu Qiang, I checked directly with the people in production and it's normal, during the forming overpressure, the below will take it's final for with the constraints applied and can slightly shift a bit, with checking the pictures there is nothing unusual. but if they are worried, the best is to test it around the setp and check accuracy and repeatability. if he product is out of calibration, then please return it via the quality return process (but I bet everything is fine, those products are carefully checked). Best, Nicolas
INC282932 | Workgroup | Incident | MS5-4D30.J22/4710-11091633 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-04 12:57 | 2023-08-04 15:54 | 0 | 0 |  | Hi Daniel, you can ask for it to be reopen, nothing blocks technically. the only thing is you need to check with your sales contact if the dial need to be recreated (possible fee) Best, Nicolas
INC282928 | Workgroup | Incident | need step. drawing | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-04 11:54 | 2023-08-04 15:55 | 0 | 0 |  | Hello Sascha, it's not available, but there is the possibility to ask the sales when you place the order for the creation of the step or dimensional drawing (possible fee). Best, Nicolas
INC282899 | Workgroup | Incident | scale for MEX3 S22/(ammonia) | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-02 23:27 | 2023-08-03 14:02 | 1 | 0 |  | Hi Angel, with the 0 at 12o'clock and -1...25bar range it'll be anyway a creation. otherwise we have the standard -1...24bar/C NH3 (black/red) available (and the options for connexion at 3/9/12 o'clock). Best, Nicolas
INC282895 | Workgroup | Incident | thread of DRO 80 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-02 11:57 | 2023-08-02 13:57 | 0 | 0 |  | Hello Xu Qiang with only a few exception, threads should only be mounted on their specific conter parts, he must have a G1/2F to mount the gauge in G1/2M. I don't see any thread adapter on his panel. he can verify the thread he has with a Screw pitch gauge. Best, Nicolas
INC282892 | Workgroup | Incident | FW: Enquiry For Contact Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-02 11:06 | 2023-08-03 15:49 | 1 | 0 |  | Hi Bee Tin, I can't see any number or get back to any SAP codes with the pictures and many things are weird (blue back vent, stamp, colored filling etc...) You'll have to redefine with the customer and his needs I'm afraid. Best, Nicolas
INC282889 | Workgroup | Incident | Y913 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-08-02 10:39 | 2023-08-02 11:14 | 0 | 0 |  | Hi Daniel, I think it's more because many people are on holidays right now, only 1 person is in charge of the order confirmations instead of like 6. but if there is an emergency, you can push it via the hierarchy to the head of the site (Stephane Hennequin) with the details of the order and the line you need a confirmation on in emergency. Best, Nicolas
INC282851 | Workgroup | Incident | E913 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-31 13:45 | 2023-07-31 15:59 | 0 | 0 |  | Salut Veronique, On ne va pas pouvoir être sur a 100% mais vu qu'on ne voit aucun filet ni teflon, ce n'est pas du NPT et visuellement le diamètre correspondrait pour moi a du 1/2. donc on peut être presque sur (à moins que ce soit du spécifique comme un 3/8 (très rare)) que c'est du G1/2. Best, Nicolas
INC282816 | Workgroup | Incident | Certificat Ferroviaire et feu fumé Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-28 11:09 | 2023-07-28 14:33 | 0 | 0 |  | Bonjour Jacques, ci dessous les documents, il y a 2 feu/fumée selon la connexion électrique et l'EN est sur le certificat de conformié CE global. Best, Nicolas
INC282809 | Workgroup | Incident | SEPARATEUR YOKO | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-27 16:15 | 2023-07-27 17:14 | 0 | 0 |  | Salut Véronique, je ne m'occupe pas du tout des prix, il faut voir avec le commerce. Best, Nicolas
INC282805 | Workgroup | Incident | WG: Quote | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-27 14:21 | 2023-07-27 17:39 | 0 | 0 |  | Hi Felix, please find the details of this old product in the enclosed datasheet. the placement product will probably be an RP2. Best, Nicolas
INC282798 | Workgroup | Incident | MEX5 With Range 1...24/R134A | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-27 11:06 | 2023-07-27 17:53 | 1 | 0 |  | Hi Bee Tin, yes it exist, here is an example: 11146060
INC282776 | Workgroup | Incident | MEX3-D21.B35 to confirm dial scale with Minimum scale of 2 bar per division / EXIM | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-27 06:21 | 2023-07-27 17:48 | 0 | 0 |  | Hi Bee Tin, here is what we would print on a MEX3xxxC35 and it seems to be what you're looking for: [Image] Best, Nicolas
INC282768 | Workgroup | Incident | BAES : demande de faisabilité pour mano sur bride | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 14:27 | 2023-07-26 14:50 | 0 | 0 |  | Cyrille, Ok pour: MEX5-D33.F35/0760_0751 + D418-210.0L2.025.4Q0.M0A.000_09 MEX5-D33.F35/0760_0751_D418-210.0L2.026.4Q0.M0A.000_09 Best, Nicolas
INC282760 | Workgroup | Incident | WG: EIL - Anfrage-Nr.: 01/20230726 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 12:39 | 2023-07-26 14:31 | 0 | 0 |  | Hi Felix, usually I would have translated it to a newer reference, but all people that could validate the mounting are out in holidays for the 3 next weeks. I think for this time the only way to get something price and code creation) would be to stick to the existing codes. I have some existing in SAP with the FUD but without the 0661/9710 and I have a DROE with those options but without he FUD. keep in mind also that the specific dial will need to be created and validated. see codes below as examples. 10344971 10350235 10344972 11112222 Best, Nicolas
INC282758 | Workgroup | Incident | TR: Flughafen Kloten_T14_Kälte_Notkühlung | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 11:54 | 2023-07-26 14:50 | 0 | 0 |  | Nathalie, avec seulement une vue de face je vais avoir du mal à définir tous les détails surtout que c'est du Nuova Fima. on pourrait offrir du MEX3-B21.X2P/9712 et MEX3-B21.X1L/9712 en bar/C R410A. il faudra créer les cadrans mais ca doit être possible a partir des photos avec confirmation des égalités C-bar par le client. Best, Nicolas
INC282751 | Workgroup | Incident | FW: Shell Flanged gauges | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 10:30 | 2023-07-28 16:29 | 1 | 0 |  | Then, sorry but without more information, I'll not be able to find a replacement. Best, Nicolas
INC282744 | Workgroup | Incident | WG: PO PT2021/080019-1 1849641 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 09:27 | 2023-07-26 09:36 | 0 | 0 |  | Hi Felix, we have the diaphragm seal DANC that can be fitted with a gauge and the DP100 with mechanical transmission. Best, Nicolas
INC282740 | Workgroup | Incident | WG: Information // 30169414 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-26 08:45 | 2023-07-26 09:45 | 0 | 0 |  | Hi Felix, for the transmitter, the E91 was the phase out non-ATEX version of the Y91. 1545/1310/1359 are relative to a diaphragm seal. the replacement would be Y91333B76R/1545/1310/1359 /9010 (idem 11261649 but in B76 and without the span adjustment option). Best, Nicolas
INC282728 | Workgroup | Incident | RT2E/N42223C6 / temperatur switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-25 11:36 | 2023-07-25 12:17 | 0 | 0 |  | Hi Sascha, yes, there is a mistake in the code, maybe the customer missed something, I can't define it, please see with the customer. Best, Nicolas
INC282724 | Workgroup | Incident | WG: 202307241818 RE: Inquiry of Replace of Wika | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-25 09:00 | 2023-07-25 09:50 | 0 | 0 |  | Hi Felix, this is a Schneider reference (MG are usually monoflange but you have M/F thread so I think it's something else). you have the direct contact to Schneider, you should ask them directly. Best, Nicolas
INC282722 | Workgroup | Incident | WG: Anfrage Kunke | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-25 08:36 | 2023-07-25 09:53 | 0 | 0 |  | Hi Felix, special threads are usually not impossible but it needs to go through a request to our supplier with the complete code and quantities via the sales team. What I can tell you is that below 10units, it'll be extremely expensive. also the question is why high pressure connections on a differential limited to 400bar ? Best, Nicolas
INC282717 | Workgroup | Incident | TR: Request for price of CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-24 15:06 | 2023-07-25 09:55 | 0 | 0 |  | Nathalie, non, pour une étendue basse et reglable il faudra se tourner vers le PBMN basse pression (et en tout cas la suisse qui sauront mieux que moi). Best, Nicolas
INC282715 | Workgroup | Incident | Demande de renseignements pressostat différentiels | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-24 14:48 | 2023-07-25 10:07 | 1 | 0 |  | Aurelien, oui c'est possible, voir pour les connections (cristal) a bien valider avec le client. Best, Nicolas
INC282713 | Workgroup | Incident | Cadran creation | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-24 14:45 | 2023-07-25 12:18 | 1 | 0 |  | Hi Jürgen, please take the code of a 400mbar for definition and check with sales for the price. Best, Nicolas
INC282711 | Workgroup | Incident | MEX3-F50.H27/0764 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-24 14:33 | 2023-07-25 17:59 | 1 | 0 |  | Hi Daniel, the P2F was based on if the dial is EN or not which was not the right way of seing it. the EN considers that only safety gauges (MEP for example) should be used on Oxygen so MEP will have the option 0765 and the MEX will not. of course this is based on the EN requirement but you can ask for psi ANSI, then your dial will not have EN marking but you still respect in principle the EN so 0765. Best, Nicolas
INC282705 | Workgroup | Incident | FW: Stainless Steel Capillary | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-24 11:21 | 2023-07-24 11:55 | 0 | 0 |  | Double, I found the first one sorry, I'll answer on it. Best, Nicolas
INC282692 | Workgroup | Incident | Stainless Steel Capillary | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-21 12:42 | 2023-07-24 11:56 | 1 | 0 |  | Hi Daniel, we will not sell capillary without the connections, sorry. but they can order it with the connections as an AKPL with the desired connections. Best, Nicolas
INC282675 | Workgroup | Incident | TR: Zapytanie ofertowe DM/1318/2023 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 15:36 | 2023-07-20 16:06 | 0 | 0 |  | Bonjour Anne, c'est un raccord rayonnant de MEX8 avec un filetage spécifique. ça ne peut donc pas être vendu seul car ils ne peuvent pas le souder sur le boîtier, souder le tube etc... Best, Nicolas
INC282673 | Workgroup | Incident | Capteur de pression sur séparateur | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 15:12 | 2023-07-20 16:11 | 0 | 0 |  | Salut Cyrille, nous on ne va avoir que le D055 en G1/2 membrane affleurante HC276 et la limite de montage est 100bar mini. mais si on part sur un Y923 ou un PBMN Flush par exemple la pression le raccord et la membrane affleurante seront OK, il faut juste faire la membrane en HC276 et je pense que c'est possible mais ca devra passer par un SPR. Best, Nicolas
INC282668 | Workgroup | Incident | WG: Customer request: parts | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 13:57 | 2023-07-20 16:01 | 0 | 0 |  | Hi Jorg, it'll simply be an RPPN7-A3.204 as far as I can see. Best, Nicolas
INC282665 | Workgroup | Incident | DRS VALEURS TEMPERATURE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 10:21 | 2023-07-21 10:01 | 1 | 0 |  | Veronique, voici les conclusions du rapport labo sur la reduction de temperature sur des montages dont un DRS: [Image] Best, Nicolas
INC282663 | Workgroup | Incident | Séparateur en Hastelloy C22 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 09:24 | 2023-07-20 10:13 | 0 | 0 |  | Bonjour Cyrille, je pense surtout que C22 est une erreur, ca doit être B22 ou C276. après, tu trouvera ces 2 nuances sur de nombreux séparateurs. si le C22 existe vraiment et que je ne le connais pas, il faudra définir ton séparateur avec un X et faire une SPR. Best, Nicolas
INC282657 | Workgroup | Incident | DRS VALEURS TEMPERATURE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-20 08:06 | 2023-07-20 10:14 | 0 | 0 |  | Bonjour Veronique, oui, cette temperature max vient des essais fais avec eux pour la validation du montage. Best, Nicolas
INC282653 | Workgroup | Incident | DR 63/452.111 0..6BAR 0201 - SAP 10319886 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-19 16:36 | 2023-07-21 09:21 | 0 | 0 |  | Hi Daniel, here is the complete characteristics: Stainless 63mm gauge Panel mounting center back connection Cu Alloy tube G1/4 0...6bar over pressure safe measuring element I would then advise a MEX3G20B20, with the Stainless tube it'll be more resistant overall to over pressure and under pressure than a DR63. do you have it's potential std pressure, overpressure and underpressure ? Best, Nicolas
INC282650 | Workgroup | Incident | TR: PBMN-24B79RA1440320200/FÜD175/25/1/3_LRS1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-19 14:42 | 2023-07-21 16:05 | 0 | 0 |  | OK pour le montage: PBMN-24B79RA1440320200/FÜD175/25/1/3_LRS1 Best, Nicolas
INC282642 | Workgroup | Incident | FW: Fuel Oil Centrifuge Pressure Switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-19 10:36 | 2023-07-19 11:44 | 0 | 0 |  | Hi Bee Tin, quite a difficult question, I think it's a GA 201 serie from wika but the code is specific and the diaphragm seal is not clear either (enclosed the datasheet from wika). I suggest to go standard and offer an RP2N with 2 hermetical contacts and say we have the option to fit with a diaphragm seal. Best, Nicolas
INC282636 | Workgroup | Incident | Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-19 08:39 | 2023-07-19 10:03 | 0 | 0 |  | Hi Kerstin, we receive this request from several brokers in the world, here is the answer I sent to Jorg about it: Hi Jörg, not the easiest question but here are the translations (oldest to old to newest): ZPN205SH = RPPNCA3205 = RPPN-7C3.205 ZPN209SH = RPPNCA3209 = RPPN-7C3.209 ZPN203GHHINT = RPPNV(GSHH)Y3203 = RPPY-7W3.203 ZPN204GSHHINT = RPPNVY3204 = RPPY-7W3.204 ZPN201GSH = RPPNDA3201 = RPPN-7D3.201 ZDPN224SIINT = RDPNAY3224 = RDY8A3224 ZDPN234SHH = RDPNWA3234 = RDN-8W3.234 this is assuming: the codes were correct INT means INTrisically safe connection is G1/2 (was the standard at the time). Best, Nicolas
INC282632 | Workgroup | Incident | ARIN | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 16:30 | 2023-07-18 16:46 | 0 | 0 |  | Bonjour Veronique, c'est normal, c'est un adaptateur qui du G1/2M en G1/2M donc il ne sert a rien a part a ajouter la longueur d'un ARIN. Pourquoi tu as besoin de ca ? Best, Nicolas
INC282627 | Workgroup | Incident | PBMN-24B79RA1440320200/FÜD175/25/3_LRS1 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 14:24 | 2023-07-19 11:45 | 0 | 0 |  | Perfect, thanks Kerstin. tell me if I can do anything else. Best, Nicolas
INC282626 | Workgroup | Incident | 10344971 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 14:21 | 2023-08-29 10:31 | 0 | 1 |  | Dear Colleague, the article 10344971 is still possible to produce, I just sent the costs to sales, you can send them the request for pricing. I'm sorry it was so long to get a price. Best, Nicolas
INC282625 | Workgroup | Incident | Cal Certificates | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 14:00 | 2023-07-18 16:43 | 0 | 0 |  | Hi Jürgen, you're right, everything that is done elsewhere and managed in stocks cannot have a certificate (MIT, MAT, MTR) some that are done externally are still ordered on request and can have a cal cert (MFT, TSF, TSS,...) everything that is done here or finished here can have one (including the MMD). another consideration is the price, not many people ask for a calibration cert on a 5 gauge and if accuracy and long term use is something, they should go to the MEX line. Best, Nicolas
INC282622 | Workgroup | Incident | Nace conformity certificate 90000275 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 11:48 | 2023-07-18 11:51 | 0 | 0 |  | Hi Jürgen, for this, please contact your commercial contacts (Anne and Jean-Marie I guess) I can't give prices. Best, Nicolas
INC282609 | Workgroup | Incident | CAD drawing of MR5-4D63-B19R | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-18 07:54 | 2023-07-18 09:12 | 0 | 0 |  | Hi Xu Qiang, I have it only in 1/2G (enclosed) sorry. Best, Nicolas
INC282596 | Workgroup | Incident | Kindly confirm | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-17 14:57 | 2023-07-17 15:15 | 0 | 0 |  | Hi Anil, the DT1 is limited to 160bar so it'll not work with over range, you need a DT8 MEP5-463.B33/0191/0678/DT8.2N200.2.60.02 (max pressure of the complete fitting 280bar). for the option 0678, We do not advise it as the gauge is filled, the customer will not be able to open the gauge and set the pointer (it makes it a lot more complicated) and the filling can lubricate the mechanism of the pointer and cause drift. Best, Nicolas
INC282561 | Workgroup | Incident | FW: Baumer pressure gauge with contact_TA / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-13 06:24 | 2023-07-13 17:40 | 0 | 0 |  | Hi Bee Tin, I can't get all informations from a picture but I can already tell you a few things, capillary and flange connections are not from us. the contact gauge itself is a MS5, MR5 or MG5 with 1 normally open contact. the accuracy is 1% now (change from a couple years back) and the tube is in 1.4404 instead of 1.4571 on the dial. Best, Nicolas
INC282559 | Workgroup | Incident | FW: Baumer pressure gauge with contact_TA / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-13 05:19 | 2023-07-13 09:19 | 0 | 0 |  | Double sending
INC282557 | Workgroup | Incident | FW: Baumer pressure gauge with contact_TA / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-13 04:12 | 2023-07-13 09:19 | 0 | 0 |  | Double sending
INC282556 | Workgroup | Incident | FW: Baumer pressure gauge with contact_TA / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-13 03:39 | 2023-07-13 09:19 | 0 | 0 |  | Double sending
INC282553 | Workgroup | Incident | MEX5 + D820 | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 17:06 | 2023-07-12 17:16 | 0 | 0 |  | Salut Aurelien, C'est OK pour le montage (c'est le minimum) et pour le prix, tu as le prix du MEX dans le P2F et le D820 il faudra le demander a ton contact vente. [Image] Best, Nicolas
INC282552 | Workgroup | Incident | Info sur tube immersion TBHI | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 16:51 | 2023-07-12 17:19 | 0 | 0 |  | Aurelien, sur la datasheet, tu pourras trouver le nota ci-dessous, donc ce n'est pas possible, pour moins de 100mm tu es obligé d'avoir un tube de 8mm. [Image] Best, Nicolas
INC282550 | Workgroup | Incident | TR: BAIT ¦ Tech information | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 14:30 | 2023-07-12 15:30 | 0 | 0 |  | Anne, le code ci-dessous est obsolete, il faudrait le remplacer par: DRCG101.211.133.535/3272 + DANC01-2L2B.01/Q003_Q008 le certificat FDA du LRS1 est dispo et le Ra doit être demandé sur le 3.1 matière. Best, Nicolas
INC282540 | Workgroup | Incident | SDI 439 - Lans Nivotherm (21107104) - PO RW-99999-14166 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 11:15 | 2023-07-13 09:33 | 1 | 0 |  | Hi Rudi, please ask for the codification of 10262331 + 9936, it'll be a laser marking on the dial. Nadia will revert with the new code and the price. Best, Nicolas
INC282539 | Workgroup | Incident | ED701 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 11:03 | 2023-07-12 15:32 | 1 | 0 |  | Bonjour Cyrille, désolé, on ne pourra pas, c'est trop ancien par contre il y a sûrement des details sur le corps du transmetteur si tu arrive à avoir des photos. Best, Nicolas
INC282518 | Workgroup | Incident | ref 11181062 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-12 09:09 | 2023-07-12 09:36 | 0 | 0 |  | Bonjour Yannick, il y a une petite confusion, le code SAP que tu as est bien un Schneider: 11181062 SAFFSA-N4N4-A mais le code en jaune (2VT...) est un code indien aussi pour un manifold 2 voies. ci dessous les dimensions et caractéristiques du Schneider: [Image] [Image] [Image] Best, Nicolas
INC282511 | Workgroup | Incident | Demande de certificat | Archive > BHFR Technical Support | Middle |  |  | Closed | Jay, Aurélien [auj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-11 17:03 | 2023-07-11 17:14 | 0 | 0 |  | Voici le certificat CE un MFE, après il faudra voir selon le besoin comme on a fait avec philippe sur l'affaire en cours. Best, Nicolas
INC282506 | Workgroup | Incident | Détails configuration 95000084 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-11 14:33 | 2023-07-11 16:02 | 0 | 0 |  | Salut Jacques, ce code est un code dummy qui pouvait contenir différentes configurations, je ne pourrais savoir qu'avec le numéro de commande désolé. On n'utilise plus ces code, maintenant on utilise que des codes uniques. Best, Nicolas
INC282505 | Workgroup | Incident | VALDIATION REFERENCE POUR YOKO | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-11 12:30 | 2023-07-13 15:11 | 1 | 0 |  | Veronique, voici le séparateur pour montage DRS HT que nous pouvons proposer pour une échelle de 100mbar (70mmHg) absolu et jusquà 300C ? D910L100T(ou U)8C0 conditions habituelles pour conditionnement et certificats selon contrat TOTAL YOKO Best, Nicolas
INC282499 | Workgroup | Incident | TR: BAIN / We required Y92 250 Bar, 4 to 20 mA | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-11 10:06 | 2023-07-11 14:09 | 0 | 0 |  | Anne, c'est ok pour créer ces codes. Best, Nicolas
INC282467 | Workgroup | Incident | Capteur de pression sur séparateur | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-07 16:18 | 2023-07-07 16:52 | 0 | 0 |  | Bonjour cyrille, ci-dessous les details du code pour t'aider, soit un PBMN soit n Y9 pourra être monté sur un séparateur 1626. le Y9 sera un poil moins precis mais le PBMN plus cher. Best, Nicolas
INC282462 | Workgroup | Incident | BARFLEX - periodic calibration | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-07 13:51 | 2023-07-07 16:49 | 0 | 0 |  | Hi Rudi, we don't advise calibration periods, it depends greatly on the usage, the history of the instruments and the policy of the user. what I can tell you is what we do internally for a comparable product this is the lifecycle of a reference gauge we use and you can see we started with 1year cycle to then expand it to 2years after the first verification. [Image] Best, Nicolas
INC282454 | Workgroup | Incident | MEP5-460.B19 Kempstons | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-07 10:39 | 2023-07-07 11:32 | 0 | 0 |  | Hi Daniel, sure, the codes of the overpressure protector and the intermediate connector are below, you can find it in SAP or the P2F and the datasheets are available on the website. AORP381+ARIN2606 Best, Nicolas
INC282437 | Workgroup | Incident | 10309304 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-07 08:30 | 2023-07-07 09:12 | 0 | 0 |  | Veronique, il est en Z10 CNF 18-09 (compo ci-dessous) qui est bien un Inox: [Image] Best, Nicolas
INC282436 | Workgroup | Incident | MEX5 With Option 0078 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-07 08:18 | 2023-07-07 11:30 | 0 | 0 |  | Hi Bee Tin, if you order 0078 during creation it'll automatically add the option 0751, both options are necessary. Best, Nicolas
INC282433 | Workgroup | Incident | Raccord réfrigération | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-06 16:39 | 2023-07-13 09:36 | 0 | 1 |  | Bonjour Cyrille, je clos le ticket pour l'instant, si tu as des nouvelles, nhésite pas a revenir ver moi. Best, Nicolas
INC282432 | Workgroup | Incident | Ohne Betreff | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-06 16:21 | 2023-07-07 16:58 | 1 | 0 |  | Yannick, pourquoi pas un RP2E ? en 0...4 ou 0...6bar ? toujours commencer par voir si le RP2 est possible avant d'envisager un RPP. sinon ton montage me parait bon, les conditions de services sont clairement fausses mais je pense que tu les interprètes bien (les valeurs dans hysteresis sont en fait les valeurs déchelle). Best, Nicolas
INC282426 | Workgroup | Incident | TED ModBus | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-06 14:36 | 2023-07-12 09:25 | 0 | 1 |  | Bonjour Cyrille, oui, tu peux y aller: [Image] Best, Nicolas
INC282425 | Workgroup | Incident | RP2 setpoint change | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-06 14:15 | 2023-07-06 17:55 | 0 | 0 |  | Sure, they have it on the user instructions, please find a copy attached in the ticket. Best, Nicolas
INC282423 | Workgroup | Incident | MAJ REFERENCE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-06 13:36 | 2023-07-06 17:51 | 0 | 0 |  | Bien sur, c'est maintenant un RP2NL355 + D0432L03, par contre je ne vois rien sur le remplissage, a voir si BH1 ok. Best, Nicolas
INC282391 | Workgroup | Incident | EMBALLAGE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-05 11:57 | 2023-07-12 16:03 | 1 | 1 |  | Bonjour Veronique, voici les conditionnements cartons qui seraient utilisés pour chaque (pou le TSF c'est une estimation basée sur le MEX5 qui a aussi 10m de capillaire et un diamètre de 100mm). MEX5 + séparateur + 10m de capillaire 400x200x160mm (10306864) MX7 + 2 séparateurs + 10m de capillaire 500x400x400mm TSF + 10m de capillaire + bulbe de 140mm 400x200x160mm TBI plongeur de 250mm 285x103x103mm (11147434) Best, Nicolas
INC282389 | Workgroup | Incident | TR: CoC missing | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-05 10:48 | 2023-07-06 17:31 | 0 | 0 |  | Anne, so the answer is "we don't have" and if it's important for the customer, we need to know which CE Directive they want us to be conform to. Best, Nicolas
INC282386 | Workgroup | Incident | RE: Kind Reminder - IC prijzen voor DP100 pressure gauges - Agidens Life Sciences | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-05 10:27 | 2023-07-05 14:06 | 0 | 0 |  | Anne, le Ra est la Rugosité d'une surface, en standard elle est inférieure à 0.8 mais en option (SPEZ) elle peut être demandée à 0.4 (exemple ci-dessous) ca aura un impact sur le prix (que je ne connais pas). [Image] Best, Nicolas
INC282385 | Workgroup | Incident | TR: Request for offer | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-05 10:21 | 2023-07-05 14:17 | 0 | 0 |  | Anne, il semble que ce soit un DRC100xxx525 monté sur un 154x mais le document word avec la photo ne suffit pas pour définir le reste. il va falloir redéfinir le code avec le client. Best, Nicolas
INC282374 | Workgroup | Incident | FW: aanvraag ED 701 820. 113. A25. 19 | Archive > BHFR Technical Support | Middle |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-04 11:39 | 2023-07-04 14:05 | 0 | 0 |  | Hi Kees, I can find out what it was but it was for mounting on a diaphragm seal which I'll not be able to find (to old, it was before SAP). ED701: transmitter 8: Male G1/2 for diaphragm seal mounting 2: 4....20mA 0: CE Conform 1: 3 wires cable IP65 1: 0.4% accuracy from -10 to 125C 3: Gauge pressure A25: 0...5bar 1: stainless steel 9: no filling oil (the oil filling will be the one of the diaphragm seal). Attached in the ticket the datasheet. we should be able to replace with a Y913 + diaphragm seal but I can't define the seal Best, Nicolas
INC282373 | Workgroup | Incident | Technische vragen aan Vendome | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-04 11:36 | 2023-07-04 14:06 | 0 | 0 |  | I think I'm in copy for the other request, I close this one. Best, Nicolas
INC282362 | Workgroup | Incident | step files | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-04 10:30 | 2023-07-04 14:15 | 0 | 0 |  | Hello Kerstin, here is the closest I was able to find: TBH100-212.161 -> Datasheet TBH (dimensional drawing) MEX5-D30.B20 -> OK T 09143.0005 -> Datasheet TBH (dimensional drawing) Best, Nicolas
INC282353 | Workgroup | Incident | RHODIA HYA 2087057 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-04 08:42 | 2023-07-04 09:34 | 0 | 0 |  | Voici le plan de ce qu'ils avaient, si ils sont sur du femelle droit coté process il serait opportun de tout passer en G1/2, sinon si ils ont du conique, on a pas de solution standard, il faudra voir pour un adaptateur selon le nombre de pièces. [Image] Best, Nicolas
INC282348 | Workgroup | Incident | Pls check scale angle | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-03 18:45 | 2023-07-06 10:51 | 0 | 0 |  | Hi Jörg, it won't be possible with a standard movement as we have a max of 295 which means it'll be impossible to calibrate and will provide 0 movement at 0 and max scale. but if you think it's important, Nicolas thinks we could do something with using the MVX7 movement and make a test in production. tell me what you think and we can organise the next steps. Best, Nicolas
INC282340 | Workgroup | Incident | Manifold 5VR-S-CT-P-44-FF-N | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-03 14:27 | 2023-07-07 15:44 | 0 | 1 |  | Hi Daniel, please send an SPR request to Nathalie to trace the request, I'll ask India for a price and we will make you an offer from France (seen with Jörg). Best, Nicolas
INC282338 | Workgroup | Incident | Reliance Industries - details of procedure being followed to attach Tantalum diaphragm to SS saddle | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-03 13:45 | 2023-07-03 14:32 | 0 | 0 |  | Hi Anil, this is really production informations, we can't disclose it. if they worry about something, our products are tested according to all the standard european norms, ISO9001 and 9100 for the quality system and respect the PED. If they have a specific question, maybe we can answer them but we won't disclose a complete internal procedure. Best, Nicolas
INC282337 | Workgroup | Incident | D621 Seal / Bitax | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-07-03 12:15 | 2023-07-03 14:47 | 0 | 0 |  | Hi Bee Tin, there is missing information but if we start by offering the most basic, I would offer: MEP5460B20 (B22 for the second one) with either: Order code: D425-2N60.213.G6002 old Order / SAP code: Flange material: Stainless steel (1.4404) 2 Instrument connection: ►1/2 NPT Female N Diaphragm material : Hastelloy C 276 (HC) 6 Diaphragm Coating : ►Without 0 Standard Flange: ►ANSI B 16-5 2 PN/ DN: Class 150 / 3/4" 13 Flange face : ►Raised face RF 1,6 (Type B1) EN1759-1/ANSI B16-5 G Bottom part material : Hastelloy C 276 (HC) 6 Bottom housing Coating : ►Without 0 Oil Filling : LRS2 (-60+150C) Silicon DC 510 02 or SAP No. Order code D621-2N60-213-G6002 D621 Flange material Stainless steel (1.4404) 2 Instrument connection 1/2 NPT Female N Diaphragm material Hastelloy C 276 (HC) 6 Diaphragm Coating Without 0 Standard Flange ANSI B 16-5 2 PN/ DN Class 150 / NPS 3/4" 13 Flange face Raised face RF 1,6 (Type B1) EN1759-1/ANSI B16-5 G Bottom part material Hastelloy C 276 (HC) 6 Bottom housing Coating Without 0 Oil Filling LRS2 (-60+150C) Silicon DC 510 02 Best, Nicolas
INC282289 | Workgroup | Incident | TR: High pressure gauges MFPx - ExxonMobil - 0 ... 4000 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-29 10:09 | 2023-06-29 17:00 | 1 | 0 |  | MFP est un produit qu'on achète chez STIKO et qui est géré par Stockach. il faut faire la demande a Jorg et la commande sera pour BHDE. Best, Nicolas
INC282267 | Workgroup | Incident | LYONDELL – 20524122- POSTE 60 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-28 12:30 | 2023-06-29 09:59 | 1 | 0 |  | Bonjour Yannick, il n'y a pas de question, je ne sais pas ce que tu veux... Best, Nicolas
INC282219 | Workgroup | Incident | question tech DP SOLYRO - SANOFI | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-27 10:09 | 2023-06-29 09:56 | 1 | 0 |  | Bonjour Yannick, oui, une fluctuation de la précision avec la température tu le retrouvera sur à peu pres tous les instruments. ca donne: 145-20=125, 125/10=12.5, 12.5x0.8=10% it means we are allowed to have 11.6% accuracy. of course we are way below, but I think 0.2 is acceptable. Best, Nicolas
INC282206 | User | Incident | WG: Richiesta di offerta PSR.04490 | Archive > BHFR Technical Support | Low |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-06-27 09:12 | 2023-07-04 10:07 | 1 | 1 |  | Hi Felix, not the easiest question but here are the translations (oldest to old to newest): ZPN205SH = RPPNCA3205 = RPPN-7C3.205 ZPN209SH = RPPNCA3209 = RPPN-7C3.209 ZPN203GHHINT = RPPNV(GSHH)Y3203 = RPPY-7X3.203 (microswitch to be redefined) ZPN204GSHHINT = RPPNVY3204 = RPPY-7X3.204 (microswitch to be redefined) ZPN201GSH = RPPNDA3201 = RPPN-7D3.201 ZDPN224SIINT = RDPNAY3224 = to redefine (RDY) ZDPN234SHH = RDPNWA3234 = RDN-8W3.234 this is assuming: the codes were correct INT means INTrisically safe connection is G1/2 (was the standard at the time). Best, Nicolas
INC282187 | Workgroup | Incident | Industrial Temperature Switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-27 06:19 | 2023-06-27 10:08 | 0 | 0 |  | Hi Bee Tin, please check again on the website, I think it might have been a bug: https://www.bourdon-instruments.com/medias/__secure__/Baumer_RTNE3_EN_DS.pdf?mediaPK=8905754378270 Best, Nicolas
INC282177 | Workgroup | Incident | Mano sur appli cryo | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-26 09:51 | 2023-06-27 09:48 | 1 | 0 |  | Adrien, réponse via teams, temperature trop basse l'inox ne tiendra pas. Best, Nicolas
INC282171 | Workgroup | Incident | Gauges and seals | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 17:45 | 2023-06-27 09:56 | 1 | 0 |  | Hi Andrew, the diaphragm seals mounting is possible, no problem: [Image] for the AGF, I think this is what you need, the G face type doesn't exist. 270mm no problem, take 300 for reference Tapered is possible but you're at the minimum thickness so you can only do bigger Order code: AGFB-0227L.1G300.P AGF old Code / SAP code: Flange Type: With raised face fl ange (all types B, B1, B2) B Extension: Without 0 Material: Stainless steel 1.4404 2 Flange Diameter: Class 300 / 2" 27 Thread on instrument side: G1/2" L Inner Diameter Thermowell: 7 mm 1 Outer Diameter / outer diameter: 13 mm G Plunger lenght: 300 mm 300 Design: Straight shank P Best, Nicolas
INC282168 | Workgroup | Incident | TR: BAIT ¦ ERREL FQ 9608MC | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 15:12 | 2023-06-23 16:33 | 0 | 0 |  | Anne, ce sont des produits indiens, je ne sais pas si ils veulent / peuvent avoir accès au catalogue indien. ci-joint la datasheet pour commander en inde ou définir le remplacement. Best, Nicolas
INC282167 | Workgroup | Incident | TR: Need Lifecycle Status, RoHS 3 Status and Replacement | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 14:39 | 2023-06-23 16:34 | 0 | 0 |  | Anne, d'habitude j'aurais recherché mais comme je ne suis pas dispo, peux tu voir avec Paul pour les sujets qualité cette semaine stp ? Best, Nicolas
INC282166 | Workgroup | Incident | FW: MGA RFQ 108094 | Archive > BHFR Technical Support | Middle |  |  | Closed | Goldsworthy, Sally [saha] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 14:36 | 2023-06-23 16:38 | 0 | 0 |  | Hi Sally, this will not be enough information to define a switch, can you please see with the technical sales people in UK if they can define it with the customer. I would have been in more details usually I'm really sorry but I'm full week ou to train people and I'll not be able to give as much support as usual. please don't think is always this kind of answers and continue to use the ticketing system starting week 27. Best, Nicolas
INC282165 | Workgroup | Incident | equivalence | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 14:09 | 2023-06-23 16:41 | 0 | 0 |  | Yannick, ca me semble bien, tu peux prendre les trous avec bouchon, pas besoin de vannes. demande bien au client de valider le montage. le D418 est presque surdimensionné mais c'est le plus simple pour avoir les trous de purge et event. Best, Nicolas
INC282148 | Workgroup | Incident | boucle Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-23 09:18 | 2023-06-23 16:50 | 1 | 0 |  | Pour les barrières ATEX je ne suis pas un spécialiste et on a pas vraiment de spécialiste. le plus important est de verifier que tu peux avoir une alim en 11 - 28V DC pour le Y913 a partir de la barrière et respecter les definitions de la barriere selon les zones ATEX d'utilisation (voir ci-dessous). habituellement j'aurais regardé un peu plus loin pour te dire si c'etait bon mais je ne peux pas cette semaine. [Image] Best, Nicolas
INC282140 | Workgroup | Incident | ETTN - Surface flange - -200...+200°C - Jessa Ziekenhuis | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-22 16:48 | 2023-06-22 17:02 | 0 | 0 |  | Yes, it is possible, you can go ahead Best, Nicolas
INC282118 | Workgroup | Incident | Y513 Géoservices | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-22 10:15 | 2023-06-22 10:53 | 0 | 0 |  | Bonjour Cyrille, on ne fait plus du tout le Y513, mais nous avons une gamme complete de transmetteur pour répondre et si cette connexion spécifique est toujours nécessaire, il faudra faire une SPR. Best, Nicolas
INC282102 | Workgroup | Incident | Informations commande 1917929 URGENT | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-22 07:48 | 2023-06-22 09:29 | 0 | 0 |  | Salut Véronique, Voici les details: 11251178 / D918-MD0.3F967.0050000000/Q001_Q003_++++ BP Capillary: protected and white ATEX plastic covered 3m LRS9 11251189 / D920-D3F.6259.Q67.09/Q001_Q003_Q005_+++ HP Capillary: protected and white ATEX plastic covered 3m LRS9 et nos réponses: il y a une légère différence de volume entre le D918 et le D920 du fait de la membrane et de la voûte qui sont différentes mais cette différence ne devrait pas être significative. les capillaires sont identiques et les chambres du transmetteur le sont aussi. les Huiles de remplissage (LRS9) sont les memes. le tirage au vide pour les 2 est fait de la meme façon donc identique a la precision de nos appareils prêt. Pas de précaution particulière (vider en injectant un peu d'air et fermer pour éviter une accumulation d'eau et donc les possibles désagréments a long terme avec la corrosion est possible). Notre conseil est si possible de protéger du soleil les parties sensibles (séparateurs et instrument). les capillaires sont assez protégés par la couche de plastic blanc. si ils veulent aller plus loin dans ce problème, il faudra inclure Jérôme et Christophe Charpentier. Best, Nicolas
INC282099 | Workgroup | Incident | BAIT ¦ Customer Vecom: Geovita> Convertion from E+H | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 15:27 | 2023-06-21 15:37 | 0 | 0 |  | Perfect question, clear and even if it's not easy you had the answer pretty much ready, here is the right mounting: PFMN-54.BC3R.C1571.33212.2501 + D820-10L.42GK.B2001 was there a reason for LRS8 ? Best, Nicolas
INC282098 | Workgroup | Incident | FW: DDP - DFI : BARKSDALE Pressure Switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 14:54 | 2023-06-21 15:32 | 0 | 0 |  | Hi Rudi, it's a pressure switch from barksdale on a SE426 which is a specific seal with the adaption flange enclosed. the material should be A37S with special paint (SP2042). we will not do that anymore as it's so far from standard, but you can find solutions with D4xx or D6xx with exotic materials and coatings. it'll be less expensive for the customer, will be easier to replace (standardised) and can be fitted with an RP2 from us, in terms of quality I replaced this mornings 3 units from 2003 and 1998 still working. Best, Nicolas
INC282090 | Workgroup | Incident | WG: Des-1a - gauge for internal use / offer 20522509 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 13:21 | 2023-06-21 15:41 | 0 | 0 |  | Probably not, as 600kg/cm is equal o 60Mpa (with the small approximation) but it should stop at the same level which is 600/60. Best, Nicolas
INC282089 | Workgroup | Incident | Gauge Protector (Schneider Reference S005-50-222 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 13:00 | 2023-06-21 15:45 | 0 | 0 |  | Hi Daniel, I would like to tell you to get it as an AORP or from our sister company in india but it's probably not the easiest way yet. you can get AORP + ARIN2606 or S005.50.222 2-6 BAR/SETP 11148208 Best, Nicolas
INC282085 | Workgroup | Incident | RPPN-AA8-204 / Gilang | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 10:24 | 2023-06-21 15:47 | 1 | 0 |  | Hi Bee Tin, here it is: Order code: RPPN-7A8.204 RPPN old Code / SAP code: Sensing element and diaphragm: Type: PPN - Medium range -1 ... 600 bar 7 Microswitch: Microswitch 1 standard changeover switch A Hydraulic connection: 1/4" NPT female 8 Pressure range: Pressure range (bar) 0.5 ... 10 bar 204 Best, Nicolas
INC282080 | Workgroup | Incident | WG: Valmet Karawang project inquiry | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-21 09:21 | 2023-06-21 09:34 | 0 | 0 |  | Hi Felix, sure, this is the every angle stem and we have it on TBI for example : [Image] Best, Nicolas
INC282065 | Workgroup | Incident | FW: Rfq | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-20 11:42 | 2023-06-21 09:08 | 0 | 0 |  | Old request.
INC282057 | Workgroup | Incident | TR: TARIF-123828-Baumer SAS-ADV | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-19 17:03 | 2023-06-20 10:21 | 0 | 0 |  | Bonjour Jacques, j'ai bien la question de Gregory, on est en train de regarder, car le code a été fermé car plus de mouvement depuis 2015 mais pas à cause d'une décision de larrêter donc on doit regarder ce qui est possible de faire. Best, Nicolas
INC282054 | Workgroup | Incident | Baumer India Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-19 16:24 | 2023-06-19 16:29 | 0 | 0 |  | Hi Daniel, please find the datasheet in the ticket, the equivalent will be a MEX5. Best, Nicolas
INC282053 | Workgroup | Incident | Tri Clamp connection - Agidens RFQ PALPAL | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-19 15:27 | 2023-06-23 15:59 | 0 | 0 |  | Hi Rudi, I close the request for now, you tell me when the customer give you a feedback. Best, Nicolas
INC282028 | Workgroup | Incident | WG: EIL - Anfrage-Nr.: 05/20230614 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 15:42 | 2023-06-16 16:13 | 0 | 0 |  | Bonjour Felix, You're perfectly right, a 1547 can e mounted with a gauge DN100 starting from 25bar. it'll be 2.5bar with a 1547 which is fine. Best, Nicolas
INC282027 | Workgroup | Incident | TR: [EXTERNAL]: Présentation gamme de capteurs Rail / Baumer | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 12:36 | 2023-06-22 11:31 | 0 | 1 |  | Bonjour Jacques, désolé pour la réponse tardive mais c'est une semaine un peu intense et je voulais vraiment verifier avec pascal avant de répondre car les infos disponibles sont un peu contradictoires. la référence la plus fiable est le PBSN qui a le meme type d'alimentation que l'EF6 et la réelle tension d'alimentation est 8-32V DC. c'est une bonne nouvelle car cela couvre et le besoin de 8-30V DC et le besoin de 10-32VDC. [Image] Tu peux prendre cette valeur pour tous les EF6 et PBSN et je vais demander que les documents relatifs soit corrigés (ce ne sera pas immédiat). Best, Nicolas
INC282025 | Workgroup | Incident | COUPLES DE SERRAGE D920 / 30163895 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 11:42 | 2023-06-16 16:29 | 0 | 0 |  | Bonjour veronique, sur les D920, c'est une bride avec des trous non-filetés et nous ne fournissons pas la boulonnerie. nous ne pouvons donc pas donner d'instructions de montages car elles dépendront de la boulonnerie choisie pas le client. Best, Nicolas
INC282016 | Workgroup | Incident | TR: Fa. Baumer Wasserstoffsensoren | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 09:39 | 2023-06-16 16:29 | 1 | 0 |  | Nathalie, voici un exemple de 2.2 (Q002) avec test helium. [Image] Best, Nicolas
INC282012 | Workgroup | Incident | MCX5-460.N10/0693_0751_9941+ARIN2606AORP0501/SETP / Sakura Proj | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 09:00 | 2023-06-16 14:52 | 1 | 0 |  | Hi Bee Tin, it is possible in this order and it's an AORPB as you are in low pressure: MCX5-460.N10/0693_0751_9941+ AORPB501/SETP + ARIN2606 Best, Nicolas
INC282007 | Workgroup | Incident | MCX5-D60.X69/0751 / Theerasak | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-16 06:15 | 2023-06-16 14:48 | 1 | 0 |  | Hi Bee Tin, sorry, it doesn't exist yet. Best, Nicolas
INC282002 | Workgroup | Incident | Demande de faisabilité / Faiveley | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-15 14:00 | 2023-06-22 14:25 | 0 | 1 |  | Greg, le LBO est validé, Stéphane travaille sur un prix que Jacques recevra directement ou via le service commercial. Best, Nicolas
INC282001 | Workgroup | Incident | Demande MFX3 / Faiveley | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-15 12:18 | 2023-06-15 16:13 | 0 | 0 |  | Salut Greg, étant donné que c'est un code fermé sans mouvement depuis 2023 on ne peut pas le rouvrir juste comme ca. par contre des MFX sont encore fabriqués, donc le but va être de verifier que le plan de celui-ci peut toujours être fabriqué et créer un code au format actuel avec un prix actualisé. merci de faire une SPR demandant la réouverture du code avec les infos importantes de l'affaire. Best, Nicolas
INC281984 | Workgroup | Incident | ROBIVAL - BODENNEC | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 16:57 | 2023-06-14 17:29 | 0 | 0 |  | Bonjour Yannick, l'eau Osmosée n'est pas un probleme. par contre la resistance a la soude et a l'acide nitrique ne se fait pas avec les memes materiaux. vu la faible concentration, je pense que le 316L est le meilleur compromis a proposer. Bien sur, le client reste le garant de la decision finale. un revetement PTFE serait envisageable pour les 75-65 C Best, Nicolas
INC281976 | Workgroup | Incident | TR: Anfrage Friotherm Angebot für 2 Stk Manometer | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 11:51 | 2023-06-14 15:41 | 0 | 0 |  | Bonjour Nathalie, non, on prefere passer le plus possible en MEX3 pour a terme pouvoir enlever les DRO. ce sera un cadran special qui sera peut être à créer par contre pour le R134A surtout en Bar bourdon (sûrement X81).
INC281972 | Workgroup | Incident | Certificat CE ted | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 10:45 | 2023-06-14 11:32 | 0 | 0 |  | Yannick, voici les certificats CE pour les TED 5 et 6. Best, Nicolas
INC281968 | Workgroup | Incident | FW: MOC-20-P0-019 - DDP manomètres - Offre de prix Baumer - TotalEnergies 0020521646 - (RFQ MOC-20-P0-019) | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 10:09 | 2023-06-19 17:52 | 0 | 1 |  | Hi Rudi, I checked with Jerome and we won't give a specific document for this. we remind them that in the REQ ad THR, they rely on our internal document Q1201/CERTIFEP still applicable and accepted today. Best, Nicolas
INC281967 | Workgroup | Incident | MEP 3 -- projet H2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 10:03 | 2023-06-14 11:54 | 0 | 0 |  | Bonjour Yannick, la question a été posée la semaine dernière à David et il a répondu que les 600/1000bar MEP3 seraient débloqués sur l'exercice 2023/2024 donc raisonnablement je tablerai sur debut 2024. Best, Nicolas
INC281965 | Workgroup | Incident | séparateur à membrane et bain d'huile | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 09:48 | 2023-06-14 11:37 | 0 | 0 |  | Adrien, ci-dessous le lien de notre separateur 1020 pour l'industrie du papier, a voir si ca peut convenir au client: https://www.bourdon-instruments.com/fr/en/product-overview/diaphragm-seals/specific-diaphragm-seals/1020/p/37866 Best, Nicolas
INC281959 | Workgroup | Incident | WG: Anfrage Herr Horber | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 08:51 | 2023-06-14 11:43 | 0 | 0 |  | Hi Felix, specific connection is always a possibility but the price entirely depends on the qties. we have an existing drawing for the 1/4VCR-M, to be validate with the customer, then you need to submit an SPR with the qties. [Image] Best, Nicolas
INC281956 | Workgroup | Incident | TR: inquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-14 08:42 | 2023-06-15 09:07 | 0 | 0 |  | Nathalie, j'ai demandé a Franck et on ne va plus proposer les AN4x. Best, Nicolas
INC281950 | Workgroup | Incident | Baumer GmbH - Request for User Manual - ZED545/11D.001 | Archive > BHFR Technical Support | Middle |  |  | Closed | Braun, Cedric [brce] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-13 13:27 | 2023-06-13 15:32 | 0 | 0 |  | Hello Cedric, here is the datasheet (in the ticket attachement), I will not have anything else, this is very old. Best, Nicolas
INC281929 | Workgroup | Incident | WG: Enquiry - New Customer | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-13 07:39 | 2023-06-13 11:04 | 0 | 0 |  | Hi Felix, not with an MTX5 which only for very standard application. but it's possible with an MCX5. Best, Nicolas
INC281919 | Workgroup | Incident | Liquide de remplissage avec séparateur DAxx | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-12 15:27 | 2023-06-12 16:38 | 0 | 0 |  | Salut Cyrille, tu tombe en pleine polémique sur ce sujet, on est pas d'accord entre la qualité et moi. pour l'instant je te propose de donner la fiche de données de sécurité du LRS1 qui contient tous les details sur la composition et les différentes normes auxquelles le fluide répond. si ca ne suffit vraiment pas au client, il faudra faire la demande a la qualité. Best, Nicolas
INC281913 | Workgroup | Incident | WG: 3D model for MEX5 Pressure Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-12 13:30 | 2023-06-12 14:20 | 0 | 0 |  | Hi Felix, sorry it doesn't exist for now, the closest I have is the 1/2NPT. Best, Nicolas
INC281891 | Workgroup | Incident | TR: BAIT ¦ DIREMA offerta MANUF. PART :MTA5-D20.N07 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-09 15:42 | 2023-06-13 15:37 | 0 | 0 |  | Anne, il n'existe pas en standard en G1/4 et il faudra donc faire une commande spécifique de raccords, il faut donc faire une SPR. si il y a peu de qté, ce sera cher et un MCX5 en G1/4 sera existant et sûrement moins cher. Si vous avez de grandes quantités, faire une SPR sur le MTA peut être intéressant. Best, Nicolas
INC281890 | Workgroup | Incident | TR: Global Marine Supplies Spa : REQ 23-439101-SN | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-09 11:57 | 2023-06-12 18:01 | 0 | 0 |  | Anne, après revue avec Christophe, ce serait le code: 11185076 dont voici les details: [Image] Best, Nicolas
INC281878 | Workgroup | Incident | 3D drawing of Y91-3.3.9.B19.R/2160 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-09 07:00 | 2023-06-09 09:10 | 0 | 0 |  | Hi Xu Qiang, I'm sorry but this drawing is not available yet. you can still order it via sales if really necessary. Best, Nicolas
INC281849 | Workgroup | Incident | TR: VALOREL -- DUNCAN | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 14:48 | 2023-06-08 16:17 | 0 | 0 |  | Yannick, seul la membrane est en HC276 (6) ou en 1.4435 (2) selon le choix. le corps (qui comporte la portée de joint) est toujours en 1.4404. sinon il faut faire une SPR pour avoir un corps aussi en matière spécifique. Best, Nicolas
INC281847 | Workgroup | Incident | Fiche Technique level Méca LB3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 13:51 | 2023-06-08 14:17 | 0 | 0 |  | Bonjour Adrien, ce n'est pas de chez nous, c'est indien mais voici la datasheet. Best, Nicolas
INC281846 | Workgroup | Incident | BH5 avec un mano 0…400 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 13:42 | 2023-06-08 14:02 | 0 | 0 |  | Adrien, pas de problème, tu peux y aller. Best, Nicolas
INC281845 | Workgroup | Incident | Y913- G 1/4 A DIN 3852-E (G50) | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 12:54 | 2023-06-08 14:05 | 0 | 0 |  | Hi Daniel, please find enclosed in the ticket a 3D drawing of a Y9 with this connection. Best, Nicolas
INC281844 | Workgroup | Incident | Demande de faisabilité / BOSTIK | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 11:45 | 2023-06-13 15:56 | 0 | 0 |  | Hi Greg, sorry but 2.5bar range is too low for this mounting but it would work with a D435 if material and design are ok for the customer: M31-4A3.B18H 11083753 AKPL-BF5301 10264888 Order code: D435-2L60.4FF.B2001 old Order / SAP code: Flange material: Stainless steel (1.4404) 2 Instrument connection: G 1/2 Female L Diaphragm material : Hastelloy C 276 (HC) 6 Diaphragm Coating : ►Without 0 Standard Flange: EN 1092-1 4 PN/ DN: PN 25 / DN 32 FF Flange face : Raised face RF (Type B) EN 1092-1 B Bottom part material : ►Stainless steel 316L (SS) 2 Bottom housing Coating : ►Without 0 Oil Filling : ►LRS1 (-15+150C) Vaseline 01
INC281835 | Workgroup | Incident | 3D drawing of MEX3-B21.B35/9690_9691_9690 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-08 08:21 | 2023-06-08 10:09 | 0 | 0 |  | Hi Xu Qiang, please find enclosed in the ticket the 3D. but the dial details will not appear on it as it's for mounting information Best, Nicolas
INC281790 | Workgroup | Incident | FDA approved | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-06 13:24 | 2023-06-08 15:36 | 1 | 0 |  | Hi Rudi, for the LRS1, the MSDS enclosed state all properties of the fluid including the conformity to FDA. Best, Nicolas
INC281768 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 15:48 | 2023-06-08 16:43 | 0 | 1 |  | Nathalie, vu avec Giorgio au tel, on a l'accord de maxime pour le faire en 200mm tel que ci-dessous, je te laisse voir pour le prix. TBL100/232.151.70T/9003_0200 Best, Nicolas
INC281763 | Workgroup | Incident | TR: PBMN-2.4.B24R.ASIP.D040.2N06.09/9010- SAP11189714 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 14:00 | 2023-06-05 14:24 | 0 | 0 |  | Anne, pour proposer au plus proche en remplacement, je proposerais un: PBMN-24B24RA11405201000 + ASIPD1N6 + D0432N06 + LRS9 (option montage séparateur et 0.5% de precision du montage). Il y a aussi maintenant des alternatives avec refroidisseur intégré et membrane flush a regarder mais pas en 1/2NPT. Best, Nicolas
INC281762 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 13:45 | 2023-06-05 14:14 | 0 | 0 |  | Nathalie, le TBL est spécifique aux conduites d'air est n'est disponible que dans une configuration spécifique (165mm) mais un TB devrait leurs convenir: https://www.bourdon-instruments.com/medias/__secure__/Baumer_TB100_EN_DS.pdf?mediaPK=8905716727838 Best, Nicolas
INC281759 | Workgroup | Incident | TR: crimped housing mex3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 11:06 | 2023-06-05 11:56 | 0 | 0 |  | Anne, le MEX3 crimped version (0133) est dispo sur le site: https://www.bourdon-instruments.com/fr/en/product-overview/pressure-gauges/industrial-pressure-gauges/mex3/p/38718 il n'existe normalement qu'en type D ou F, pour le type B ou G, ils doivent passer sur le MEX3 standard (ou faire une SPR si il y a des qtés sachant que la production pourrait être assez longue). Best, Nicolas
INC281758 | Workgroup | Incident | Manomètre contact 230502 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 11:03 | 2023-06-05 14:56 | 0 | 0 |  | Bonjour Yannick, on a rien dans la gamme en monel pour le remplacer, il faudra proposer un séparateur. si tu as des quantités fais une SPR, on pense que passer un MS5 en monel en utilisant le tube du MEM5 devrait être possible mais ca demande quand meme quelques essais et une FAI. Best, Nicolas
INC281757 | Workgroup | Incident | TR: BAIT - Comm tech - Y92 + ASIP | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 10:48 | 2023-06-08 16:41 | 0 | 1 |  | Anne, vu au tel avec Francesco, il va proposer un Y91 monté sur un AKPL2F5306, c'est juste pour un problème de température. Best, Nicolas
INC281756 | Workgroup | Incident | Mano - belmar | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 10:39 | 2023-06-05 14:10 | 0 | 0 |  | Bonjour Adrien, on pourrait proposer quelque chose, il va nous falloir les details mais ce serait un MX7 (échelle quadratique à créer) avec 2 ASIP et 2 séparateurs DT3 à priori. Best, Nicolas
INC281748 | Workgroup | Incident | WG: EIL - Anfrage-Nr.: 01/20230526 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-05 09:12 | 2023-06-05 09:39 | 0 | 0 |  | Hi Felix, you're right to ask me. the mounting you need B18 (0...2.5bar) is not possible. as you can see on the email below, I sent you the limits of the mouting and the min pressure range would be 0...6bar (B20). Best, Nicolas
INC281719 | Workgroup | Incident | WG: Help | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 16:03 | 2023-06-07 14:06 | 0 | 1 |  | Hi Felix, then I close the ticket, you contact me when you have a feedback. Best, Nicolas
INC281716 | Workgroup | Incident | TR: BAIT - Comm tech - Y92 + ASIP | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 14:36 | 2023-06-01 16:14 | 0 | 0 |  | Anne, A répondre au client: Y91 + ASIP is ok, but Y92 + ASIP makes no sense, we could offer Y91 + ASIP + D055 then or a PBMN flush with cooling neck. Best, Nicolas
INC281709 | Workgroup | Incident | Pressostat instrumia | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 11:51 | 2023-06-01 14:02 | 0 | 0 |  | Adrien, des essais ont étés fait par Philippe Perrichon et la difference entre un montage vertical et horizontale est assez négligeable, elle ne pose pas de problème. un montage vertical à lenvers par contre peut assez influencer le setp. Best, Nicolas
INC281707 | Workgroup | Incident | WG: 0678 OPTION FOR TSF100 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 11:18 | 2023-06-01 13:59 | 0 | 0 |  | Hello Felix, the option 0678 is the adjustable micrometric pointer (same as on the gauges). but it's no external, you have to open the bezel ring to access it and turn the small screw directly on the pointer to set it. [Image] Best, Nicolas
INC281700 | Workgroup | Incident | PBMN Low Pressure sur séparateur | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 10:24 | 2023-06-01 12:17 | 0 | 0 |  | Bonjour Cyrille, a devrait le faire mais je te propose plutôt un: PBMN-24B17RA15403201000_ASIPD1L3_D055-2L1C_LRS10 Attention, il y aura les options habituelles pour ces montages avec la précision déclassée a 0.5%. Best, Nicolas
INC281695 | Workgroup | Incident | WG: Inquiry_k 202306011714183 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 09:30 | 2023-06-01 16:40 | 1 | 0 |  | Hi Felix, I think the customer needs an AGFR (1/2NPT F/M) Order code: AGFR-0260N.3K060.P AGF old Code / SAP code: Tpye: Stainless steel bored Thermowell R Extension: Without 0 Material: Stainless steel 1.4404 2 Male Threat for flange: 1/2" NPT 6 Thread on instrument side: Fixed 0 female thread on instrument side: 1/2" NPT N Inner Diameter Thermowell: 9 mm 3 Outer Diameter / outer diameter: 16 mm K Plunger lenght: 60 mm 060 Design: Straight shank P This is possible and if the flow is static, it'll withstand 350bar, but if there is flow and the customer need to verify the resistance, we need to make a wake frequency calculation test that you can order via sales. Best, Nicolas
INC281689 | Workgroup | Incident | FW: Pressure switch RFQ | Archive > BHFR Technical Support | Middle |  |  | Closed | Ngiam, Kent [ngik] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-06-01 08:12 | 2023-06-01 12:11 | 0 | 0 |  | Kent, direct customer request, Answered via email to Kent. Best, Nicolas
INC281682 | Workgroup | Incident | MIX7 - lunette avec collerette type B | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-31 15:45 | 2023-06-01 12:07 | 1 | 0 |  | Bonjour Adrien, c'est plus une question pour les ventes, de mon coté les pieces existent, la question est de savoir si votre distri peut les assembler. normalement on ne fait pas de pieces détachées mais on l'a fait avec des gros distris pour les US, par contre ils avaient du materiel de chez nous, étaient formés et cétait un projet construit ensemble. Best, Nicolas
INC281679 | Workgroup | Incident | Manomètre pour mesure niveau dans cuve | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-31 15:00 | 2023-06-05 10:08 | 1 | 0 |  | Salut Adrien, j'ai eu Bernard sur ce sujet, en gros, une solution avec manomètre est possible mais sera très imprecise et aléatoire, il va revoir pour d'autres alternatives je pense. Best, Nicolas
INC281664 | Workgroup | Incident | TR: Urgent request | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-31 09:48 | 2023-06-01 10:06 | 1 | 0 |  | Anne, vu ensemble, si c'est 160mbar pas possible d'avoir une aiguille à maxima, si c'est 0 à 16bar alors possible sur les manomètres standards. Best, Nicolas
INC281647 | Workgroup | Incident | MEX8 Special request | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 13:24 | 2023-06-01 11:46 | 1 | 0 |  | Hi Jürgen, in standard we don't have above 1600bar, the only gauge would be the MFP (High pressure gauge). it's on the catalogue (enclosed) but I don't have any other information, if you can define what you want I can make a request to Jorg. Best, Nicolas
INC281646 | Workgroup | Incident | Question RP | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 13:00 | 2023-06-01 11:40 | 1 | 0 |  | Hi Kerstin, B is "Baisse" in french for Falling Best, Nicolas
INC281645 | Workgroup | Incident | TR: EIL - Anfrage-Nr.: 01/20230526 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 12:09 | 2023-06-01 11:38 | 1 | 0 |  | Nathalie, le mini de montage est 6bar (voir ci-dessous) si il faut étudier une pression inférieure il faut faire une SPR. Best, Nicolas
INC281642 | Workgroup | Incident | Presse-étoupe pour RP2E | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 11:36 | 2023-06-01 11:33 | 1 | 0 |  | Bonjour Cyrille, il est bien inclus dans la livraison, c'est ce code: 10304162 PE 3/4NPT ADE1F2 A BRIDE N5 CAP038777V1 Best, Nicolas
INC281637 | Workgroup | Incident | Pressostat ATEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 11:21 | 2023-06-01 11:35 | 1 | 0 |  | Bonjour Cyrille, c'est le micro S qui doit bloquer, il est possible mais il faut faire la demande pour verifier le prix et le délai. si tu veux un point de consigne spécifique, il faut le preciser. Best, Nicolas
INC281636 | Workgroup | Incident | 0678 Aiguille micrométrique est réalisable sur un MEX5 monté sur DT2 ? | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 11:21 | 2023-06-01 11:23 | 1 | 0 |  | Adrien, pas de problème pour moi tu peux y aller. Best, Nicolas
INC281632 | Workgroup | Incident | TR: Replacement ¦ Zapytanie i 0006011 BAUMER | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 09:54 | 2023-06-01 11:20 | 1 | 0 |  | Anne, voici le remplacement que je proposerais (peut être aussi disponible en PP20H): PBMH-24B19RA11452222000 Best, Nicolas
INC281630 | Workgroup | Incident | TR: RoHS certificate of CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-30 09:45 | 2023-06-01 10:54 | 1 | 0 |  | Nathalie, pour le RoHS, on a cette declaration globale au niveau du groupe: https://www.baumer.com/medias/sys_master/downloads/downloads/h1b/hd2/9014995550238/Baumer-Material-Compliance-Customer-Information-EN-RoC-2021-81128726.pdf et c'est également noté sur le certificat de conformité. Best, Nicolas
INC281613 | Workgroup | Incident | TR: rfq DT2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-26 12:00 | 2023-05-26 12:14 | 0 | 0 |  | Anne, oui c'est possible par contre le code pour le 1"NPTM est le "E". Si c'est pour peu de pieces ca va peut-être être cher ou long a produire, on pourra proposer une option avec adaptateur (connexion standard + ARIN) Best, Nicolas
INC281604 | Workgroup | Incident | TR: Commande EDF N° 5500-4310703865 - div5583 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-26 07:09 | 2023-05-26 10:14 | 0 | 0 |  | Bonjour Veronique, pour leur application le produit est bon (Pstat 20 et delta de 4) le seul problème que je vois est qu'on switch trop pres du 4. il faut leurs rappeler que pour un bon fonctionnement le point a la baisse ne doit pas être plus de 3.68 de delta P. sinon tout me parait bien défini. si la precision n'est pas crucial pour leur application, je conseillerais peut être de passer sur un 245 ca leurs donnerais plus de marge sur le delta P et plus de sécurité pour ne pas abîmer l'appareil en cas de chute trop fort de la pression en LP. Best, Nicolas
INC281556 | Workgroup | Incident | TR: EIL - Anfrage-Nr.: 04/20230523 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-24 16:21 | 2023-05-25 17:40 | 0 | 0 |  | Nathalie, 11263445 est OK 11263445 mais en B77 est OK PBMN-24B77RA14403202/1627/1359/1310 est OK Best, Nicolas
INC281555 | Workgroup | Incident | Deadband of RDN8B6 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-24 16:12 | 2023-05-25 11:27 | 0 | 0 |  | Dear Zhang Wei, it is perfectly possible to set it at: RDN 8B6 224/SETP :H 0.65 bar deadband:0.25bar I checked with rising and falling setp (0.65-0.90reset and 0.65-0.4reset) and both are OK. Best, Nicolas
INC281528 | Workgroup | Incident | MEX5 With R134A Range / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-24 03:30 | 2023-05-24 12:12 | 0 | 0 |  | Hi Bee Tin, here are the dial existing (only the -1...15 exists): [Image] Best, Nicolas
INC281526 | Workgroup | Incident | Barflex | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-23 16:48 | 2023-05-25 08:48 | 0 | 0 |  | Bonjour Cyrille, désolé, la connection intermédiaire UNF n'est plus disponible en standard, si il y a des quantités il faudra faire une SPR a Franck Ferrarese. ci-dessous les codes des connexions possibles: [Image] Best, Nicolas
INC281504 | Workgroup | Incident | STP | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-23 10:30 | 2023-05-23 10:39 | 0 | 0 |  | Hi Jürgen, Here it is: Best, Nicolas
INC281478 | Workgroup | Incident | 08052023 20516822 ¦ Of 00011-0299 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-22 12:21 | 2023-05-23 10:24 | 1 | 0 |  | Veronique, tu peux decliner si il n'y a pas de qtées. Best, Nicolas
INC281477 | Workgroup | Incident | RT2E réglage hystérésis | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-22 12:09 | 2023-05-24 12:14 | 1 | 0 |  | Adrien, je ne vois pas ce que tu veux dire pare repere, on peut s'appeler demain pour ca si tu veux. Best, Nicolas
INC281474 | Workgroup | Incident | TED5 + Seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-22 11:33 | 2023-05-23 10:11 | 1 | 0 |  | Hi Daniel, yes, you're right something is wrong in the code you have for the D2xx but I would have offered the same code as you in replacement. just check they don't have an overpressure we need to take care of because the D208 was a reinforced version up to 100bar (DT8). but right now anyway he's bound to respect the max pressure of the TED B15 which fits a DT2. Best, Nicolas
INC281473 | Workgroup | Incident | MEX3 STP-file | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-22 11:27 | 2023-05-23 10:02 | 1 | 0 |  | Hi Jürgen, here are the 3D drawings in type D, in 1/4NPT and G1/4. Best, Nicolas
INC281451 | Workgroup | Incident | MEX5-D30.B76.DANC01.2L2A.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-17 13:48 | 2023-05-17 17:45 | 0 | 0 |  | Bonjour Adrien, non, il te faut une étendue de 6bar pour un MEX5+DANC01 mais par contre, ca passe avec un MEX3 ou un MX35 (MEX5 avec tube et éléments de MEX3). MEX3-D20.B76.DANC01.2H2A.01 ou MX35-D30.B76.DANC01.2L2A.01 Best, Nicolas
INC281449 | Workgroup | Incident | Mano 40mm laiton | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-17 12:24 | 2023-05-17 17:51 | 1 | 0 |  | Adrien, je ne vois pas vraiment ce qu'on pourrait avoir avec ces options mais on peut sûrement offrir quelque chose (MAT/MIT/MPT/MMX1) il faudrait verifier avec Jörg que ca ne lui dit rien sur une affaire spécifique mais je ne peux pas le faire avant Mardi, avec le pont et un jour de congés. je te conseil quand meme de consulter Jorg au cas ou, ensuite si tu as des problemes a definir un remplacement, a partir de mardi je pourrais t'aider. Best, Nicolas
INC281441 | Workgroup | Incident | Certificate of Origin | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-17 10:42 | 2023-05-17 11:33 | 0 | 0 |  | Hi Daniel, the Declaration of origin can be issued by the chamber of commerce and has a cost, we can issue a certificate of origin internally (see with sales). the EUR1 and the declaration of preferential origin cannot be issued for now, Marting schneider is working on something to have the origin of all component and their weight in the product cost under control to do it. It's still better if you confirm with sales before offering they'll have the details and costs. Best, Nicolas
INC281431 | Workgroup | Incident | 10301232 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 17:03 | 2023-05-16 17:45 | 0 | 0 |  | Erreur de ticketing
INC281432 | Workgroup | Incident | 10301233 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 17:03 | 2023-05-16 17:45 | 0 | 0 |  | Erreur de ticketing
INC281430 | Workgroup | Incident | 10301205 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 17:00 | 2023-05-16 17:45 | 0 | 0 |  | Erreur de ticketing
INC281429 | Workgroup | Incident | 10301194 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 16:57 | 2023-05-16 17:46 | 0 | 0 |  | Erreur de ticketing
INC281422 | Workgroup | Incident | RE: rfq | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 14:39 | 2023-05-16 16:28 | 0 | 0 |  | Anne, ci-dessous les validations de montage: 1 / MEP6-433.V21/0693_Q008_0192_Q003_9941 FITTING TO D2952L2021L OK 2 / MEP6-433.V25/0693_Q008_0192_Q003_9941 FITTING TO D2952L2021K OK 3 / MEP6-433.V25/0693_Q008_0192_Q003_9941 FITTING TO D2952L2021L OK tu as validé que les MEP6 étaient dispos a la vente (et en triplex ou en polycarbonate) ? Best, Nicolas
INC281419 | Workgroup | Incident | Y91 sur séparateur, | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-16 12:15 | 2023-05-16 14:00 | 0 | 0 |  | Bonjour Cyrille, oui, ca passe la precision est déclassée comme tu peux voir dans les commentaires: [Image] Best, Nicolas
INC281399 | Workgroup | Incident | MIRION PROJET MEX3 URGENT | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-15 17:06 | 2023-05-16 11:09 | 0 | 0 |  | Véronique, la faisabilité n'est pas un problème (ce n'est qu'un usinage de raccord et la pression de 10bar avec du OD est O). le mini que nous faisons pour l'instant est du OD6, du OD10 est possible le OD 12 est en général plus utilisé sur le diamètre 100, voir si ca peut passer sur du carré de 14. Peux tu faire une SPR à Céline et on pourra valider le prix, le délai. Best, Nicolas
INC281395 | Workgroup | Incident | TR: ZDPN 211 GSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-15 16:39 | 2023-05-16 11:04 | 0 | 0 |  | Nathalie, ci-joint la datasheet qui fait le lien entre les ZP (anciens) et les RPP (nouveau) et le databook qui contient beaucoups de documentation dédiée au nucléaire. attention des parties sont obsoletes comme l'ED752 ou le DX. les datasheets des produits standard (MEX, AGF, TBI etc...) sont plus a jour sur le site web. Best, Nicolas
INC281393 | Workgroup | Incident | Plan et characteristique techine E913 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-15 15:54 | 2023-05-24 11:58 | 0 | 1 |  | Bonjour, Voici les details pour les 2 codes et la spec en PJ au ticket 102068097 : spécif n 3199007 et plan n 71-059-01 REP101 102068098 : spécif n 3199007 et plan n 71-059-01 REP102 Ces deux références ne sont plus réalisable depuis le passage aux cellules Metallux. Best, Nicolas
INC281390 | Workgroup | Incident | Case TBI100 - UTT Procurement PO-2791 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-15 14:42 | 2023-05-16 10:19 | 1 | 0 |  | Hi Rudi, why not the same in all angles ? it's a more standard case and if a bit more expensive, it might be faster to produce. Best, Nicolas
INC281366 | Workgroup | Incident | EQUIVALENCE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-15 08:21 | 2023-05-16 11:48 | 1 | 0 |  | Véronique, je dois voir avec Jérôme et Christophe charpentier pour cette demande, je l'ai passé sur les ail, et tu es en copie donc je ferme le ticket. est ce que tu pourrais avoir le document client en meilleure qualité stp ? j'ai du mal a tout lire. Best, Nicolas
INC281359 | Workgroup | Incident | TR: manomètre absolu à contacts | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-12 15:24 | 2023-05-12 16:54 | 0 | 0 |  | Bonjour Yannick, le ZPN215SI est un ancien pressostat aujourd'hui remplacé par les RPPxxA3xxx Attention, le code 215 n'a jamais existé. attention a bien revalider toute les caractéristiques, presque tous les composants ont changé depuis 50 ans. Best, Nicolas
INC281353 | Workgroup | Incident | EQUIVALENCE | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-12 12:33 | 2023-05-12 16:33 | 0 | 0 |  | Bonjour veronique, on a pas de transmetteur differentiel, meme coté Baumer. par contre on pourrait probablement fournir les séparateurs. Best, Nicolas
INC281347 | Workgroup | Incident | MTA5-D20.N08 - 10294156 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-12 09:42 | 2023-05-12 16:32 | 1 | 0 |  | Hi Daniel, both are possible, I think it'll be more a question of price. I can't have access to your price, but the 1/4" version might be 1.7x more expensive than the 1/2". Best, Nicolas
INC281335 | Workgroup | Incident | TBI100 - Rain Carbon | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-11 12:36 | 2023-05-17 10:30 | 1 | 1 |  | Hi Rudi, I checked with productionand the stem is double crimped inside and outside the case and it's done on a machine. so no reall reason it untightened. but it also means it's no welded and with strength and if it was before subjected to vibrations etc... it is technically possible to move it. the best is anyway to do a return procedure so that we can take a look. Best, Nicolas
INC281334 | Workgroup | Incident | TR: Y92 2 9G B29 R | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-11 11:42 | 2023-05-11 17:40 | 1 | 0 |  | Nathalie, on peut faire, ci-dessous la validation de Pascal avec le CCR: Bonjour Nicolas Oui on faire Voici le CCR2 , malgré quil soit bloqué les prix ont lair correct. [Image] Salutations Pascal Baudet
INC281328 | Workgroup | Incident | TR: RFQ ARIN | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-11 10:36 | 2023-05-11 11:17 | 0 | 0 |  | Anne, Ca va être du hors standard, il faut faire la demande de prix a sylvain je suppose avec les qtés: ARIN2X0L (3/4NTM) ARIN2X0L (1NPTM) Best, Nicolas
INC281326 | Workgroup | Incident | TR: price MPE6 -30INHG/0/-1/0BAR | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-11 10:27 | 2023-05-11 11:19 | 0 | 0 |  | Anne, ils veulent du MPE6 et je crois que tu as regardé dans le MEP6. le MPE est remplacé par le MPG6 et tu trouveras soit le D soit le A, on en a gardé qu'un. Best, Nicolas
INC281320 | Workgroup | Incident | MEX5-B + AKPL + D055 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-11 09:06 | 2023-05-11 09:48 | 0 | 0 |  | Salut Adrien, oui c'est possible: [Image] Best, Nicolas
INC281314 | Workgroup | Incident | code Bourdon 38377-2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-10 16:51 | 2023-05-11 10:02 | 0 | 0 |  | Adrien, c'est un plan de membrane seule. qui date de 1965 Best, Nicolas
INC281309 | Workgroup | Incident | TR: replacement ¦ BAPL ¦ PFSX ¦ odp ¦ zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-10 15:03 | 2023-05-10 15:25 | 0 | 0 |  | Anne, ce sera le TED5 le plus proche, il a aussi des SETP isolés galvaniquement. les réglages pourront être les mêmes. je t'ai ajouté la datasheet obsolète du PFSX et la notice d'installation du TED. Best, Nicolas
INC281303 | Workgroup | Incident | MPE6-D62.B22 SAP 11249815 replacement | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-10 11:21 | 2023-05-10 15:17 | 0 | 0 |  | Hi Daniel, you can offer an MPG6D62C22, it's the bar option to fit with the ANSI norm the gauge is compliant to. Best, Nicolas
INC281296 | Workgroup | Incident | TR: Pressure switch enquiry 03/05/23 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 16:39 | 2023-05-10 15:14 | 1 | 0 |  | Nathalie, ci-dessous ma réponse, ce sont des très vieilles références et beaucoup de choses étaient manquantes, a bien valider avec le client: Item MODEL/PN/TYPE Description Brand QTY New reference Notes 1 ZPN206SH "BOURDON" MEDIUM RANGE PRESSURE SWITCHES BOURDON 1 RPPY 7C6 206 New standard case is blue and black, please check range, ATEX and connection as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. TYPE ZP, ALL FLUIDS COMPATIBLE WITH THE MEASURING ELEMENT -50 TO +80 DEG.C (STD) -50 TO +200 DEG.C (ZPN-ZPX-ZPC) OPERATING AMBIENT TEMP. -10 TO +55 DEG.C, STORAGE TEMP -50 TO +70 DEG.C, REPRODUCIBILITY +/- (% OF FULL SCALE, READING ACCURACY +/- 5% OF FULL SCALE MODEL : ZPN206SH SCALE ADJ. DEADBAND : 3 -10 BAR RANGE AT TOP AT BOTTOM : 2.5-10BAR 2 "BOURDON SEDEME" THERMOSTATS BOURDON 2 RTNA B 408 2A0C6 Bulb diam. 10mm -> 9.5mm Bulb length 310mm -> 325mm connection, ATEX etc to check with customer if new version fits. SERIES ZT RAGE AT BOTTOM AT TOP : +70/+150 DEG.C. ZTN408SIIINT ,BULB LEN. 310 MM, CAP. LEN 10M 3 ZDPN224SH BOURDON" MEDIUM RANGE DIFF. PRESS. SWITCHES BOURDON 3 RDN8C6224 Please check range, ATEX as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. TYPE ZDP2, ALL FLUIDS COMPATIBLE WITH THE MEASURING ELEMENT FROM -50 TO +80 DEG.C ZDP AND -50 TO +200 DEG.C (ZDPN), OPERATING AMB. TEMP. -10 TO +55 DEG. C, STORAGE TEMPERATURE -50 TO +70 DEG.C, REPRODUCIBILTY: +/- 1 % RANGE, DISPLAY ACCURACY: +/-5 % OF RANGE MODEL SCALE ADJ. DEABAND 0.15-1 RANGE AT BOTTOM 0 0.65-1.3 AT TOP 0.75-1.3 4 ZDPN234SHH 0.15-4 BAR BOURDON RDN8W6234 Please check range, ATEX as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. 0.65-2 0.8-2
INC281290 | Workgroup | Incident | TR: BAIT ¦ INTECH I: Rif. 208467, richiesta di offerta BOURDON (BAUMER) 03/05 0920 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 12:15 | 2023-05-09 14:43 | 0 | 0 |  | Doublon.
INC281289 | Workgroup | Incident | TR: Anfrage 759697 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 11:54 | 2023-05-09 14:53 | 0 | 0 |  | Nathalie, cela depend de si ils veulent une sortie cable ou connecteur, les remplacements possibles sont: ZDPN211CGM Qualified connector / Ultra sensitive* ZDPN211SGM Cable / Ultra sensitive / qualified Best, Nicolas
INC281285 | Workgroup | Incident | TR: BAIT ¦ INTECH Rif. 202268, richiesta di offerta BOURDON (BAUMER) 05/04 1500 7264 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 10:48 | 2023-05-10 15:15 | 1 | 0 |  | Anne, ci-dessous ma réponse, attention ce sont de très vieilles références et beaucoup d'éléments sont manquants, donc bien valider avec le client. Item MODEL/PN/TYPE Description Brand QTY New reference Notes 1 ZPN206SH "BOURDON" MEDIUM RANGE PRESSURE SWITCHES BOURDON 1 RPPY 7C6 206 New standard case is blue and black, please check range, ATEX and connection as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. TYPE ZP, ALL FLUIDS COMPATIBLE WITH THE MEASURING ELEMENT -50 TO +80 DEG.C (STD) -50 TO +200 DEG.C (ZPN-ZPX-ZPC) OPERATING AMBIENT TEMP. -10 TO +55 DEG.C, STORAGE TEMP -50 TO +70 DEG.C, REPRODUCIBILITY +/- (% OF FULL SCALE, READING ACCURACY +/- 5% OF FULL SCALE MODEL : ZPN206SH SCALE ADJ. DEADBAND : 3 -10 BAR RANGE AT TOP AT BOTTOM : 2.5-10BAR 2 "BOURDON SEDEME" THERMOSTATS BOURDON 2 RTNA B 408 2A0C6 Bulb diam. 10mm -> 9.5mm Bulb length 310mm -> 325mm connection, ATEX etc to check with customer if new version fits. SERIES ZT RAGE AT BOTTOM AT TOP : +70/+150 DEG.C. ZTN408SIIINT ,BULB LEN. 310 MM, CAP. LEN 10M 3 ZDPN224SH BOURDON" MEDIUM RANGE DIFF. PRESS. SWITCHES BOURDON 3 RDN8C6224 Please check range, ATEX as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. TYPE ZDP2, ALL FLUIDS COMPATIBLE WITH THE MEASURING ELEMENT FROM -50 TO +80 DEG.C ZDP AND -50 TO +200 DEG.C (ZDPN), OPERATING AMB. TEMP. -10 TO +55 DEG. C, STORAGE TEMPERATURE -50 TO +70 DEG.C, REPRODUCIBILTY: +/- 1 % RANGE, DISPLAY ACCURACY: +/-5 % OF RANGE MODEL SCALE ADJ. DEABAND 0.15-1 RANGE AT BOTTOM 0 0.65-1.3 AT TOP 0.75-1.3 4 ZDPN234SHH 0.15-4 BAR BOURDON RDN8W6234 Please check range, ATEX as I had to guess from the pictures. The references where very old so please check all details with the customer before validation. 0.65-2 0.8-2
INC281271 | Workgroup | Incident | FW: Enquiry For MAR5A30A20MN62, 0-6 bar Pressure Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 04:00 | 2023-05-11 17:32 | 0 | 1 |  | Hi Bee Tin, MAR5A30A20MN62 becomes MAR5A30B20/0840/0151/0412/0670/0751/XXXX the flange is not yet defined and is to be created by R&D, it may induce a slight shift in the overall dimensions in depth. Best, Nicolas
INC281270 | Workgroup | Incident | FW: Enquiry For RP2E+Diaphragm Seal / Sriprapa | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-09 03:39 | 2023-05-09 12:22 | 0 | 0 |  | Hi Bee Tin, I found the first number (only related to the RP2E not the diaphragm seal), and only one code is created in SAP for an RP2E + DT3: RP2-EL.642.DT3-2N20.0.260.01 11163355 (maybe not related to your request) C3039 = RP2E-U06-41 [Image] Sorry, that's all I found. Best, Nicolas
INC281262 | Workgroup | Incident | TMH AMS - MVD9 - SAP 10263042 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-05 15:27 | 2023-05-09 10:52 | 1 | 0 |  | Bonjour Adrien, on a plus du tout de 250mm chez BHFR, en remplacement, on peut offrir un MVE ou MVX. Le cadran spécifique sera a redéfinir pour impression. Best, Nicolas
INC281259 | Workgroup | Incident | MCX5 with range 1 – 3 kPa, G1/2" connection, with sealing ring suitable for Pharma Application | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-05 13:48 | 2023-05-05 15:34 | 0 | 0 |  | Hi Bee Tin, the G1/2 aluminium gasket also exists in 316L, here is the code: 11011130 we could integrate it in the code of an MCX5 but then we need to create an option via switzerland and PM validation. if it's more for a one shot, I suggest to ask the price for a packet of like 100units and provide it next to the order s an accessory. Best, Nicolas
INC281257 | Workgroup | Incident | WG: AW: Pressure Switch RP2Y Cable Entry | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-05 12:54 | 2023-05-05 14:34 | 0 | 0 |  | Hi Sascha, yes, it's this PG9 in polycarbonate: 10309992 PE PG9 POLYCARBONATE Best, Nicolas
INC281256 | Workgroup | Incident | MEX5 With Dial Scale : -760 mmHg to 3 KG/CM2 / Theerasak | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-05 12:36 | 2023-05-05 14:47 | 0 | 0 |  | Hi Bee Tin, it's not existing but it's possible (X76) Best, Nicolas
INC281251 | Workgroup | Incident | TR: RFQ-WARSAK DTT requirement-URGENT | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-05 09:39 | 2023-05-05 11:20 | 0 | 0 |  | Yannick, merci pour le plan, ca veut dire que quelque chose existe (j'avais dit le contraire à Christophe hier) le problème est que meme avec les plans et le numéros, pour l'instant je ne trouve rien et les personne qui ont peut être géré ca ne sont plus la. je vais continuer mes recherches, de ton coté, toute info que tu peux avoir du client sera bonne a prendre. ce que je conseille pour la suite est un meeting avec les achats (vu que les TSFE sont fait par STIKO), la R&D (David) et la qualité (Paul). pour voir si on retrouve quelque chose et determiner la marche a suivre. on peut gérer ca par mail vu que je ne peux pas le faire tout seul, donc je ferme le ticket. Best, Nicolas
INC281232 | Workgroup | Incident | TR: EDF - Mise en œuvre de l'écoconception | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-04 12:03 | 2023-05-09 11:24 | 0 | 1 |  | Bonjour Jacques, on a pas ca encore, mais tu peux en faire la demande à la qualité. Best, Nicolas
INC281209 | Workgroup | Incident | WG: TBI100-251-142.11T | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-03 11:03 | 2023-05-03 12:16 | 0 | 0 |  | Hi Felix, here it is, you can find it pretty easily on the website in the temperature accessories: [Image] best, Nicolas
INC281196 | Workgroup | Incident | MCX5 With Range -0.25 to 0.4 kg/cm2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-03 03:36 | 2023-05-03 09:04 | 0 | 0 |  | Hi Bee Tin, sorry, it's not existing yet, and the code will be F69. Best, Nicolas
INC281191 | Workgroup | Incident | Contacts Tropicalisés - pressostats | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-02 14:33 | 2023-05-03 09:00 | 0 | 0 |  | Bonjour Adrien, nous n'avons plus de contact tropicalisé, je suis désolé. mais normalement nous couvrons autant de plages d'utilisation et voltages. Best, Nicolas
INC281180 | Workgroup | Incident | WG: Customer request: info | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-02 13:06 | 2023-05-03 09:07 | 1 | 0 |  | Forwarded to Jerome Bredon via Email.
INC281175 | Workgroup | Incident | WG: Friction pointer | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-02 11:48 | 2023-05-03 15:28 | 1 | 0 |  | Hi Sascha, indeed it's possible, you order the max pointer and adjust one full turn from "after the pointer" to "before the pointer". it'll from this moment indicate you the lowest reached pressure between 2 resets. Best, Nicolas
INC281174 | Workgroup | Incident | mano diam 160 en -1 ; 1.5 bar avec séparateur clamp DN38mm | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-02 11:03 | 2023-05-02 11:28 | 0 | 0 |  | Bonjour Adrien, tu étais pas loin, les 2 solutions sont: -1...5 bar en MEX8D30B77 / DANC012L2A / LRS1 (DN38) -1...1.5 bar en MEX8D30B74 / DANC022L2A / LRS1 (DN51, aussi possible en DN50) Best, Nicolas
INC281171 | Workgroup | Incident | RP2EL.346 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-05-02 08:45 | 2023-05-02 11:31 | 0 | 0 |  | Hi Bee Tin, you can see this in the marking, as nothing is engraved on the setting area on the device, it means nothing was specified at the creation of the code. it's still a good idea to check with them if they want us to set a setpoint (a code will be created ith the specific setp required). Best, Nicolas
INC281166 | Workgroup | Incident | TR: Temperaturfühler RT2E Atex | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-28 14:24 | 2023-04-28 14:30 | 0 | 0 |  | C'est ce code, mais il faudra le recréer avec la definition du SETP: 11014628 RT2-EU.43.C03C6/SETP Best, Nicolas
INC281153 | Workgroup | Incident | WG: AW: C3/ Low Steam Pressure Alarm | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-28 11:36 | 2023-04-28 11:55 | 0 | 0 |  | Hi Felix, it's possible but G3/8 is a specific connection, you need to ask for the price. Best, Nicolas
INC281145 | Workgroup | Incident | RACCORD TOURNANT | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-28 10:12 | 2023-04-28 14:19 | 0 | 0 |  | Je vois, c' est le pire cas de figure, pas la meme norme, pas la meme taille et pas M/F. je pense que prendre un raccord tournant + 2 ARIN va être cher et il n'est jamais bon de multiplier les accessoires. J'ai regardé chez Schneider mais meme eux ne font pas ce raccord tournant. il faudrait prendre un GS-MMSP-N2N4 (Schneider) + ARIN230N (l'ADV pourra demander un prix) mais à mon avis elle à aussi vite fait de racheter un manifold. Best, Nicolas
INC281139 | Workgroup | Incident | DDP 2304RSR116 solyro | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-28 09:09 | 2023-05-09 11:25 | 0 | 1 |  | Yannick, sans réponse, je cloture le ticket mais je reste dispo dès que tu auras des nouvelles. Best, Nicolas
INC281136 | Workgroup | Incident | BAPL ¦ RP2NR644 ¦ setting the set point ¦ Pytanie - set point RP2NR644 | Archive > BHFR Technical Support | Middle |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-27 18:12 | 2023-04-28 16:20 | 1 | 0 |  | please find enclosed the manual, there is instructions on how to set but the scale is really a rough indication, it's not a good idea to set it visually. our advise is to set it with calibrated pressure as it should be done. Best, Nicolas
INC281133 | Workgroup | Incident | request for a differential pressure gauge ¦ Saint Gobain - Manometr różnicowy | Archive > BHFR Technical Support | Middle |  |  | Closed | Furmańczyk, Anna [anf] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-27 16:06 | 2023-04-27 16:17 | 0 | 0 |  | Hi Anne, both devices have 316L (1.4404) parts in contact with the fluid. this is a common material used in the oil and gas market and even if it's complicated to say without having the concentration, state, temperature etc..., it seems like a good usual choice for this type of applications. Best, Nicolas
INC281122 | Workgroup | Incident | 3D STP CAD Files required | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-27 10:15 | 2023-04-27 14:06 | 0 | 0 |  | Hi Daniel, here is a 3D of the MA7, the MDF7 is phased out and old so no 3D but I could probably find a drawing if I had the complete codifications. Best, Nicolas
INC281118 | Workgroup | Incident | Certificate of conformity | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-27 09:09 | 2023-04-27 14:10 | 1 | 0 |  | Hi Felix, for a connector like this, a Q001 2.1 conformity can be requested or a Q003 3.1 material certificate. please find enclosed the certificates overview for a bit more details. Best, Nicolas
INC281100 | Workgroup | Incident | Spare part for MEX3 10358001 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-26 11:42 | 2023-04-27 14:34 | 1 | 0 |  | Hi Felix, here is a picture with the code of the clamp and the 2 screws, what it looks like and how it's packaged. Nathalie will be able to define a price for this. Best, Nicolas
INC281095 | Workgroup | Incident | DDP DEXIS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-26 08:12 | 2023-04-26 09:10 | 0 | 0 |  | Bonjour Véronique, daprès ce que je vois, ce serait un simple manomètre mais avec graduation spéciale (double bar/tonne) ce qui est possible sans problème, et comme ils le demandent en 1/4, je commencerais par proposer un MEX3. Best, Nicolas
INC281093 | Workgroup | Incident | TR: Prix IC ? - MR5-4D33.B19.D621-2L7.24CC.BF002 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-25 17:39 | 2023-04-25 18:02 | 0 | 0 |  | Nathalie, oui c'est bon pour moi. Best, Nicolas
INC281091 | Workgroup | Incident | TR: Cross old thermometer | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-25 17:00 | 2023-04-25 17:47 | 0 | 0 |  | Nathalie, I have the same request from India and didn't found anything in the system, web or datasheets. I'm sorry it doesn't come from here and anyway we don't have a replacement. Our TE2, ETTN or other could match (with a different input). Best, Nicolas
INC281087 | Workgroup | Incident | Square Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-25 15:21 | 2023-04-25 17:52 | 0 | 0 |  | Hi Daniel, we don't do anymore square gauges. And I check in the Indian catalogue and only found this, not sure it fits: [Image] Best, Nicolas
INC281086 | Workgroup | Incident | Activate Kmat no. 96003138 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-25 15:09 | 2023-04-25 15:17 | 0 | 0 |  | Hi Rudi, I didn't understood the request, I cannot do anything on the KMAT numbers to make it available to SUs. I thought you had some configurations of RP2 + Dxx to verify Best, Nicolas
INC281068 | Workgroup | Incident | WG: Lastenheft H2-Manometer / P-ELO-22556 - WEH GmbH | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-25 08:39 | 2023-04-25 14:10 | 0 | 0 |  | Hello Felix, the H2 applications are a bit tricky because no norm is regulating it for now. what we can tell customer is that the thickness of the tube and the whole 316L construction seems to not leak and the MEP5 is approved by air liquide for H2 applications. the tube connection will not be possible because we can't test it with high pressure. I don't have a solution for now on orientable connection, we would advise Gas connection or metal to metal cone tightness (SAE connection for example). At the end, for now, only the customer will be responsible for the choice of what is used on H2 Best, Nicolas
INC281061 | Workgroup | Incident | ted6 0...1bar + d055 (1''1/2) | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-24 14:45 | 2023-04-24 14:49 | 0 | 0 |  | Bonjour Adrien, je pense qu'on va pouvoir accepter, mais ca reste hors specs et je ne sais pas si on va préférer passer par un TED flush ou TED+D055. peux tu faire une SPR vite fait et me l'envoyer pour qu'on fasse une validation officielle stp ? Best, Nicolas
INC281058 | Workgroup | Incident | MR5 With D425TA_D621 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-24 13:30 | 2023-04-24 14:15 | 0 | 0 |  | Hi Jürgen, the MR5 seems fine, and I would suggest to offer the D6xx below. if the customer is working with ANSI/ASME flange norm, I also have a corresponding flange but we will have an issue with the 4bar gauge for the mounting. Also I can't be sure of the flange material and potential coating I don't see, this is to validate with the customer: MR5-4D33.B19 + D621_2L7.24CC.BF002 LRS2 to validate, I'm not sure why you preferred it ? Best, Nicolas
INC281040 | Workgroup | Incident | manomètres décalés | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-24 09:09 | 2023-04-24 14:26 | 1 | 0 |  | Bonjour Cyrille, la premiere chose à faire va être de les remettre à la verticale, d'enlever lévent et de tapoter le boîtier pour recaler l'aiguille. Si cela ne recale pas l'aiguille sur le 0, il faudra voir pour un SAV ou une re-calibration selon la garantie avec l'adv. Best, Nicolas
INC280983 | Workgroup | Incident | AW: BAPL ¦ replacement ¦ 10329534 ¦ ED 711/L20.414.A25.11A ¦ AKP_INQ_20230419_0745_Ntech | Archive > BHFR Technical Support | Middle |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-20 14:48 | 2023-04-20 14:57 | 0 | 0 |  | solved
INC280981 | Workgroup | Incident | Rückruf: BAPL ¦ replacement ¦ 10329534 ¦ ED 711/L20.414.A25.11A ¦ AKP_INQ_20230419_0745_Ntech | Archive > BHFR Technical Support | Middle |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-20 14:45 | 2023-04-20 14:58 | 0 | 0 |  | Doublon
INC280976 | Workgroup | Incident | AW: BAPL ¦ replacement ¦ 10329534 ¦ ED 711/L20.414.A25.11A ¦ AKP_INQ_20230419_0745_Ntech | Archive > BHFR Technical Support | Middle |  |  | Closed | Ammann, Christof [amc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-20 13:00 | 2023-04-20 14:59 | 0 | 0 |  | Solved
INC280975 | Workgroup | Incident | raccord BARFLEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-20 12:00 | 2023-04-21 16:43 | 0 | 0 |  | Yannick, désolé pour la réponse tardive, ci-joint les photo des flexible disponible, on pourra faire un point lundi avec les plans Best, Nicolas
INC280959 | Workgroup | Incident | TR: BAPL ¦ replacement ¦ 10329534 ¦ ED 711/L20.414.A25.11A ¦ AKP_INQ_20230419_0745_Ntech | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-19 16:57 | 2023-04-19 18:28 | 0 | 0 |  | Celine, c'est un transmetteur flush, il pourra être remplacé par un PBMN flush ou meme peut être un PP20, je ne connais pas bien ces produits (les PP20). si ca ne leurs va pas, dit leurs de m'appeler directement. Best, Nicolas
INC280956 | Workgroup | Incident | TR: Certificat CE | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-19 16:03 | 2023-04-19 16:39 | 0 | 0 |  | voici les dernières en date, tu peux les trouver ici: Z:\Public\Conception\Homologations produits\Déclaration CE Best, Nicolas
INC280950 | Workgroup | Incident | item 10344097 | Archive > BHFR Technical Support | Middle |  |  | Closed | Korner, Carla [koca] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-19 12:57 | 2023-04-20 15:21 | 0 | 0 |  | Hi Carla, I discussed with Sascha this morning and made the analyse and we can't replace the DRO by another product. this is due to the very specific product which mean that the new MEX3 will be special also and we need to run it through the codification of a complete specific product in SAP and a customer validation. this is a lot of work and time and with checking the lead times, it'll not be much of a time gain for you at the end. Best, Nicolas
INC280920 | Workgroup | Incident | TR: Updated Pricing // 1623 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-18 10:27 | 2023-04-18 12:00 | 0 | 0 |  | Celine, oui, ça correspond, le code en STUA est juste pour le corps et le FUD après est avec l'option 1359 1623 = J70439.0351 Best, Nicolas
INC280916 | Workgroup | Incident | WG: inquiry_k 202304180916 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-18 09:51 | 2023-04-18 12:03 | 0 | 0 |  | Hi Felix, The principle is that we only sell the complete instrument + diaphragm seal mounting. with some distributors that have the capabilities for vacuum filling on their own we offer only the seal but it needs to be validated by sales management (Carlo maybe, or Jörg). Best, Nicolas
INC280911 | Workgroup | Incident | BSE/TSE Certificate | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-17 16:33 | 2023-04-17 18:08 | 0 | 0 |  | Hi Andrew, our DANC are mettalic solid parts, so not concerned, and the filling fluid isn't related to animals in any way so also not concerned. Best, Nicolas
INC280908 | Workgroup | Incident | FW: aanvraag | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-17 16:03 | 2023-04-18 14:36 | 0 | 0 |  | Hi Jürgen, in terms of low cost / small contact gauges we don't have anything to compete. for the 0...76cmHg, then it's weird because the height of Hg is supposed to be used exclusively for vacuum. I think there is a misunderstanding somewhere. Best, Nicolas
INC280884 | Workgroup | Incident | WG: Anfrage Anschlussplan MR5 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-17 14:00 | 2023-04-17 15:03 | 0 | 0 |  | Hello Felix, please find enclosed the manual. Best, Nicolas
INC280857 | Workgroup | Incident | TR: P/N 16923149 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-17 09:03 | 2023-04-17 12:14 | 0 | 0 |  | c'est ce joint: 10305228
INC280854 | Workgroup | Incident | WG: Pürfung Flanschschutzrohr gemäß DIN 43772 Labom | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-17 08:33 | 2023-05-03 12:07 | 1 | 1 |  | Hi Sascha, I send the offer from the supplier to Nathalie so she can give you a price. Best, Nicolas
INC280853 | Workgroup | Incident | TR: P/N 16923149 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 16:30 | 2023-04-14 16:59 | 0 | 0 |  | Nathalie, c'est le joint alu pour le 1/2G a mon avis, ci-joint le plan a leurs faire valider et voici le code: 10305228 [Image] Best, Nicolas
INC280852 | Workgroup | Incident | doigt de gant AGF | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 16:09 | 2023-04-14 17:02 | 0 | 0 |  | Salut Cyrille, je ne vois aucun problème, tu reste au dessus des 3mm dépaisseur et le reste est standard, c'est bon pour moi. Best, Nicolas
INC280848 | Workgroup | Incident | TR: Inqury_k 202306120850 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 12:09 | 2023-04-14 15:27 | 0 | 0 |  | Nathalie, de ce que je comprends il ne trouve pas les prix, la dessus je te laisse faire. oui en standard le remplissage sera fait en LRS1. Best, Nicolas
INC280847 | Workgroup | Incident | D820 Diaphragm Seals | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 12:09 | 2023-04-14 16:46 | 1 | 0 |  | Yes, it seems that wika can go lower than we do on this type of seals. please use the D821 below 10bar. I'm sorry, we are investigating our diaphragm seals competition and the input you sent was really valuable (we also didn't had the low pressure limit on their 990.27). Best, Nicolas
INC280844 | Workgroup | Incident | Certificat CE Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 11:06 | 2023-04-14 11:18 | 0 | 0 |  | Bonjour Jacques, voici le certificat CE du Y91. Best, Nicolas
INC280843 | Workgroup | Incident | Wetted Part Material For MCD7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-14 11:03 | 2023-04-14 15:25 | 0 | 0 |  | Dear Bee Tin, yes the wetted parts are in 316L. Best, Nicolas
INC280825 | Workgroup | Incident | TR: EIL - Anfrage-Nr.: 01/20230413 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 15:48 | 2023-04-18 16:04 | 0 | 0 |  | Nathalie, it's ok for the fittings, here is the price of the 175 23 1 5 (Clamp DN20): 10333965 690 CCR2 (basé sur 2 pieces) Best Nicolas
INC280823 | Workgroup | Incident | TR: Anfrage ED701L20/B5-525-11 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 15:15 | 2023-04-13 15:36 | 0 | 0 |  | Hello Kevin, please find enclosed the datasheet of the ED701, I would suggest to replace it by a PBMN flush or PP20H. Best, Nicolas
INC280822 | Workgroup | Incident | TR: gaskets for PBSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 15:03 | 2023-04-13 15:31 | 0 | 0 |  | Hi Nathalie, the one we offered is the external gasket for the 1/4 connection, the internal one is this one: JT TOR 12,48X1,78 EPDM SVT NT1935 10303427 Best, Nicolas
INC280817 | Workgroup | Incident | TR: gaskets for PBSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 13:48 | 2023-04-13 14:09 | 0 | 0 |  | Nathalie, the connection gasket on the 11166832 is the : 10303406 Joint ED DIN3852 forme EEPDM peroxyd Best, Nicolas
INC280816 | Workgroup | Incident | TR: Inqury_k 202306120850 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 13:45 | 2023-04-13 14:16 | 0 | 0 |  | Hi Nathalie, Jessica, we could offer 2 solutions, with Y913 or with PBMN. Y91333B19R/2160 / 1621 (1%) PBMN-24B19RA15303201000 / 1621(0.5%) Y913 solution is cheaper and with quick delivery time but less accurate. Best, Nicolas
INC280785 | Workgroup | Incident | MR5-2163.D22+AKPL-10N602+D415-2N2.0214.A10/9030 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 08:30 | 2023-04-13 11:56 | 0 | 0 |  | Hi Bee Tin, this is OK, but with a D425, the D415 is not available in 1" 150. Best, Nicolas
INC280782 | Workgroup | Incident | Maximum Pressure Rating for DT1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-13 07:45 | 2023-04-13 11:01 | 0 | 0 |  | Hi Bee Tin, you can find it on the codification page of the datasheet: [Image] Best, Nicolas
INC280772 | Workgroup | Incident | 3D CAD Step Files | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-12 14:33 | 2023-04-12 16:56 | 0 | 0 |  | Hi Daniel, please find enclosed the 3D file for the MEX, we don't have it for the TB (you can find the dimensional on the datasheet). Best, Nicolas
INC280771 | Workgroup | Incident | TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-12 14:12 | 2023-04-12 16:58 | 0 | 0 |  | Bonjour Cyrille, je suis désolé mais on a pas de 3D disponible pour celui ci, tu trouveras le dimensionnel sur la datasheet. Best, Nicolas
INC280767 | Workgroup | Incident | MX7-D30.B20K - option 0059 : Aiguille à maxima, avec remise à zéro Bouton fixe | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-12 11:00 | 2023-04-12 11:18 | 0 | 0 |  | Bonjour Adrien, Non, je suis désolé, il n'y a pas l'option 0059 de disponnible, j'ai l'option 0052 qui est une aiguille index (pas a maxima) reglable par un bouton sur le voyant polycarbonate. Best, Nicolas
INC280756 | Workgroup | Incident | Low pressure Hg scale | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-12 09:21 | 2023-04-12 11:26 | 0 | 0 |  | Hi Jürgen, a MEX3 is possible for -76cmHg but the dial needs to be created (fee) but it exists in -760...0mmHg (X59) already. concerning the differential we could do a -5...5bar (X22) but the dial needs to be created, a -4...4 would need to go to R&D for approval. Best, Nicolas
INC280751 | Workgroup | Incident | MEX 3 With Special Dual Scale | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-12 04:33 | 2023-04-12 10:47 | 0 | 0 |  | Dear Bee Tin, you can go ahead, the dial will need to be created but it's possible, the code will be X1E Please make sure you ask for all the necessary details for the creation (colors, equivalence, fluid number, each scale etc...) Best, Nicolas
INC280749 | Workgroup | Incident | BAIT - Intech ¦ RDO RIF 11041107 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-11 17:48 | 2023-04-11 18:19 | 0 | 0 |  | Hi Erald, I don't give prices but I can give you the SAP code for the AKPL and I corrected your D4xx, it can't be a D415 at this rating it's a D425 but none is existing in steel so a price request via Nathalie is necessary.. 11006102 AKPL10N602 D425-2N2.0214.A10 Best, Nicolas
INC280748 | Workgroup | Incident | TR: MIX7/D820 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-11 16:09 | 2023-04-11 16:19 | 0 | 0 |  | Oui, le code et le montage sont OK.
INC280746 | Workgroup | Incident | ExxonMobil (55203668) - Angled 60 degree connection | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-11 15:42 | 2023-04-11 16:16 | 0 | 0 |  | Hi Rudi, for 1/2 NPT, here it is it's the 138, but for the 1/2BSP with cone, it would need to be created so via an SPR if necessary because as far as I can see it's not existing in the drawings. [Image] Best, Nicolas
INC280740 | Workgroup | Incident | WG: FW: AKP INQ 20230125 1655 Viess 20503941_Rectus | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-11 12:51 | 2023-04-11 16:22 | 0 | 0 |  | Hi Carlo, here are the .stp files for MEX5 D3X and MCX5D3X Best, Nicolas
INC280710 | Workgroup | Incident | RP2-EU.654 - Louwers Mechanical Parts | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-07 14:24 | 2023-04-07 16:31 | 0 | 0 |  | Hi Rudi, with this type of issue, I would advise to first send them a procedure for doing the setting and if they can't do it, we can plan a video call but only directly with the lab guy to guide him. after that if there still is an issue, we can start a return procedure for calibration but there will be a cost. Best, Nicolas
INC280703 | Workgroup | Incident | 160 mm gauges, type G | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 17:21 | 2023-04-11 09:01 | 0 | 0 |  | Hi Jörg, there is a misunderstood we can use our MEX8 and the back piece of the indian one, their 150 has the same back size as our MEX8. Best, Nicolas
INC280701 | Workgroup | Incident | SDI 439 / Theerasak | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 12:33 | 2023-04-06 14:25 | 0 | 0 |  | Hi Bee Tin, those products are really difficult to define as it's all the same except the dial and the identification is behind the dial (genius idea). but to help you I can give you the codes that could match depending on the dial and their details (see file enclosed). otherwise, you can ask the customer for marking's pictures behind the dial. Best, Nicolas
INC280700 | Workgroup | Incident | TR: BAPL ¦ DPC 100/813.133.116 (0…60 mbar) ¦ DPC 100/816.133.116 (0…60 mbar) ¦ availability | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 12:12 | 2023-04-06 14:45 | 0 | 0 |  | Nathalie, seule la première est possible (c'est sur la datasheet). DPC 100/813.133.116 (060 mbar) pour le prix si tu as besoin des references similaires sont en RL: 11104180 DPC100-816.133.116./0078 Best Nicolas
INC280697 | Workgroup | Incident | niveau de sécurité SIL sur un M33 / 20503732 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 11:36 | 2023-04-06 11:54 | 0 | 0 |  | Bonjour Adrien, nous n'avons pas demandé l'homologation SIL (sur aucun de nos produits). en terme de securité, le M3 est ATEX, ci-joint le certificat ATEX. Best, Nicolas
INC280689 | Workgroup | Incident | Scale on pressure gauges, thermometers? | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 09:15 | 2023-04-06 10:22 | 0 | 0 |  | Hi Rudi, You don't know, that's one of the gaps we have in the product informations. I worked on something and will present it shortly for approval but then it have to go through marketing so it's not for tomorrow. here are your scales in bar EN for diameter 100mm [Image] -1...1 is non standard, please confirm me the gauge you need this on and I'll ask dial creation (via teams). Best, Nicolas
INC280687 | Workgroup | Incident | 10273212 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 08:12 | 2023-04-06 11:07 | 0 | 0 |  | Bonjour Véronique, oui je suis au courant de l'existance de ce code, mais il est bloqué. De plus vu la date de création et la dernière vente, le cadran ne peut pas avoir été créé sur l'imprimante donc il serait de toute façon a recréer. Une capsule existante pourrait être utilisé mais nous devrons valider la classe de précision, il faudra peut être faire une SPR avec les détails du marché (je pourrais voir avec nico selon les capsules on pourrait peut être faire sans). Best, Nicolas
INC280686 | Workgroup | Incident | WG: CAD Files | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-06 08:12 | 2023-04-06 11:26 | 0 | 0 |  | Hello Felix, the TB is not existing in 3D with bottom connection but please find enclosed the MEX drawing. Best Nicolas
INC280682 | Workgroup | Incident | Connecteur DIN PBSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-05 16:18 | 2023-04-05 16:55 | 0 | 0 |  | Bonjour Jacques, oui il est bien livré avec. Best, Nicolas
INC280681 | Workgroup | Incident | TR: Request for IC price (Busse 02/20230321) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-05 16:12 | 2023-04-19 10:48 | 0 | 1 |  | Nathalie, voici les prix fait avec sylvain: PBMN-24B18RA14403201000 / 1627_1359_LRS1 PBMN: 234.3CHF IC web 1627.1359: 11223084 282 LRS1: 34.8 IC P2F PBMN-24B79RA14403201000 / FÜD 175/19/1/5_ LRS1 PBMN: 234.3CHF IC web FÜD 175/19/1/5: 10340664 CCR2: 343.15 LRS1: 34.8 IC P2F Best, Nicolas
INC280679 | Workgroup | Incident | Doigt de gant foré | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-05 14:51 | 2023-04-05 15:11 | 0 | 0 |  | Adrien, on peut probablement le faire, il faudra envoyer la demande à nos 2 fournisseurs avec les quantités à produire. ils valideront et nous donneront une offre. à partir de ca, nous pourront t'offrir un prix avec une désignation (probablement spécifique) pour que tu puisse faire une offre. pour un petit nombre de pièce, les prix seront très élevés, je te conseillerais donc de partir sur un equivalent standard (AGFR02EN1G070(100)P) qui se montera normalement en lieu et place (les longueurs d'insertion sont a confirmer car j'ai compensé la difference dépaisseur de fond). l'hexagone passe de 36mm à 27mm et lépaisseur du plongeur est 3mm au lieu d'environs 1.5mm. Best, Nicolas
INC280671 | Workgroup | Incident | TR: replacement ¦ BAPL ¦ PFSX ¦ odp ¦ zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-05 12:30 | 2023-04-05 14:04 | 0 | 0 |  | Nathalie, l'alternative serait un TED-5.3.5.B77.R (ou TED6) ci-joint la Datasheet du PFSX si le client veut comparer. Best, Nicolas
INC280666 | Workgroup | Incident | TR: Is this a standard older product for which there is a replacement, I can't see that it exists today and would you be able to make a new type that replaces the old no MR-305. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-05 11:27 | 2023-04-05 11:32 | 0 | 0 |  | Nathalie, the MR designation is not an MR5 from us, it's a model from REFCO, here are the details, a MEX3 could be a replacement: [Image] Best, Nicolas
INC280655 | Workgroup | Incident | BAIT ¦ intech Rif. 201311, richiesta di offerta BAUMER 31/03 1118 19676 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-04 16:39 | 2023-04-05 10:03 | 0 | 0 |  | Hi Erald, here is the code and it's still open: 10296548 RDN-4E6.131/SETP Best, Nicolas
INC280646 | Workgroup | Incident | RE: EF6 Faiveley | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-04 10:06 | 2023-04-04 10:23 | 0 | 0 |  | Answered via email.
INC280634 | Workgroup | Incident | DEMANDE ORANO 20511717 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-03 16:15 | 2023-04-04 09:05 | 0 | 0 |  | Bonjour Veronique, ce n'est pas possible, il faut proposer soit un N68 (-160...250mbar) soit un N69 (-250...400mbar). Best, Nicolas
INC280607 | Workgroup | Incident | TR: Y9223GB29R02010-100 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-03 10:42 | 2023-04-03 14:18 | 0 | 0 |  | Nathalie, nous pouvons mais des choses ont changées depuis l'ancienne datasheet et la nouvelle est pas encore prête. elle peut le définir avec la datasheet existante avec son client et vérifier avec Franck qui donnera son approbation (ou via moi). Best, Nicolas
INC280598 | Workgroup | Incident | 3D drawing of Y91363B22A | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-04-01 07:00 | 2023-04-03 14:22 | 1 | 0 |  | Hello Xu Qiang, I have it in G1/2 (3) not in 1/2NPT(6). we work with step files, which is the most common standard, we don't have other yet. Best, Nicolas
INC280596 | Workgroup | Incident | BAIT ¦ Valvitalia | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-31 16:09 | 2023-03-31 16:38 | 0 | 0 |  | Hi Erald, yes we have the MFE for example, but most of the products for railway are specifically design with the customer. please send me he information you have via email to see how to handle that. Best, Nicolas
INC280586 | Workgroup | Incident | certification IEC for y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-31 14:57 | 2023-03-31 16:08 | 0 | 0 |  | Hi Angel, please find enclosed the certificate: Best, Nicolas
INC280552 | Workgroup | Incident | FW: [Inquiry] Pressure gauge manufactured by “bourdon-instruments”/ Mizusawa | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-31 04:39 | 2023-03-31 09:05 | 0 | 0 |  | Hi Bee Tin, I'm pretty sure it's this code except I don't see the marking for x0.1MPa, was it added ? if you need the marking, you can ask fir it to be added to the code as a "technical marking" otherwise this is the code: 11171107 Best, Nicolas
INC280551 | Workgroup | Incident | 3D drawing of Y91333B79R | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-31 03:51 | 2023-03-31 09:07 | 0 | 0 |  | Hi Xu Qiang, here it is: Best, Nicolas
INC280550 | Workgroup | Incident | BAIT ¦ INTECH 199969, richiesta di offerta BOURDON (BAUMER) | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 19:51 | 2023-03-31 11:03 | 0 | 0 |  | Hello Erald, I had this also from singapore, from the picture, here is what I can tell (my answer to BASG below), but your customer's definition seems fine: Dear Bee Tin, here is what I can get from your pictures: Diaphragm seal: D4xx series 1" 150 RF ANSI 1.4404 body Contact gauge: 100mm 2 contacts 0...1000kpa (10bar) 2m capillary with stainless steel protection (maybe plastic coating) and 1/2G in between the components. I miss those important points that we can't deduct from the picture: types of contacts and NO/NC diaphragm material, coating this will need to be confirmed with the customer Best, Nicolas
INC280542 | Workgroup | Incident | TR: Baumer YTED pressureswitch question | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 14:33 | 2023-03-30 17:54 | 0 | 0 |  | Hi Angel, no the YTED cannot be reactivated but the TED is still available even if no more advertised on the website (see note from frank). if any doubt on this don't hesitate to use the technical support directly. Best, Nicolas
INC280535 | Workgroup | Incident | CN1 NAVAL | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 11:00 | 2023-03-30 11:32 | 0 | 0 |  | Véronique, sur le Naval, Nicolas est bien plus à meme que moi de répondre. de mémoire, je ne suis pas surpris par ces options il me semble que les manos marine étaient proche du standard et ces options sont typiques. j'ai quand meme relancé Nico et je vais passer le voir. Best, Nicolas
INC280529 | Workgroup | Incident | PLANS 3D | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 09:27 | 2023-03-30 10:18 | 0 | 0 |  | Bonjour Veronique, je suis désolé mais on a pas de 3D pour les plans d'ensemble. sachant que je vois par contre une ennorme augmentation des demandes de plan 3D, je vais faire remonter le sujet, on est surement pas encore assez préparés. Best, Nicolas
INC280527 | Workgroup | Incident | FW: Inquiry- BOURDON HAENNI /JWCT | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 09:03 | 2023-03-30 10:25 | 0 | 0 |  | Dear Bee Tin, here is what I can get from your pictures: Diaphragm seal: D4xx series 1" 150 RF ANSI 1.4404 body Contact gauge: 100mm 2 contacts 0...1000kpa (10bar) 2m capillary with stainless steel protection and 1/2G inbetween the components. I miss those important points that we can't deduct from the picture: types of contacts and NO/NC diaphragm material, coating this will need to be confirmed with the customer Best, Nicolas
INC280525 | Workgroup | Incident | PLAN DAO | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-30 08:06 | 2023-03-30 10:19 | 0 | 0 |  | Bonjour Veronique, le voici. Best, Nicolas
INC280521 | Workgroup | Incident | DANC01-2L2B | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 15:30 | 2023-03-29 16:48 | 0 | 0 |  | Hi Andrew, I confirmed with the girls it's possible to do, but you need to ask Nathalie for this. Best, Nicolas
INC280519 | Workgroup | Incident | TR: DELPHARMS TOURS - BAUMER - Docs et certifs | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 14:12 | 2023-03-29 15:29 | 0 | 0 |  | Bonjour Yannick, je ne comprends pas, l'electropolissage et le Ra qui est la valeure critique apparaissent sur le Q002: [Image] Best, Nicolas
INC280515 | Workgroup | Incident | ETTN sans silicone | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 12:33 | 2023-03-30 17:52 | 0 | 0 |  | Bonjour Cyrille, je suis désolé mais du silicone est utilisé pour fixer l'élément sensible dans la gaine. On ne peut donc pas faire d'ETTN avec option 0832. Best, Nicolas
INC280514 | Workgroup | Incident | TR: price and delivery - MFT5-P03.X10X.N2/9704 (for Tim's enquiry for CDL) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 12:03 | 2023-03-29 14:13 | 0 | 0 |  | Nathalie, en me basant sur la photo et leur définition, c'est bien possible et ce sera un: pas de code sap existant, et cadran a valider via Fantinelli ainsi que le prix. Order code: MFT5-P03.X10Q.N2/9704/9709 MFT old Code / SAP code: Nominal size: 100 mm 5 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection with mounting bracket for 2"pipes P Type of case: Case without partition of safety 0 Oil filling: BH3, silicone oil 3 Pressure unit: Special pressure units X Pressure range: Pressure range for mbar 0 ... 250 mbar 10 Static pressure: 250 bar Q Process connection: AISI 316L (1.4404), stainless steel 1/2" NPT female N2 Echelle quadratique (photo) 0=0 et 200mbar=11800kg/h 9704/9709
INC280513 | Workgroup | Incident | ETW 099 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 11:27 | 2023-03-29 11:46 | 0 | 0 |  | Bonjour Cyrille, Je ne trouve pas beaucoup d'infos mais les codes des ETW sont lies à 2260 ou à 2110 ce qui veut dire que c'est fait ou acheté en Allemagne, il faudrait demander a un contact la bas ? Best, Nicolas
INC280511 | Workgroup | Incident | MEP5-460.B25/0765_0771 (0...20 bar) / Chemicoat | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 10:45 | 2023-03-29 10:57 | 0 | 0 |  | Hi Bee Tin, yes it's OK, it'll be the new design. Best, Nicolas
INC280510 | Workgroup | Incident | MEX3 0...250Bar, 2 bar/Division /Wai Seng | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 10:39 | 2023-03-30 11:10 | 0 | 0 |  | Dear Bee Tin, here is the dial in Bar Bourdon (small divisions) every 2bar. [Image] Best, Nicolas
INC280496 | Workgroup | Incident | Please Please create the following item MEx5-D31.B77.DANC01.2L2A.01/Q003_9900_9704 | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 07:48 | 2023-03-29 09:06 | 0 | 0 |  | Hi Jens, for code creation, you must ask the func material data email, here you reached the technical support. Best, Nicolas
INC280493 | Workgroup | Incident | MQ7-460.B24J/0765 With 3-valve Manifold / Chemicoat | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-29 05:21 | 2023-03-29 09:13 | 0 | 0 |  | Dear Bee Tin, mounting of an MQ7 (or MZ or MFT) is usually done on an AMFD05N (5ways), I never saw a mounting on 2x3ways manifolds. here is an example with the accessories to add between the gauge and the manifold: [Image] Best, Nicolas
INC280486 | Workgroup | Incident | TBW | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-28 12:00 | 2023-03-29 17:24 | 0 | 0 |  | Hi Jurgen, it's the same sensitive element as the TBX which is existing, so yes, you can go ahead with this code: TBW100-212.111.14T Best, Nicolas
INC280476 | Workgroup | Incident | Y91-3.3.3.B22.R/0592+ASIPD2L3+DT8-2L20.0.230.09 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-28 09:42 | 2023-03-28 14:42 | 1 | 0 |  | Hi Xu Qiang, why not a DT2 that will be a lot cheaper ? at 10bar, you don't need a DT8: Y91-3.3.3.B22.R/0592+ASIPD2L3+DT2-2L20.0.230.09 Best, Nicolas
INC280475 | Workgroup | Incident | TR: RP2EL | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-28 09:33 | 2023-03-28 14:39 | 1 | 0 |  | Nathalie, NH3 is not particularly aggressive so mainly I would say yes, but the customer is the only person responsible for the choice of material as compatibility depends on many other factors (temperature, pressure, geometry of the parts, concentration etc...) Best, Nicolas
INC280474 | Workgroup | Incident | TR: BAPL ¦ PFSX ¦ odp ¦ zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-28 09:30 | 2023-03-28 14:37 | 1 | 0 |  | You could offer a TED in replacement but it's not SIL2 (main reason for the PFSX existence). Old datasheet enclosed for reference. Best, Nicolas
INC280459 | Workgroup | Incident | Query for model MEP3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-28 06:51 | 2023-03-28 14:35 | 1 | 0 |  | Hi Anil, you can go ahead with up to 100 000kPa on the MEP3 (D41). considering the Q017 it's a standard certificate and it's available for stainless gauges as our connections and tubes are in 316L. please when you'll ask for the lead time, specify that you need the Q017. Best, Nicolas
INC280451 | Workgroup | Incident | WG: PROFESS - Pressure gauge with chemical separator | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-27 15:57 | 2023-03-28 14:29 | 1 | 0 |  | Hey Carlo, considering the pressures, you'll be less expensive and it'll be easier to define a MEX on DT. I suggest a lining of the diaphragm at least and the connection can be M20x1.5 (if there is enough qty so it's a viable financial option) but as they have an adapter we could also offer G1/2 if it's what the adapter is or an ARIN to replace the one they have. Best, Nicolas
INC280429 | Workgroup | Incident | TED-6.3.3.B24.R+DADL05.2.L2A.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-27 06:27 | 2023-03-27 09:19 | 0 | 0 |  | Hi Bee Tin, yes the construction is ok and as the code of the TED is existing there is no need to ask Franck. (lease specify it's OK with NDO in your request, I validated with her). Best, Nicolas
INC280428 | Workgroup | Incident | Y91-3.3.3.B24.R.DADL05.2.L2A.01 /Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-27 06:24 | 2023-03-27 09:20 | 0 | 0 |  | Hi Bee Tin, yes, the configuration is OK, but for the lead time, please ask commercial. Best, Nicolas
INC280426 | Workgroup | Incident | TR: Bourdon quote | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-24 15:57 | 2023-03-27 09:22 | 0 | 0 |  | Nathalie, je pense que tu cherche ce code SAP 11198166 Best, Nicolas
INC280424 | Workgroup | Incident | Y91 spécifique | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-24 15:03 | 2023-03-27 12:05 | 0 | 0 |  | Bonjour Cyrille, voici la réponse que j'ai faite aux US, il y a peu de chance que vous l'ayez, la demande sera surement répondue directement des US ou du Canada: vous pouvez toujours avoir ce produit, mais sans l'option 0839 qui n'est plus disponnible et nous conseillons une connection standard (aussi DIN43650 mais avec PE standard) qui sera bcp plus avantageuse au niveau prix et tres probablement compatible. Y913 59 H37R/2158/2151/2154. Best, Nicolas
INC280419 | Workgroup | Incident | CAPILLAIRES D912 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-24 12:21 | 2023-03-24 13:56 | 0 | 0 |  | Bonjour Veronique, pas vraiment pour l'instant mais c'est quelque chose sur lequel on travail. Best, Nicolas
INC280415 | Workgroup | Incident | Demande de création / OP20509971 / PMC / MEX3-D20.F74/0679.XDSE:52642-02/LRS1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-24 11:30 | 2023-03-24 12:26 | 0 | 0 |  | Gregory, tu as du te tromper, il faut l'envoyer a la codif. Best, Nicolas
INC280411 | Workgroup | Incident | Worley - Shell C&T | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-24 10:24 | 2023-03-24 12:25 | 0 | 0 |  | - Can you specify the material used for the blow-out protection on the top of the MEX5? [Image] - Can you specify the grade of stainless steel used for the movement in a MEX5? (It is not specified on the datasheet) St. St. 1.4301 - When the customer asks for an "Inspection TEST Plan", what do they usually mean by that? an ITP is usually part of a complete Manufacturing report that states the inspection details (does an external inspector comes, how much of the order is tested, is it visual inspection or are tests possibles etc...) - What information is usually stamped into the pressure gauge socket? A 2 digits code that states the time of manufacturing. - The customer wants both the material used and the pressure range stamped into it it's possible but complicated to define (to see with R&D and codification) but the options for technical marking on the case or dial are available. Best, Nicolas
INC280400 | Workgroup | Incident | Plan EF6 pour Sncf | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 16:18 | 2023-03-23 16:37 | 0 | 0 |  | Sure, you can find it in the ticket: Code Sap Plan 11151997 SP3415014 11150576 SP3415010 Ind A 11151998 SP3415015 Ind B 11151991 SP3415011 Ind A 11151995 SP3415013 Ind A 11151993 SP3415012 Ind A Best, Nicolas
INC280394 | Workgroup | Incident | MEM5-G61.F15 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 12:42 | 2023-03-23 14:07 | 0 | 0 |  | Hi Daniel, here is the code, I created it (you might need to ask for it to be open for your SU to material data). 11263322 Best, Nicolas
INC280393 | Workgroup | Incident | Transmetteur IP3972 pour Novintec | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 12:18 | 2023-03-23 14:23 | 0 | 0 |  | Cyrille, je ne vais pas etre le mieux placé pour aider la dessus, je te conseille de passer directement par une procedure de non qualité, je sais que la qualité a géré beaucoup de choses sur ces produits et ils seront mieux placés une fois l'alerte recué pour orienter la suite a donner. Best, Nicolas
INC280387 | Workgroup | Incident | MMD dial | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 11:00 | 2023-03-23 16:48 | 0 | 0 |  | Hi Kerstin, here is the dial of the standard MMD3 10bar EN: [Image] Best, Nicolas
INC280384 | Workgroup | Incident | TR: inquiry for Y923 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 10:51 | 2023-03-23 14:19 | 0 | 0 |  | Nathalie, tu peux y aller, on en a fait des speciaux pour l'H2 il y a pas longtemps (avec des options en plus mais la base est la meme). Best, Nicolas
INC280354 | Workgroup | Incident | MEX3 With Special Scale | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-23 07:54 | 2023-03-23 16:46 | 1 | 0 |  | Hello Bee Tin, please find hereafter the standard dials, otherwise special ones needs to be created: [Image] Best, Nicolas
INC280351 | Workgroup | Incident | RE: EF6 Faiveley | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 16:45 | 2023-03-22 17:43 | 0 | 0 |  | Cyrille, ca va pas etre tellement moi du coup, je conseille de demander le délai Céline martin et preciser quel code SAP de la 3409001 ? [Image] Après pour un produit qui necessiterait une nouvelle carte potentiellement, il faut demander à Franck Ferrarese et au PM Ferro (peut etre par une SPR) Best, Nicolas
INC280349 | Workgroup | Incident | Demande de fiche technique MEP8 | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 15:51 | 2023-03-22 17:47 | 0 | 0 |  | Hi Greg, I just had Jorg and it's a no, he doesn't have it released yet. Sorry for this. Best, Nicolas
INC280348 | Workgroup | Incident | 11159224 step file | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 15:33 | 2023-03-22 17:59 | 0 | 0 |  | Hi Kerstin, here we don't have it but I still asked Jorg in case he's got something. I close it because I think there is few chances, but I'll eep you update if we get something. Best, Nicolas
INC280346 | Workgroup | Incident | BAIT ¦ Terranova instruments ¦ Richiesta offerta 4229 PG MEP5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 14:51 | 2023-03-24 11:18 | 0 | 0 |  | Hi Erald, the MEP code is from Baumer, but the seal SEA06 is not from us, it's from Terranova. we don't have inconel diaphragm on our D04x, if you need us to o the seal it needs to go through an SPR. for the certificates, you have the Q00x, it should not be an issue. Best, Nicolas
INC280340 | Workgroup | Incident | 30153082 BRD912 SIEMENS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 12:09 | 2023-03-24 15:12 | 0 | 0 |  | Salut Véroniques, les brides BRD912 ne seront pas compatibles on a un systeme prorietaire entre la bride et la BRD uniquement pour les D912 il faudra prendre les BRD ci-dessous, ci-joint le plan à faire valider au client pour le montage. 53288-03-AS-03-0V 1"150RF 53288-03-AS-12-0V 2"150RF 53288-03-AS-06-0V 1"300RF 53288-03-AS-15-0V 2"300RF Best, Nicolas
INC280338 | Workgroup | Incident | conditionneur Sedeme TM26 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 11:36 | 2023-03-22 11:51 | 0 | 0 |  | Salut Cyrille, désolé, je suis remonté jusque dans les années 80 dans les catalogues mais ca devait etre sedeme avant bourdon sedeme, je n'ai rien trouvé. Best, Nicolas
INC280334 | Workgroup | Incident | OPTIONS Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 09:42 | 2023-03-22 11:40 | 0 | 0 |  | Veronique, oui c'est possible, et je te joints la lettre explicative pour la difference entre les 2 options. Best, Nicolas
INC280327 | Workgroup | Incident | CONFIRMATION DEFINITION D912 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-22 08:06 | 2023-03-22 12:19 | 0 | 0 |  | Veronique, le LRS8 est bien obligatoire pour les mesures en Vide et Abs. la limite en température est de 300C On est donc un peu bloqué, si le client a un montage en T ou une extension pour reduire la température a 300 ce serait OK, sinon il prendra un risque de dilatation du liquide et de creation de bulles qui peuvent fausser les mesures, voir endommager la membrane. Best, Nicolas
INC280325 | Workgroup | Incident | FW: Baumer Process. | Archive > BHFR Technical Support | Middle |  |  | Closed | Moran, Andrew [am] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 16:54 | 2023-03-22 11:37 | 0 | 0 |  | Hi Andrew, the mountings you choose are fine, the MEP5+D425 min range for fitting is 1bar and MEP5 + D820DN25 is 10bar. adding a capillary of 10m will work also, it doesn't change the mountings properties, just be careful tu use 1/2NPT everywhere if you start with the 1/2NPT gauge. Best, Nicolas
INC280323 | Workgroup | Incident | MEX5 avec capillaire et séparateur | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 15:51 | 2023-03-22 14:14 | 1 | 0 |  | Cyrille, Il manque juste le LRS et les options seront peut etre appliquées sur le mano mais pas de probleme pour moi sinon. pas de problème, le MEX5 + D425 se monte à partir de 4Kg/cm et les 2m de capillaires ne changerons rien. par contre, je ne peux pas donner de prix, il faut voir avec l'administration des ventes. Best, Nicolas
INC280319 | Workgroup | Incident | TR: inquiry MFT7-P10.X11Q.82100 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 14:54 | 2023-03-22 11:56 | 1 | 0 |  | Nathalie, la codif dans le config sera telle que ci-dessous donc avec le digit 1 pour ATEX, je vais faire remonter pour correction du P2F. [Image] il faudra verifier avec fantinelli si le cadran est a creer et si il y a des frais. Best, Nicolas
INC280317 | Workgroup | Incident | 30152869 EDF | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 11:48 | 2023-03-21 14:30 | 0 | 0 |  | Bonjour Véronique, non, ce n'est pas tout jeune, probablement des années 90, je ne retrouverais pas la référence avec le n de série. On peut proposer en remplacement un TSF à collerette avant, mais si c'est pour EDF, il faudra redéfinir le produit et probablement faire valider un plan. Désolé de ne pas pouvoir aider plus, mais c'est trop ancien. par contre, EDF a peut etre un plan de l'époque qui nous aiderait à redéfinir le produit. Best, Nicolas
INC280314 | Workgroup | Incident | TED 5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 11:36 | 2023-03-21 14:35 | 0 | 0 |  | Cyrille, pour l'option c'est bon, par contre les codes de TED sont en statut PN, donc si le code est a creer il faudra demander l'autorisation au PM. ca ne devrait pas poser de probleme normalement. Best, Nicolas.
INC280299 | Workgroup | Incident | MCX5-460.N06/0693_0751_9941+AORP0502/SETP+ARIN2606 /Sakura | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-21 08:51 | 2023-03-21 10:18 | 0 | 0 |  | Hi Bee Tin, I suggest to change the AORP for a 501 that will allow you to cut the pressure around 100/150mbar (thats the min). it means that from 40mbar to 150mbar you rely on the safety valve internal to the MCX5 and reaching 150mbar, the AORP closes and you're safe up to 400bar. It is not ideal, because the internal MCX5 overpressure safety is a releif valve so can only handle short terms overpressure, it means the after reaching 40mbar and before reaching the 150mbar setp of the AORP, you potentially have gas release in the case (which could be fine if it's air or any safe gas). Best, Nicolas
INC280291 | Workgroup | Incident | STEP file for contact gauges required | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-20 15:36 | 2023-03-31 15:57 | 0 | 1 |  | Nicolas will check this out with Jorg as the drawing needs to be done.
INC280282 | Workgroup | Incident | Kindly confirm if the model Barflex EA4 is still available. | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-20 10:30 | 2023-03-20 15:04 | 0 | 0 |  | Hi Anil, we can do it (similar to 11192329) but you need the approval of Franck or Mario to create the code (I also will forward you their email from this morning. Best, Nicolas
INC280270 | Workgroup | Incident | Liquid Filling For O2 Gas Applicaiton / Chemioat | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-20 08:27 | 2023-03-20 15:07 | 1 | 0 |  | Hi Bee Tin, yes your selection is fine for the compatibility. Only O2 can react directly with organic material. Best, Nicolas
INC280269 | Workgroup | Incident | FW: quotation requests of MDJ7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-20 07:54 | 2023-03-20 15:12 | 1 | 0 |  | Hi Bee Tin, you can offer the MFT below as replacement, there will probably be a fee for the dial creation: Order code: MFT7-B00.X11Q.N2 MFT old Code / SAP code: Nominal size: 150 mm 7 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Back connection, front flange, 3 mounting holes B Type of case: Case without partition of safety 0 Oil Filling: no oil filling 0 Pressure unit: Special pressure units X Pressure range: Pressure range for mBar 0 ... 400 mbar 11 Static pressure: 250 Bar Q Process connection: AISI 316L (1.4404), stainless steel 1/2" NPT female N2 Best, Nicolas
INC280268 | Workgroup | Incident | FW: RFQ- BAUMER/ Temperature Gauge / Unibel | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-20 07:36 | 2023-03-21 12:13 | 1 | 0 |  | Dear Bee Tin, I can't realy see the serial number but I read either 6191 or 5191, and both match the same part numbers 5191 TBH080-212.152.52T 10321301 6191 TBH080-212.152.52T 10321301 Best, Nicolas
INC280267 | Workgroup | Incident | TR: OP Renault España SA Montaje Vall 15032023 20508309 ¦ RP2-YM.363+DANC01.2L2A.07/8991_0831_0794 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 17:30 | 2023-03-24 15:18 | 1 | 1 |  | Celine, on a du regarder de plus près mais c'est ok pour Sylvain et Christophe, attention, le montage sera limité par la pression de montage max du clamp (25 à 60bar selon le collier utilisé par le client). RP2-YM.355 + DANC01.2L2A.07/8991_0831_0794 Best, Nicolas
INC280258 | Workgroup | Incident | TR: Prix-IC ? - ExxonMobil (55203668) - RFQ 20230317 - MEX8-DX1.B29/9690_C026 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 14:36 | 2023-03-17 15:33 | 0 | 0 |  | Oui, c'est le raccord 138 pour exxon mobile voici un exemple en RL et le plan du raccord: 10286264 MEX5-DX1.B29/0751 [Image] Best, Nicolas
INC280257 | Workgroup | Incident | FW: Re: BARFLEX(R) - Manomètre électronique portatif - Offre de prix Baumer - Air Liquide Medical 0020488747 - (RFQ BARFLEX).pdf | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 14:15 | 2023-03-21 15:34 | 0 | 0 |  | Hi Jurgen, no good news for this, the only way to connect to a computer is the Irda and this is not available anymore. I think you were not aware of this, but a couple years ago, the Irda was already outdated and the software was only running like on windows XP and now the Irda is not available anymore. I'm sorry, if you were not aware I guess it puts you in an akward position with the customer. Best, Nicolas
INC280249 | Workgroup | Incident | TR: OP Renault España SA Montaje Vall 15032023 20508309 ¦ RP2-YM.363+DANC01.2L2A.07/8991_0831_0794 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 12:18 | 2023-03-17 15:55 | 0 | 0 |  | Celine, le montage demandé ci-dessous ne se monte pas pour 2 raisons, la membrane est trop petite et hautes surpressions pour lesquelles sont faites les modeles 61-66 ne seront de toute facon pas admissible par le clamp. RP2-YM.363 + DANC01.2L2A.07 on peut par contre passer sur un RP2-YM.355 (25bar 100bar surpression au lieu de 200) qui pourrait etre monté avec un clamp DN51: RP2-YM.355 + DANC02.2L2A.07/8991_0831_0794 les options sont OK Best, Nicolas
INC280248 | Workgroup | Incident | TR: Baumer Process. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 12:06 | 2023-03-17 14:41 | 0 | 0 |  | Nathalie, le D820 DN25 avec un MEP5 se monte a partir de 10bar donc seul le B22 passe. pour un remplacement, il faudra regarder en D435/25. Best, Nicolas
INC280239 | Workgroup | Incident | leaking TED-5 - Air Liquide | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 10:57 | 2023-03-17 11:03 | 0 | 0 |  | Hi Rudi, didn't heard of a reccuring issue like this, you'll have to make a return (probably with replacement units. we'll need to know the exact gas and what the TED is displaying, because a leak there means a leak on te cell so it's probably damage. I would say the first thing that comes to mind here is overpressure that damaged the cell and caused the leaks. I would advise the customer to add a gauge saver to his process below the TED. Best, Nicolas
INC280235 | Workgroup | Incident | 3D Step Data | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 10:27 | 2023-03-17 10:56 | 0 | 0 |  | Hi Felix, I'm sorry, I don't have the 3D drawing available on this one. you can request it via sales but I don't know if there is a cost. Best, Nicolas
INC280232 | Workgroup | Incident | M31-2D6.B18L/0764_0835 / PT Air Liquide/ PGS | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 10:12 | 2023-03-17 11:26 | 0 | 0 |  | Hello Bee Tin, the 0835 is the laboratory cleanliness and is available on M31 (marking on dial "Laboratory cleanliness") the 0765 is the cleaning for O2 applications (marking on the dial "OXYGEN USE NO OIL" but you can't have both on the same device, if the usage is really on O2, I suggest the 0765, if the cleaning must be deep but not focus on H2 (for different gases including O2 in laboratory environment), I suggest the 0835. Best, Nicolas
INC280229 | Workgroup | Incident | 3D STEP Data | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 10:09 | 2023-03-17 16:09 | 0 | 0 |  | Hi Felix, again same answer, I'm sorry but on the old german gauges and the TSS from stiko, we're quite poor in 3D drawings. Best, Nicolas
INC280220 | Workgroup | Incident | Customer label for pressure swicht | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 08:45 | 2023-03-17 10:43 | 0 | 0 |  | Answered via teams, pictures of the products and markings sent, needs to go through an SPR for price and specific executions. if only for the marking on the RP2N, we can directly make an offer I thiink. Best, Nicolas
INC280216 | Workgroup | Incident | 3D TED-6.18.3.B26.R | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-17 08:12 | 2023-03-17 09:36 | 0 | 0 |  | Sure, here it is in the ticket: Best, Nicolas
INC280202 | Workgroup | Incident | TR: RE: | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 14:54 | 2023-03-16 15:07 | 0 | 0 |  | Yannick, je ne pourrais pas plus retrouver de reference, mais on a bien fait des séparateurs S6, le design a changé, je proposerais un D4xx. la référence ci-dessous devrait leurs convenir: D435-2L2.7CD.D20 Best, Nicolas
INC280199 | Workgroup | Incident | TEDx & BARFLEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 14:09 | 2023-03-16 15:11 | 0 | 0 |  | Hi Rudi, TED and barflex are not closed, there is only a blocking to create new product to controle what we create. everything they are used to order is fine, nothing changes for them. Best, Nicolas
INC280197 | Workgroup | Incident | TR: IC Pricing | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 13:54 | 2023-03-16 14:49 | 0 | 0 |  | Nathalie, ci-dessous mes commentaires: 2pcs RP2-NL.346/SETP mounted to D820-10L.2216.G20 Pas possible en 1"1/2, le mini de montage sera du 2" 2pcs RP2-NL.346/SETP mounted to D820-10L.2217.G20 OK (Surpression max: 50bar et mini de reglage: 1bar) 17pcs RP2-NL.346/SETP mounted to DANC01-2L2B Pas possible en DN38, le mini de montage sera du DN50/51 5pcs RP2-NL.346/SETP mounted to DANC01-2L6B Pas possible en DN38, le mini de montage sera du DN50/51 4pcs RP2-NL.346/SETP mounted to DANC02-2L2B OK (Surpression max: 40bar et mini de reglage: 1bar) Best, Nicolas
INC280190 | Workgroup | Incident | EF6 Faiveley | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 12:03 | 2023-03-16 14:58 | 0 | 0 |  | Cyrille, ci-dessous une spec d'EF6 pour faiveley de 2023 sur un EF6 equivalent. on est en 10...32VDC et si ils recommandent des produits, je suggere de refaire une specs avec des références recentes: [Image] Best, Nicolas
INC280186 | Workgroup | Incident | TR: BAIT - Savio - CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 11:18 | 2023-03-16 14:24 | 0 | 0 |  | Hi Nathalie, we can't go that low in pressure with the CTX, for this application, I suggest to llok at the PBMN, PP series with low pressure and possible programming. Best, Nicolas
INC280183 | Workgroup | Incident | PROBLEME JOINTS MANOS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 10:54 | 2023-03-16 14:22 | 0 | 0 |  | Bonjour Véronique, la température et le Bore ne pose pas de problème. nous utilisons en standard le joint Alu ci-dessous avec un couple de serrage de 80Nm 10305228 JOINT 1/2G ALU SP2490 1050 A [Image] [Image] Best, Nicolas
INC280156 | Workgroup | Incident | PBSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-16 09:00 | 2023-03-16 14:00 | 1 | 0 |  | Bonjour Cyrille, ci-joint le MTTF du PBSN Best, Nicolas
INC280144 | Workgroup | Incident | Demande d'identification produit / LEAD 30152191 | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-15 16:45 | 2023-03-16 15:29 | 0 | 0 |  | Salut Greg, c'est un mano de chez nous pour de la Tres haute pureté qui a plus de 20ans. on ne peut plus proposer de manos UHP, on ne le fait plus. Best, Nicolas
INC280124 | Workgroup | Incident | TED-6 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-15 12:12 | 2023-03-15 14:03 | 0 | 0 |  | Bonjour Cyrille, Désolé, mais les n de commande en Y sont trop vieux, on a plus acces au système. il va falloir redefinir le raccord et le joint avec le client. Best, Nicolas
INC280082 | Workgroup | Incident | WG: acetylene plant | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-14 15:00 | 2023-03-14 15:16 | 0 | 0 |  | Hello Felix, there is no material compatibility issue apparently (see table below): [Image][Image] But as it's an extremely flammable gas so I guess there is ATEX zones and types of protection to respect. Only the customer can be aware of the zones and protection needed. Best, Nicolas
INC280069 | Workgroup | Incident | PRESSOSTAT DIFFERENTIEL | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-14 12:06 | 2023-03-14 15:27 | 0 | 0 |  | Salut Véronique, avec un RDN 234 (0.15...4bar), ils pourraient avoir un point bas a 0.6 (diff, à la baisse) et 2.6 (diff, a la hausse). ou avec un RDN 235 (0.8...4bar), 1.2 a la baisse et 3.6 a la hausse (toujours en pression diff). Best, Nicolas
INC280053 | Workgroup | Incident | Article Request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-14 09:24 | 2023-03-14 10:55 | 0 | 0 |  | Hi Sascha, please find enclosed the excel with the definitions, I was not able to do all because of the lack of information, for the rest, you need to define with the customer. Best, Nicolas
INC280050 | User | Incident | SV: Air Liquide | Archive > BHFR Technical Support | Low |  |  | Closed | Hedemark, Thomas Jydby [tjh] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Mosberger, Marco [mmos] | 2023-03-14 08:51 | 2023-03-14 11:29 | 0 | 0 |  | Hi Thomas, please make sure to send those questions to technical support, Marco had to redirect it to me. the cleaning options are as follow: 0765 O2 cleaning (conform to EN) (Procedure ST004) 0764 O2 cleaning (not conform EN, the procedure is the same but no EN marking for example on a MEX5 as the EN calls for Safety gauges) (Procedure ST004) 0829 Specific cleaning (clean and dry) for use on clean gases to avoid pollution with dust/grease particles. (Procedure ST029) we do not disclose our internal procedures except with special authorizations. Best, Nicolas
INC280032 | Workgroup | Incident | BAIT ¦ HOLTEK RICHIESTA DI OFFERTA: 2302572 | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-13 18:30 | 2023-03-15 17:29 | 0 | 0 |  | Hi Erald, I was not able to find out the original part number (too old or there is something weird in the codification) but I can help you define the seal fo sure, here is the suggested code with the capillary length to check (4.5m here): D912LD0FG82L27G0M110 it must be LRS8 for vaccum use and the max temp is 300C Also I can't see if they ordered options so here are the most common ones to offer. [Image] Best, Nicolas
INC280014 | Workgroup | Incident | BAIT ¦ homogenizer process connection | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-13 12:21 | 2023-03-13 14:33 | 0 | 0 |  | Hi Erald, please find enclosed the datasheet of the diaphragm seal you're looking for and it potentially can fit with a Y91. it would be the first code created so we need to go through the SPR process and probably make a test sample. please create and SPR and send it to your sales contact, and we will be able to take a decision from there. Best, Nicolas
INC279970 | Workgroup | Incident | 3D Step Data | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-10 14:48 | 2023-03-10 15:01 | 0 | 0 |  | Hi Felix, I only have it for D2, not D3, the rest is what you need. Best, Nicolas
INC279932 | Workgroup | Incident | RE: BAIT ¦ Italianpack: Alternativa a trasd pressione IFM ¦ 500pcs | Archive > BHFR Technical Support | Middle |  |  | Closed | Ferrarese, Franck [ffr] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 17:54 | 2023-03-10 14:15 | 1 | 0 |  | Solved via FFER. Best, Nicolas
INC279930 | Workgroup | Incident | RE: BAIT ¦ Italianpack: Alternativa a trasd pressione IFM ¦ 500pcs | Archive > BHFR Technical Support | Middle |  |  | Closed | Ferrarese, Franck [ffr] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 17:48 | 2023-03-10 14:17 | 1 | 0 |  | Solved via FFR, Best, Nicolas
INC279931 | Workgroup | Incident | R: BAIT ¦ Italianpack: Alternativa a trasd pressione IFM ¦ 500pcs | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 17:48 | 2023-03-10 14:16 | 1 | 0 |  | Solved via FFER. Best, Nicolas
INC279929 | Workgroup | Incident | BAIT ¦ Italianpack: Alternativa a trasd pressione IFM ¦ 500pcs | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 17:45 | 2023-03-10 14:17 | 1 | 0 |  | Solved via FFR Best, Nicolas
INC279925 | Workgroup | Incident | TR: 1935/2004 Declaration of compliance. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 16:15 | 2023-03-10 10:52 | 0 | 0 |  | Hi Solveig, I added to the ticket the following certificaes: 1935 for filling fluids 1935 for DANC (fitted with transmitters, sorry I don't have one with just the DAxx) 3A certificate for mounting MEX + DANC (used for the same type of applications) Best, Nicolas
INC279923 | Workgroup | Incident | 3D Step Data | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 15:51 | 2023-03-10 10:53 | 0 | 0 |  | Hello Felix, please find the STP file enclosed: Best, Nicolas
INC279901 | Workgroup | Incident | projet novasep | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 11:21 | 2023-03-09 12:08 | 0 | 0 |  | Bonjour Yannick, pour moi c'est bloqué juste parceque le code est mal construit, si tu demande la recreation, ca sera bon. le D820 passe en 821, et les options habituelles de montage sur séparateur seront ajoutées. Best, Nicolas
INC279895 | Workgroup | Incident | MR5 with option 4716 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 10:30 | 2023-03-17 16:46 | 0 | 1 |  | Hi Felix, the code is old and we will transfer it to newest devices. Here is what we can offer, we just spotted a weird thing, it seems like it's a flush diaphragm seal G3/4, but we have an adaptor in the code from G3/4 to G1/2 and it makes no sense then to have a flush mounting, better choose a DT or D04 that will be cheaper. MR5-4D30.B38/4716 + D055-2L15-01 please confirm the mounting part with the customer and check if a DT or D04 would not be a better option. Best, Nicolas
INC279879 | Workgroup | Incident | MEX5 with dial scale -760 to +760 mmHg | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-09 06:45 | 2023-03-09 12:02 | 0 | 0 |  | Hi Bee Tin, yes it exists, here is an existing code: 11174894 MEX5-D60.X73/9709_D665 Best, Nicolas
INC279870 | Workgroup | Incident | 3D Step Data | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-08 16:09 | 2023-03-08 16:34 | 0 | 0 |  | Hello Felix, the step drawing is attached in the ticket, Best, Nicolas
INC279866 | Workgroup | Incident | RV: RFQ BARFLEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-08 15:33 | 2023-03-08 15:41 | 0 | 0 |  | Hi Angel, it'll be a EA5-AF0.H30R (11100867) because there is no gasket in the EA5A version. this product is phased out so it'll go through authorization of the PM if you ask for ordering it. for the 1/4NPT connector, you need an ARINX05 (M12*1.5 female / 1/4NPT male) but for low qty it'll be expensive. Best, Nicolas
INC279827 | Workgroup | Incident | Demande de définition séparateur / BOSTIK | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-08 09:00 | 2023-03-08 14:28 | 1 | 0 |  | Greg, c'est très dur de définir un separateur avec les dimensions sans connaitre le type. Je peux te dire que c'est très petit (un D821 DN1/2 fait 89mm de diamètre). De plus avec ce type de diamètre, selon l'EN et l'ASME, ce sera plutot du 4 vis. C'est surement soit du spécifique soit un DN très petit et peu utilisé. il me faudra d'autres informations (une photo dessus/dessous). je ferme le ticket pour l'instant, tu me recontacte dés que tu a des infos. Best, Nicolas
INC279825 | Workgroup | Incident | validation SANOFI MARCY ETOILE CONCURRENCE Guterriez | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-08 08:45 | 2023-03-08 14:33 | 1 | 0 |  | Bonjour Yannick, ok pour l'option en DN51 mais pas en DN38. par contre pour le switch a rearmement manuel, il faut le valider avec philippe, si ce n'est pas fait dis le moi et je lui demanderais (normalement pas de problème). Best, Nicolas
INC279785 | Workgroup | Incident | Atlas Copco (21109244) - PO 760317 - E91 -> Y91-3.2.3.B29.R/2164 incl. 9001.0100 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-07 11:21 | 2023-03-08 15:14 | 1 | 0 |  | Hi Rudi, we don't use the option 9001/xxxx anymore, you can just order the cable separately, here are the code: 11082947 Y91-3.2.3.B29.R/2164 + 10303915 CORDON M12 10MBLINDE ATEX RKTS529810M Best, Nicolas
INC279779 | Workgroup | Incident | capteur ferroviaire E304 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-07 10:12 | 2023-03-31 16:00 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279753 | Workgroup | Incident | TR: QU3161522 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-06 17:39 | 2023-03-08 15:34 | 1 | 0 |  | Nathalie, on peut remplacer par ce code, la différence est que nous ne fournissons plus le joint. TED-6.3.3.B26.R.DADL04-2.L2A.01/0592 Best, Nicolas
INC279747 | Workgroup | Incident | WG: MSDS for MIX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-06 16:21 | 2023-03-31 16:00 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279740 | Workgroup | Incident | TR: RFQ // MEX3-D50.H30/0059-HV | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-06 15:39 | 2023-03-31 15:59 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279735 | Workgroup | Incident | ETTN | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-06 13:36 | 2023-03-31 15:59 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279732 | Workgroup | Incident | D821 + D415 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-06 13:03 | 2023-03-31 15:59 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279698 | Workgroup | Incident | MEP5-D31.B26/0693_0078_0678_9941, 2 nos /Neste | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Lecomte, David [leda] | 2023-03-06 05:15 | 2023-03-31 15:59 | 1 | 1 |  | Outdated, please contact us if any further action needs to be done. Best, Nicolas
INC279678 | Workgroup | Incident | Fichier 3D COMEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Dupin, Yannick [dya] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-03 15:45 | 2023-03-03 16:05 | 0 | 0 |  | Bonjour Yannick, désolé on a pas de 3D dispo pour le MIX7 type B. Best, Nicolas
INC279672 | Workgroup | Incident | Création codes articles plaquette de soudure 1mm et 1.5mm BOHM | Archive > BHFR Technical Support | Middle |  |  | Closed | Lecomte, David [leda] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-03 14:33 | 2023-03-03 16:07 | 0 | 0 |  | David, tu m'as envoyé les demandes de codes a la place de Nadia. je te laisse lui renvoyer, je ne peux pas tranferer de ticket a son ticketing. Best, Nicolas
INC279669 | Workgroup | Incident | TR: RFQ - BAUMER / PLNG-AMNT-23037 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-03 14:18 | 2023-03-03 16:42 | 0 | 0 |  | Nathalie, je t'ai transféré l'offre de Fantinelli pour le remplacement Best, Nicolas
INC279619 | Workgroup | Incident | MEX-5 + D821 | Archive > BHFR Technical Support | Middle |  |  | Closed | Horber, Felix [hfe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-03 10:03 | 2023-03-03 10:45 | 0 | 0 |  | Hi Felix, the PN40 limits the mounting to 40bar, I suggest a D415 with a PN100: D415-2L20.4JC.B2000 anyway, the customer cannot have a PN40 installation with 50bar in it. Best, Nicolas
INC279589 | Workgroup | Incident | RE: Demande de prix / H2O Prod | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-02 16:24 | 2023-03-02 16:27 | 0 | 0 |  | Doublon
INC279588 | Workgroup | Incident | Demande de prix / H2O Prod | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-02 16:21 | 2023-03-02 16:25 | 0 | 0 |  | Ma faute, j'oublie tout le temps que c'est traité par l'allemagne. Best, Nicolas
INC279565 | Workgroup | Incident | TR: DRO100 for R134A 12 NOS /GUE | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-02 10:24 | 2023-03-02 16:26 | 1 | 0 |  | Nathalie, oui mais je n'aurais pas assez d'info avec les photos, il faut redefinir le produit si possible en MEX5. les cadrans en R134A existent surement. Best, Nicolas
INC279559 | Workgroup | Incident | Demande de définition séparateur / PMC | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-02 09:42 | 2023-03-02 11:05 | 0 | 0 |  | Salut Greg, voici ce que je vais avoir de plus proche, et un revetement est possible (a voir lequel convient le mieux). ci-joint le plan, ce sera surement codifié en tant que D006. mais on ne l'a qu'en 10 trous, ca ne va surement pas marcher (si tout le reste est OK pour le client je peux peut etre demander a la faire faire en 8 trous). Best, Nicolas
INC279537 | Workgroup | Incident | Temperature Switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-02 05:03 | 2023-03-02 09:55 | 0 | 0 |  | Hi Anil, with the information we have I would offer them the 2 options below and send the datasheets for revision: RT-NAB.301.E00EJ/SETP (direct mounting but G3/8 fixed connections) RT-NAB.401.232E3/SETP (available with other length) the diameter of the stems is 14mm. Best, Nicolas
INC279536 | Workgroup | Incident | doigt de gant en équivalence | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 17:36 | 2023-03-01 18:01 | 0 | 0 |  | Cyrille, ca va etre du specifique je pense, le mieux c'est de faire la demande de prix avec le plan (ci-joint). on fera directement une demande au fournisseur avec ce plan. apres pour la vente, a voir avec la codif si on prefere codifier un code standard avec des X ou un code specifique avec un n de plan. Best, Nicolas
INC279535 | Workgroup | Incident | TR: Connector for pressure and differential pressure transmitters ¦ złącza EN 175301-803 A | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 17:21 | 2023-03-02 14:55 | 0 | 0 |  | Nathalie, nous avons ces connecteurs DIN43650 (EN175301) sous le code 11125168, ce sont des 3poles et PG9. le contenu du sachet est tel que photo jointe au ticket. Best, Nicolas
INC279514 | Workgroup | Incident | FW: 关于压力表MIX7A60B24/0751/AKPLB-0N6-10的停产升级替代咨询 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 13:24 | 2023-03-01 14:04 | 0 | 0 |  | Hello Zhang Wei, no problem for 10m, I don't know why they limited to 3 in the codification but it says 1 to 30m in the text so no worries. here is the code you need and it'sopen for sales: AKPL-B0N610 10264887 Best, Nicolas
INC279513 | Workgroup | Incident | AO_848326 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 12:57 | 2023-03-01 14:12 | 0 | 0 |  | Bonjour Cyrille, ces options sont dispos sur les MEX3/5: [Image] pas de remplissage par contre avec ces options. Best, Nicolas
INC279508 | Workgroup | Incident | TR: AMFD02N/0765 - 20503683 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 12:18 | 2023-03-08 12:19 | 0 | 1 |  | Adrien, oui, je confirme qu'on peut faire, ce sera surement avec des AMFD en provenance de schneider. AMFD02N/0765 ou codif schneider (je ne sais pas comment on va preferer le codifier. mais tu peux faire la demande de prix, jeconfirme que c'est faisable. Best, Nicolas
INC279472 | Workgroup | Incident | Replacement for MDX7-F | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-03-01 02:54 | 2023-03-01 12:00 | 0 | 0 |  | Hi Bee Tin, yes the new version of the MDX is the MX, please find enclosed the old datasheet. sorry we only have it in french. Best, Nicolas
INC279421 | Workgroup | Incident | FW: Rfq | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-28 08:09 | 2023-02-28 10:05 | 0 | 0 |  | Dear Bee Tin, this is an M17 2BX3.B35 the X is for the connection as I don't see it and can't define it. Best, Nicolas
INC279413 | Workgroup | Incident | WG: Angebot 20498192 / Alfalytik Anfrage 22008 Projekt | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 17:42 | 2023-02-28 09:42 | 0 | 0 |  | Sascha, it's quite a simple TW except for the option W4032 which I don't understand. I would go with this (drawing to be approved by the customer and done by our supplier for the 45mm long) Order code: AGFR-0230L.3J100.P AGF old Code / SAP code: Tpye: Stainless steel bored Thermowell R Extension: Without 0 Material: Stainless steel 1.4404 2 Male Threat for flange: G1/2" (standard for flange adaption) 3 Thread on instrument side: Fixed 0 female thread on instrument side: G1/2" L Inner Diameter Thermowell: 9 mm 3 Outer Diameter / outer diameter: 15 mm J Plunger lenght: 045 mm 045 Design: Straight shank P
INC279410 | Workgroup | Incident | TR: ED701 821.413.A35 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 17:36 | 2023-02-28 09:49 | 0 | 0 |  | Nathalie, c'était un: 0...20bar 4...20mA G1/2 EN837 pour montage separateur Atex DIN43650 0.4% -10...125C Donc attention, ca veut probablement dire qu'il nous manque le séparateur. sinon le remplacer soit par un Y913 ou un PBSN. Best, Nicolas
INC279409 | Workgroup | Incident | TR: FW: AKP 20230219 2053 20503837 Rectus.pdf; | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 17:30 | 2023-02-28 09:09 | 0 | 0 |  | Nathalie, on peut bien fournir, ci-joint l'offre de fantinelli, par contre a la codif il faudra verifier le numero d'option qu'on utilisera (surement la 0060). Best, Nicolas
INC279408 | Workgroup | Incident | TR: Zapytanie MG5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 17:24 | 2023-02-27 17:30 | 0 | 0 |  | Nathalie, Sorry but we can't turn the dials on the contact gauges so it's no. Best, Nicolas
INC279405 | Workgroup | Incident | PFMN monté sur D415 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 16:57 | 2023-02-27 17:29 | 0 | 0 |  | Cyrille, pour moi pas de soucis, par contre ce sera un D425 pour ce PN16/DN25 et il faudra nous donner la config, si vous divisez trop l'étendue de mesure le montage pourrait ne plus etre possible ou en D435. pour le code du PFMN, je ne connais pas cette connection éléc mais je vous fais confiance et ca n'a pas d'influence pour le montage. on devra aussi vous confirmer la precision globale de l'ensemble (probablement +/-0.5%) J'ai ajouté l'option de montage sur séparateur 0592. D425-2L2.14DE.B21 +9010 + 0592 + LRS4+ PFMN-54.BC3R.C1583.33200.4002 Best, Nicolas
INC279390 | Workgroup | Incident | Nouryon : Scale PSI - % | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-27 14:30 | 2023-02-27 14:44 | 0 | 0 |  | Hi Rudi, seems correct but it'll be an X14, it's a receiverr scale, it might even already exist, I ask. Best, Nicolas
INC279345 | Workgroup | Incident | Venturi pour mano différentiel | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 16:54 | 2023-02-27 16:41 | 1 | 0 |  | Bonjour Cyrille, malheureusement non, je suis désolé. le client peut trouver des orifices calibrés et nous demander l'instrument mais on n'a rien a proposer de notre coté. Best, Nicolas
INC279337 | Workgroup | Incident | BarFlex ancienne version | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 15:57 | 2023-02-27 16:40 | 1 | 0 |  | Cyrille, sur cette ancienne version, c'est equivalent au PPPP actuel qui veut dire qu'il y a eu une surpression. il doit etre endommagé si au redemarrage il ne revient pas a la normal et au mieux il faudra verifier la calibration. Best, Nicolas
INC279333 | Workgroup | Incident | KRK501_thermo-hygrograph | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 14:24 | 2023-02-24 14:51 | 0 | 0 |  | Bonjour cyrille, nous avons quelques plumes en fibre en stock (10263837). pour les papiers je ne suis pas sur, il faudrait qu'il nous disee du quel il a besoin (les codes sont dans la datasheet). Best, Nicolas
INC279326 | Workgroup | Incident | 20504846 / MFT7 VS DP370 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 12:27 | 2023-02-24 15:12 | 0 | 0 |  | Véronique, Ce sont tous des équivalent low cost au MFT7-A00.XXXN.62 La plupart seront en code 11 pour la pression (0.04MPa) Le code 51 dans leurs codes sont pour des cadrans spéciaux (pour les échelles en m3/h) Le code 58 est pour un logo client. Vu que la datasheet du DP370 est sur le site de fantinelli et avec le P2F, tu n'avais pas besoin de moi normalement. Best, NIcolas
INC279324 | Workgroup | Incident | Bloc contact - SAP code 10356999 - Nouryon (21106212) - RFQ 20230213 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 11:54 | 2023-03-08 14:01 | 0 | 1 |  | Hi Rudi, I checked the codes and it's from a supplier, you can order it but the price needs to be define by sales. I send them the new costs from the supplier and especially on the reference you need and tell them you'll ask them for the price. Best, Nicolas
INC279323 | Workgroup | Incident | EF6 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 11:45 | 2023-02-24 14:26 | 0 | 0 |  | Bonjour Cyrille, L'EF6 est en cours de phase out et il était conforme à la NF/EN 50155 le PP10R est maintenant validé et je peux fournir diverses conformités à l'EN 45545 selon le code du produit. je suppose qu'on a la même chose pour le PP20R mais c'est géré par la suisse. pour ce type de questions, je suis un peu le c*l entre deux chaises pour répondre, il vaudrait mieux t'adresser à Franck Ferrarese le PM elec, il aura surement un point de vue global. Best, Nicolas
INC279321 | Workgroup | Incident | BAUMER - Cliente: MEF -MTS Europe Srl: MTS RFQ - TSF100 - URGENTE | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-24 11:39 | 2023-02-24 14:42 | 0 | 0 |  | Hi Francesco, are you sure they didn't order a 738X option for the sliding connection ? because this adds 50mm to L=200mm. If they did it's perfectly normal. If they didn't order any type of connection, the length should just be L=200mm [Image] Best, Nicolas
INC279293 | Workgroup | Incident | Maintenance pressure gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-23 21:06 | 2023-02-24 10:39 | 0 | 0 |  | Hi Angel, the customer can be able to check the calibraion (by using a 4 times more accurate reference gauge on a test pressure bench at at least 4 points (0, mid scale, mx and back to 0). but if the gauge is out of calibration, he can only set it back with specifics pointer (micrometric, or friction) and with the necessary tools to remove the bezel ring. then the operation is setting the pointer back to zero and check again. if it's still of, most of the time a real recalibration must be done. concerning the reccurence of the calibration, the state of the art suggest that outside of extreme use, calibration once a year and once every 2 years after 2...3years of right calibration without having to do anything is sufficient. a replacement mandatory after 10 years is usual in the industry. Best, Nicolas
INC279253 | Workgroup | Incident | TBHI130-264.140.12T/0078_0751_0110_0776_9003_0460 With Liquid Filling | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-23 08:27 | 2023-02-23 09:29 | 0 | 0 |  | Dear Bee Tin, yes it's ok, I created it in 600mm without any issue so you can go ahead with the 460mm. 11262651 Best, Nicolas
INC279241 | Workgroup | Incident | Meura- PJ21000122 - Commande n° 72587 - Certificat | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-22 15:48 | 2023-02-22 17:23 | 0 | 0 |  | Sure Rudi, here are the certificates for the mountings and for the transmission fluid. Best, Nicolas
INC279215 | Workgroup | Incident | TR: 11251499 / PBMN-29B99RA11499400010/C02 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-22 09:51 | 2023-02-22 11:16 | 0 | 0 |  | Nathalie, je ne vois rien de fait chez nous, il faut demander directement a la BU. Best, Nicolas
INC279211 | Workgroup | Incident | FW: RFQ-Pressure gauge with limit switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-22 09:42 | 2023-02-27 14:29 | 0 | 0 |  | Hi Bee Tn, this is not something made in france, I think it's a BTIN product. the replacement with our products would be a: MS/R5 1D30B19 + DT2 2L*** but I miss a lot of data about the seal (materials coatings, connection ?) also our design of contact gaug is a bit differen and 100mm (I can't really tell if on the picture it's 100 or 150mm. Best, Nicolas
INC279210 | Workgroup | Incident | OFFRE 20504291 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-22 09:39 | 2023-02-22 11:33 | 0 | 0 |  | Véronique, pas de problème sur la définition par rapport a la demande client, et l'AKPL ne devrait pas poser de problème, on en a quelques uns de créés. par contre l'obligation de prendre du 2.5mm est uniquement pour les séparateurs, pour un montage direct des capillaires, tu peux proposer du 1.6 ou du 2.5mm (AKPLB0N501 ou AKPL20N501). MFT7-B00.X08K.62/9704_CM21 + AKPLB0N501 le SETP n'est en général pas nécessaire car c'est un réglage manuel en façade. Best, Nicolas
INC279198 | Workgroup | Incident | TR: Replacement ¦ ED 518-311.211/A25 ¦ Zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-21 17:51 | 2023-02-21 17:59 | 0 | 0 |  | Nathalie, oui le montage avec PBMN sera possible, a voir si le code reste comme ca. en prix il risque dêtre plus cher avec le PBMN, on peut aussi lui proposer l'option du Y913. Best, Nicolas
INC279155 | Workgroup | Incident | TR: 025004301001 ¦ odp ¦ Zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-21 10:12 | 2023-02-21 10:42 | 0 | 0 |  | Nathalie, tu peux proposer un Y91333B20R + 1537 (9010/0592) attention, DIN11851 et pas 11887 mais je crois que c'est juste une évolution de la même norme. le liquide de remplissage sera de la glycerine FDA (LRS1) pour du specifique comme l'option 1311 (huile d'amande douce) ce serait sur SPR. Best, Nicolas
INC279126 | Workgroup | Incident | TR: Bestellung Grenzsignalgeber Typ M12 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-20 15:48 | 2023-02-20 16:50 | 0 | 0 |  | Nathalie, Mouting is possible but I offered 2 versions, 1 with magnetic contact AND filled, the other with standard contact and dry. also I canged the DANC for a 13 according to the DIN norm as on the picture and added LRS1 filling. M17-7130.B24/0771 + AKPL-2F5302 + DANC13-2L2B01 M27-7133.B24/0771 + AKPL-2F5302 + DANC13-2L2B01 Best, Nicolas
INC279122 | Workgroup | Incident | TR: replacement ¦ ED 518-311.211/A25 ¦ Zapytanie | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-20 15:06 | 2023-02-20 15:48 | 0 | 0 |  | Nathalie, tu peux proposer un Y91333B20R + 1537 (9010/0592) attention, DIN11851 et pas 11887 mais je crois que c'est juste une évolution de la même norme. le liquide de remplissage sera de la glycerine FDA (LRS1) pour du specifique comme l'option 1311 (huile d'amande douce) ce serait sur SPR. Best, Nicolas
INC279110 | Workgroup | Incident | TED avec option "haute température" | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-20 12:24 | 2023-02-20 15:39 | 0 | 0 |  | Bonjour Cyrille, ci-dessous le dessin du refroidisseur avec les cotes principales: [Image] Best, Nicolas
INC279100 | Workgroup | Incident | DEFINTION ARTICLES URGENT HONEYWELL BERRI 20480666 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-20 09:57 | 2023-02-21 14:39 | 1 | 0 |  | Bonjour Véronique, après avoir revu ta demande avec Jérôme, et sachant que votre problématique principale est axée sur le prix, nous vous proposons de partir sur ces séparateurs en D912: - Les longueurs de capillaires sont à ajouter en remplacement du X. - Le choix des vannes est à valider avec le client (liste dans le fichier IC D9xx 2023 dernière Rev) les vannes décalées sont plus pour les petits anneaux - Ne pas oublier les caisses spécifiques Anneaux de rinçage pleine bride class 300 3 (montage sur D824-AXA.2229.G60.09) (x12) 316L vanne décalée longueur 100mm + bouchon D912MA0XA9602M29G0A0XX Anneaux de rinçage pleine bride class 150 3 (montage sur D824-AXA.2219.G60.09) (x8) 316L vanne décalée longueur 100mm + bouchon D912MA0XA9602M19G0A0XX Anneaux de rinçage pleine bride class 150 2 (montage sur D824-AXD.2217.G60.09) (x14) 316L vanne décalée longueur 100mm + bouchon D912LA0XA9602L17G0A0XX Best, Nicolas
INC279090 | Workgroup | Incident | Documentation CPEH | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-20 08:48 | 2023-02-20 15:26 | 1 | 0 |  | Bonjour Veronique, bien sur, c'est la balance que j'utilise aussi :) Best, Nicolas
INC279076 | Workgroup | Incident | Thermometre Bourdon sedeme | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 16:18 | 2023-02-17 17:11 | 0 | 0 |  | Hi Jurgen, sorry, not so much help here, te tag number can be TG for Thermometer Gas or for temperature gauge and it's a customer tag we can't find. the NQN n is not from us. Sorry we cant help you more. Best, Nicolas
INC279073 | Workgroup | Incident | MANO DIFF A CONTACT ATEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 15:39 | 2023-02-17 15:54 | 0 | 0 |  | Bonjour Adrien, oui le M33 est toujours disponible et toujours ATEX, ci dessous son certificat: Baumer_ATEX-MG5-M37-M33_EN-FR_CA.pdf (bourdon-instruments.com) "https://www.bourdon-instruments.com/medias/__secure__/Baumer_ATEX-MG5-M37-M33_EN-FR_CA.pdf?mediaPK=8905677078558" Best, Nicolas
INC279072 | Workgroup | Incident | Enquiry from Asia Paints - Model TED6 with diaphragm seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 15:36 | 2023-02-17 16:02 | 0 | 0 |  | Hi Anil, Not easy to read the pictures but the TED seems to be a: 11212493 TED-6.3.3.B81.R/0592_2037 the mounting: 11212494 TED-6.3.3.B81.R.DADL05-2L2A.01/2037_0592 weird thing is there is no paint application or silicon free options. Best, Nicolas
INC279071 | Workgroup | Incident | WG: Temperature Switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 15:30 | 2023-02-20 17:23 | 0 | 0 |  | Hi Carlo, I checked further the specs and I think we can potentially offer the RT2N. but there is 2 issues: our switch specs are slightly off in DC voltage they need to perform force tests and vibration at a third party that are quite heavy. please send them our switch table, if they don't accept the witch we can't get further. if they accept, there is slight changes compared to there design but it's probably acceptable as direct replacement of the switches in place. [Image] Best, Nicolas
INC279067 | Workgroup | Incident | WG: Help please | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 13:06 | 2023-02-17 16:13 | 0 | 0 |  | Sure, it's the code 10302664, you can see it in the screenprint and with the DTR n from alstom it's easy to find I can sow you. Best, Nicolas
INC279040 | Workgroup | Incident | MG5-5D60.B22/0693_0765_9941 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-17 07:30 | 2023-02-17 14:13 | 0 | 0 |  | Dear Bee Tin, everything is ok except the minimum will be 1.6bar: MG5-5D60.B16/0693_0764_9941 "https://servicedesk-sps.baumernet.org/wm/" Best, Nicolas
INC278999 | Workgroup | Incident | MEX3, 0-160 Bar With 2 bar Interval dial | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 10:45 | 2023-02-17 09:22 | 0 | 0 |  | Hi Bee Tin, with the bar bourdon, you'll have the 2bar divisions (example below): 10276727 MEX3-D20.A33 [Image] Best, Nicolas
INC278997 | Workgroup | Incident | FW: Pressure switch | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 10:30 | 2023-02-16 11:40 | 0 | 0 |  | Hi Bee Tin, this was the product and it's still ok: PBSN-1.2.B38.R.A2.44.02.2.1 11119880 Best, Nicolas
INC278998 | Workgroup | Incident | TBI avec aiguille suiveuse | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 10:30 | 2023-02-17 16:21 | 0 | 0 |  | Bonjour Cyrille, les mouvements de TBI surtout en raccord en bas et orientable sont montés via la tige et des ressorts sur l'OM. ils n'y aura pas assez de puissance dans l'axe pour une aiguille traînante, je suis désolé. Best, Nicolas
INC278990 | Workgroup | Incident | SEPARATEURS D201 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 10:06 | 2023-02-16 11:20 | 0 | 0 |  | Salut Véronique, le D201 est un DT2, voici les traductions de codes: DT2-2L20.0.230.00 DT2-2L60.0.233.00 Best, Nicolas
INC278982 | Workgroup | Incident | TR: CTX Resoultion | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 09:27 | 2023-02-21 08:57 | 1 | 1 |  | Nathalie, la resolution des CTX est de 14bits Best, Nicolas
INC278973 | Workgroup | Incident | CO2 clean option | Archive > BHFR Technical Support | Middle |  |  | Closed | Hedemark, Thomas Jydby [tjh] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-16 07:57 | 2023-02-16 11:15 | 0 | 0 |  | Hello Thomas, I should have explain, sorry. the tests for O2 usage of the Y91 were done by air liquide in 2000 according to the EN ISO 2503. it validates the resistance to adiabatic compression up to 320bar. only the FKM resisted and was validated. but the norms states that the effects are neglectable under 30bar. this means that up to 30bar, any type of seals we have are ok but over that and up to 320bar, only the FKM is validated. I cannot share the test report from air liquide but extracts (in french) could be shared if they need to see specific result or if they needed the test conditions for example. Best, Nicolas
INC278967 | Workgroup | Incident | confirm compatibility MFT5/DT8 | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 21:12 | 2023-02-16 10:14 | 0 | 0 |  | Hi Angel, I checked internally and we're blocked at 100bar top pressure, right now we can't go over it. Sorry for this, Best, Nicolas
INC278940 | Workgroup | Incident | VEO DEXIS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 11:30 | 2023-02-15 17:52 | 0 | 0 |  | Véronique, en double micro séparés on ne peut avoir qu'un écart de 3.5bar entre les deux donc une configuration possible a proposer serait: point de seuil: 2 bar A la baisse point de seuil: 5.5 bar A la montée Order code: RPPE-7G3.204/SETP RPPE old Code / SAP code: Sensing element and diaphragm: Type: PPN - Medium range -1 ... 600 bar 7 Microswitch: Microswitch 2 Separated SPDT G Hydraulic connection: G1/2" male 3 Pressure range (mbar): Pressure range (bar) 0.5 ... 10 bar 204 OPTIONS Accuracy / Setting: Set point adjustment SETP un RPPN ou Y irait aussi mais vu les designs sur la datasheet et les certifications ce sera à vérifier. Best, Nicolas
INC278936 | Workgroup | Incident | VS: Air Liquide - CO2 in Gas form | Archive > BHFR Technical Support | Middle |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 11:21 | 2023-02-15 17:31 | 1 | 0 |  | Hi Solveig, over 30bar, FKM is mandatory for us to offer it with O2 cleaning. we can offer several other options like with a diaphragm seal (example: DT8-2L20.0.230/0765) or a Y923 with a flush connection with no front sealing. Best, Nicolas
INC278933 | Workgroup | Incident | AKPL hors catalogue | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 11:12 | 2023-02-15 17:37 | 1 | 0 |  | Bonjour Cyrille, ce n'est pas impossible bien sur mais il faut designer et faire usiner la connection femelle. ca doit passer par un SPR. par contre il est aussi possible d'ajouter un raccord intermediaire: AKPL-20N601 (10264767) + ARIN-2608 (10265087). Best, Nicolas
INC278926 | Workgroup | Incident | Bride ABR | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 11:03 | 2023-02-15 17:42 | 1 | 0 |  | Cyrille, tu peux y aller et demander le prix ou la creation du code, pas de soucis sinon. Best, Nicolas
INC278904 | Workgroup | Incident | Temperature Gauge Dial Size 112 mm / LKH | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-15 03:18 | 2023-02-15 09:10 | 0 | 0 |  | Hello Bee Tin, we don't have 112 as far as I know but this is a free mounting (no flange or back mountitng) so I don't see any issue with offering 100mm. with 6mm stem diameter and 175mm length. the desing will be a blank flat dial instead of the grey one. Best, Nicolas
INC278895 | Workgroup | Incident | Montage Y923 - HRS | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-14 16:06 | 2023-02-17 14:06 | 0 | 1 |  | Bonjour Jacques, je ne conseille pas d'utiliser des raccords intermédiaire sur de lhydrogène (cette molécule est a très haut risque de fuites). pour les raccords souder, tu pourras en trouver chez Baumer: [Image] [Image] Best, Nicolas
INC278885 | Workgroup | Incident | the combination of MPG6-D30.D31+DT1-2L22.2.230.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-14 14:42 | 2023-02-14 15:07 | 0 | 0 |  | Hi Xu Qiang, this mounting is possible: MPG6-D30.D31+DT1-2L22.2.230.01 Best, Nicolas
INC278866 | Workgroup | Incident | Sliding Connector T 08406.0009 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-14 11:12 | 2023-02-14 15:03 | 0 | 0 |  | Hi Bee Tin, you can find those accessories on the website: T 08406.0009 ¦ Thermowells ¦ Bourdon Instruments (bourdon-instruments.com) "https://www.bourdon-instruments.com/fr/en/product-overview/accessories-temperature/sliding-screw-connection/t-08406-0009/p/37378" Best, Nicolas
INC278833 | Workgroup | Incident | TR: E913s and E933s | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-14 08:36 | 2023-02-14 09:25 | 0 | 0 |  | Nathalie, no problem for the Y913, see below the codes, for the Y933, there need to be a check with the customer: Y933-138-B2COR (probably SMS FL DN 38 + nut; B2C -1...4bar will be replaced by B77 -1...5bar; the O at the end I don't know) Y913-33-B19A Y933-138-B19A (not sure for the 1). Y913-33-B19A the Y933 might not be codified anymore as Y933 but Y913 + DAEL. Best, Nicolas
INC278805 | Workgroup | Incident | WG: Richiesta di informazioni tecniche su collegamento tubi ai manometroi Bourdon DN63 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-13 10:36 | 2023-02-13 11:19 | 0 | 0 |  | Hi Carlo, good practice for mounting G connections is to put the aluminum gasket at the bottom of the female thread and mount the male with a 40-45Nm torque. no need for teflon tapping. Best, Nicolas
INC278804 | Workgroup | Incident | MTBF PBSN | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-13 10:21 | 2023-02-13 11:02 | 0 | 0 |  | Please find enclosed the MTTF Best, Nicolas
INC278803 | Workgroup | Incident | MEP5, With SS304 Case Material, Range 0 to 20 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-13 10:21 | 2023-02-14 10:27 | 0 | 0 |  | Hi Bee Tin, here are our answers: 1) Is MEP5 with old case in SS304 still available ? yes, you can go ahead for the moment. 2) If not, for the MEP5 with new case, is range 0 to 20 bar dial available ? OK, it's available 3) Out of the 58 nos, 8 nos of the gauge are for Oxygen service, with range 0 to 20 bar OK, with option 0765 4) Out of the 58 nos of gauge, 5 nos are for Hydrogen gas, pls advise do we need any special option code for Hydrogen gas service ? please find the validated options from air liquide for use on H2 (no rules apply for now, please make sure the customer validates as he's the expert on his process). 5) For this project there is also 18 nos of MT7 with 3/8 NPT process connection, pls advise if we are able to offer with this process connection. to be checked with the supplier, please define the products and we will check and ask for offer. Best, Nicolas
INC278796 | Workgroup | Incident | Raccord Y923 application hydrogène | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-13 09:24 | 2023-02-13 11:28 | 0 | 0 |  | Bonjour Jacques, les raccords du Y923 en G1/2 1 et 2 joints que tu as en plans, sont disponibles avec le revêtement or pour application H2. après, sur des mixes H2 ou sur des faibles pressions ou températures, un raccord affleurant 316L ou même un raccord standard peuvent être suffisant. et il est possible via une SPR (validation rapide) de demander le revêtement or sur d'autres tailles de raccords affleurants. Il n'y a pas encore de normes sur ces applications ce qui rends les choses compliquées et nous laissons le choix et la responsabilité a l'utilisateur expert de son process en attendant un cadre réglementaire. Best, Nicolas
INC278777 | Workgroup | Incident | manomètre à contact chez FAREVA, | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 15:51 | 2023-02-13 14:47 | 0 | 0 |  | Cyrille, je ne vois pas vraiment de problème, que la résistance du circuit soit infinie quand le circuit est ouvert ça me parait normal si on est sur un contact NF (si s'en est bien un). ce qui est dommage c'est que je ne vois pas la valeur quand ils passent le contact, c'est ça qui aurait été intéressant. sinon, je conseillerais quand même d'utiliser une vis frein ou un raccord amortisseur, sinon les contacts vont charbonner en peu de temps. Best, Nicolas
INC278764 | Workgroup | Incident | Contact Gauge With Diaphragm Seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 10:15 | 2023-02-13 10:50 | 1 | 0 |  | Hi Bee Tin, this one seems indian to me so you can check this out directly with them but otherwise we can offer an MS5/MR5/MG5 fitted with a DT2 with tantalum diaphragm. Best, Nicolas
INC278762 | Workgroup | Incident | AW: Technica question | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 10:09 | 2023-02-14 17:31 | 1 | 1 |  | Answered via teams. also the switches are spdt so the wiring on one side or the other shouldn't change the outcome. Best, Nicolas
INC278741 | Workgroup | Incident | Recall: MEP5-D30.B26/0191_0679_0763 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 04:15 | 2023-02-13 10:33 | 1 | 0 |  | Solved
INC278740 | Workgroup | Incident | MEP5-D30.B26/0191_0679_0763 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 03:45 | 2023-02-13 10:33 | 1 | 0 |  | Dear Bee Tin, the new MEP5 only exists in 316L: MEP5-430.B26/0191_0679_0763 11262333 Best, Nicolas
INC278739 | Workgroup | Incident | MCF5-D30.N03/0693_0651_0078_9941+ASIPD2L3 / Neste | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-10 03:18 | 2023-02-13 10:26 | 1 | 0 |  | Dear Bee Tin, te cooling tower will not have an effect on the function, no worries. concerning the temperature, I'm affraid it might still be a bit high but the customer could validate on his side (we assure only a 400C max temp). Best, Nicolas
INC278729 | Workgroup | Incident | BAIT ¦ Fluid processing ¦ Manometer with miniclamp | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-09 15:03 | 2023-02-13 10:21 | 1 | 0 |  | Hi Erald, not for now, the request is here since long for the micro clamps but the very small diaphragm coupled with the high volume in a tube makes it very complicated. Sorry for that, Best, Nicolas
INC278725 | Workgroup | Incident | TR: Request - pressure gauges | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-09 14:27 | 2023-02-13 10:20 | 1 | 0 |  | Hello Solveig, in fact FM/UL are mainly for electrical components and will not directly apply to gauges. concerning the rest, it seems to me about our quality management system. I would reply with ISO9001 and 9100 for quality management (most commong in europe). and ATEX for the safety. It should cover with more european standard, the topics they need. Best, Nicolas
INC278710 | Workgroup | Incident | TR: RFQ // MEX3-D50.H22/0059 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-09 12:24 | 2023-02-13 10:14 | 1 | 0 |  | Nathalie, voici le code: 11122781 DRC63-811.133.075/0059_0022 Best, Nicolas
INC278698 | Workgroup | Incident | TR: 10265089 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-09 10:51 | 2023-02-13 10:10 | 1 | 0 |  | Nathalie, je trouve bien la même chose mais je ne gère pas les prix, à voir avec Nadia ou Stéphane. Best, Nicolas
INC278687 | Workgroup | Incident | TR: DDP 147675 URGENT ANNULE ET REMPLACE | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-09 09:24 | 2023-02-13 11:30 | 1 | 0 |  | Céline, je ne vais pas beaucoup pouvoir t'aider la dessus, il manque un chiffre sur le code et ce commentaire ne vient probablement pas d'un problème technique mais surement d'un retour client, c'est surement spécifique a ce code en particulier. Best, Nicolas
INC278676 | Workgroup | Incident | Error Report | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-08 18:00 | 2023-02-08 18:02 | 0 | 0 |  | testWD
INC278674 | Workgroup | Incident | Error Report | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-08 17:18 | 2023-02-08 17:59 | 0 | 0 |  | test WD
INC278668 | Workgroup | Incident | Error Report | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-08 16:33 | 2023-02-08 17:59 | 0 | 0 |  | testWD
INC278667 | Workgroup | Incident | Error Report | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-08 16:30 | 2023-02-08 17:59 | 0 | 0 |  | test
INC278605 | Workgroup | Incident | appel d'offres WH-APD | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-07 12:09 | 2023-02-08 10:11 | 0 | 0 |  | Bonjour Cyrille, concernant ta demande, on ne monte pas de PBSN sur des separateurs, par contre on peut monter des Y91. pour le revetement il n'est pas possible sur un D821 mais je te propose plutot un D4xx: exemple: Y91-3.3.3.B77.R/2164/0592 + ZPX1-007 + D435-2L2.14DC.B2010/9010 pour les TW, j'ai demandé a Jerome et christophe qui bossent sur les EPC mais ils n'ont jamais vu ca, ca nous parait tres bizarre, est ce que tu pourras revoir avec le client, c'est surement une erreur. Best, Nicolas
INC278591 | Workgroup | Incident | TR: Question Prix | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-07 09:57 | 2023-02-07 16:06 | 0 | 0 |  | Nathalie, Validé avec Charles, par contre ce sera plus cher qu'un 95mm pour peu de différence: TBH100-212.150.20T/9003/0100 Best, Nicolas
INC278573 | Workgroup | Incident | devis manomètre ferroviaire | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 17:30 | 2023-02-06 18:05 | 0 | 0 |  | Bonjour Cyrille, ce ne sera pas évident de retrouver la référence et elle sera obsolète. mais nous pouvons probablement proposer des remplacements en MFE (41 ou 51 selon le diamètre). il faudra valider l'interchangeabilité avec la datasheet, si un point bloque nhésite pas a revenir vers moi, nous avons fait aussi des exécutions spéciales. Best, Nicolas
INC278571 | Workgroup | Incident | TR: Angebotsanfrage Drucksensor | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 17:03 | 2023-02-06 17:59 | 0 | 0 |  | Nathalie, c'est possible selon le code suivant: PBMN-24B79RA11402202000/1627_1359_9010_LRS1 (accuracy 0.5%) Best, Nicolas
INC278559 | Workgroup | Incident | TR: Référence pour les douilles de différentes longueurs | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 12:03 | 2023-02-06 12:18 | 0 | 0 |  | Hi Nathalie, here is what we can do, for special length, he'll need to go to AGW or AGFR: [Image] Best, Nicolas
INC278556 | Workgroup | Incident | TR: ELCON RFQ nito | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 11:24 | 2023-02-07 09:41 | 1 | 0 |  | Ok Nathalie, tu peux offrir ce code (accord Pascal Baudet): 10270472 un code récent et similaire est dispo pour référence de prix si tu as besoin: EA4-AF9.B33R/0765 11259785 Best, Nicolas
INC278555 | Workgroup | Incident | TR: Question Monomètre | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 11:12 | 2023-02-06 12:13 | 0 | 0 |  | Nathalie, vu avec FENI, on fait extrêmement peu de ces contacts, donc a part le 4901, aucun n'est ouvert. le 5912 (magnetique) ne sera pas possible mais si il y a quelques pieces, il peut faire une SPR pour un 4912 (simple). Best, Nicolas
INC278541 | Workgroup | Incident | TR: MEX5-430.B22 + DAPH29_9912_9941 / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-06 10:30 | 2023-02-06 11:00 | 0 | 0 |  | Nathalie, the codes you have for the endpieces are right and open, they need to choose depending if it's for a pipe (11039934) or for a tank (11043353). we do not provide the clamp and seals as it's more interesting for the customer to order it in big qty as it's standard. Best, Nicolas
INC278498 | Workgroup | Incident | Demande ARD spé / RAMUS | Archive > BHFR Technical Support | Middle |  |  | Closed | Chalamet, Gregory [grch] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-03 12:09 | 2023-02-07 14:04 | 0 | 0 |  | Gregory, j'ai eu Mr Imhof qui veut que toutes les demandes de produits indiens passent par le PM, je transfère la demande àJörg. désolé de ne pas pouvoir aider. Best, Nicolas
INC278497 | Workgroup | Incident | TR: Cable pour BARFLEX - STIB | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-03 11:48 | 2023-02-03 15:41 | 0 | 0 |  | Celine, c'est le code 10309319. Best, Nicolas
INC278457 | Workgroup | Incident | 30147313 / 43418-02 | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-02 14:33 | 2023-02-02 17:28 | 0 | 0 |  | Véronique, J'ai fait faire 2 calculs de stress pour vérifier, je ne comprends toujours pas pourquoi il y a ce marquage car le TW peut résister à bien plus. Enfin en tout cas a 20C, j'ai vérifié pour 400bar et ca passe sans problème (pression statique pas de flux). Best, Nicolas
INC278436 | Workgroup | Incident | step file | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-02 08:48 | 2023-02-02 09:15 | 0 | 0 |  | Hi Kerstin, here is the drawing: Best, Nicolas
INC278426 | Workgroup | Incident | OP Centralair 01022023 20500589 ¦ RP2-NL.345.D055.2L1C.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Chicharro, Dionisio [chdi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-01 17:18 | 2023-02-03 12:18 | 0 | 0 |  | Hi Dioni, I'm sorry but this is just below the minimum for mounting. the min mounting is: RP2-NL.346.D055.2L1C.01 (min réglage 1bar, Pmax 40bar) Best, Niclas
INC278393 | Workgroup | Incident | MEX5 Special Dual Scale / Kumpanart | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-02-01 11:15 | 2023-02-03 16:12 | 0 | 0 |  | Dear Bee Tin, I'm sorry but none of the scales exist, you'll have to ask for creation. Also for thee second scale, please make sure it's right, above 0 (positive pressure) it's quite unusual to find those Hg units, it's usually bar or psi. Best, Nicolas
INC278320 | Workgroup | Incident | WG: ELCON RFQ ludan | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-31 11:57 | 2023-02-06 11:45 | 0 | 1 |  | Hi Carlo, here is the technical offer for their request. I based the description on the tank height as the setp are defined with several references we don't know and the calibrated ranges are not always matching the rest. the values of min/max SETP are estimated and must be reviewed when we know more about the values of hysteresis, setp and we could play with the type of switch depending on the power to cut. Best, Nicolas
INC278307 | Workgroup | Incident | BAIT: Olsa C22133 - Pressure Trasmitter+Manometro | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-31 10:24 | 2023-01-31 10:45 | 0 | 0 |  | Hi Erald, DRCE is a contact guge, not a 4...20mA output, but the DRCG exists for this. the mounting 4bar + DANC DN38 is not possible but here are possibilities. you can also refer to the MEX5 for the mounting in the seals guide. DRCG101-211.133.055 + DANC02_2L2A.01 (0...4bar + DN51) DRCG101-211.133.065 + DANC01_2L2A.01 (0...6bar + DN38) Best, Nicolas
INC278286 | Workgroup | Incident | ETTN ?? | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-30 16:57 | 2023-01-30 17:50 | 0 | 0 |  | Hi Rudi, the last one is open and on the other ones, the reason for closing is because there was no movement for a long time. you need to ask for the creation of the 2 first codes but it should not be a problem. PLease ask for the 2 first codes with the 3rd one for reference and you can ask for the lead time based on the last code already, it should be pretty accurate. Best, Nicolas
INC278283 | Workgroup | Incident | TR: Absolute manomètre MEX3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-30 16:27 | 2023-01-30 17:46 | 0 | 0 |  | Nathalie, it'll not be absolute, but a pseudo-absolute is possible by using a X59 (-1...0bar) and have a special dial from -1000mbar to 0. Best, Nicolas
INC278280 | Workgroup | Incident | piquage TriClamp | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-30 15:57 | 2023-01-30 17:44 | 0 | 0 |  | Bonjour Cyrille, comme ca je n'ai rien mais je sais qu'ils ont les raccords a souder chez Baumer dans les accessoires de type ZPX, je n'ai pas trouvé le DN51 mais si tu as un contact Baumer, il t'aidera plus facilement. Best, Nicolas
INC278224 | Workgroup | Incident | Confirm accuracy with case filling in model MPE6 | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-28 07:18 | 2023-01-30 10:00 | 0 | 0 |  | Hi Anil, the accuracy is : Grade 2A (according ASME B40.1, 0.5%) but the MPE6 is being phased out and if it's not already fully done, I strongly reccomend going for the new MPG6 (redesigned) with the PBT case (also UL94 V0 certified). Best, Nicolas
INC278215 | Workgroup | Incident | WG: Inquiry for pressure gauges and seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-27 15:15 | 2023-01-27 15:37 | 0 | 0 |  | Hi Carlo, for CEN-CENELEC, it seems to be equivalent to ISO and IEC that we have (see link below): ISO and IEC - CEN-CENELEC (cencenelec.eu) "https://www.cencenelec.eu/european-standardization/international-cooperation/iso-and-iec/" concerning the options, we can't have rotating connection but we can offer overpressure protection (RAS and check the seal max pressure) and we can offer restictor screw in the MEX connection. Best, Nicolas
INC278207 | Workgroup | Incident | technical Request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-27 14:06 | 2023-01-27 15:45 | 0 | 0 |  | Hi Sascha, yes it is possible, please find below an example for the codification of the special connection (N + 9414) and the rest of your code is fine. [Image] Best, Nicolas
INC278205 | Workgroup | Incident | TED-6.3.B15.R/0606_0619_2037 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-27 13:54 | 2023-01-27 15:54 | 0 | 0 |  | Hi Rudi, no problem, you can use this code as an example: 11222165 TED-6.3.5.B22.R/0606_2037_0619 Best, Nicolas
INC278198 | Workgroup | Incident | SIL 1 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-27 12:03 | 2023-01-27 15:57 | 0 | 0 |  | Hi Jurgen, sorry, SIL1 is a safety certificate that is specific, you customer need to search for SIL1 replacement only. we had 1 product that was SIL2 but we stopped the certification, we have no replacement for this customer. Best, Nicolas
INC278178 | Workgroup | Incident | Maximator MEP3 and MEP5 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-26 17:06 | 2023-02-01 16:28 | 0 | 1 |  | Sascha, je clos la demande dans le ticketing, la demande de production des pièces est faite. je tiens tout le monde au courant des qu'on a des délais. Best, Nicolas
INC278157 | Workgroup | Incident | Turndown sur CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-26 11:06 | 2023-01-26 15:09 | 0 | 0 |  | Bonjour Jacques, ce n'est pas une impossibilité technique mais les CTX sont des produits basiques standards qui ne sont pas destinés a ce genre de manipulations. de plus le changement de prix induit par l'option le rendra probablement inintéressant par rapport a un PBSN. Best, Nicolas
INC278150 | Workgroup | Incident | MEX2 With Dual Scale in Bar/kPa | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-26 10:27 | 2023-01-26 15:17 | 1 | 0 |  | Hi Bee Tin, sorry it doesn't exist in MEX2, and the code will be MEX2-D20.L19. Best, Nicolas
INC278102 | Workgroup | Incident | 10297526 step file | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-25 15:18 | 2023-01-25 15:22 | 0 | 0 |  | Hi Kerstin, sure thing, here is the .stp (in the ticket) Best, Nicolas
INC278094 | Workgroup | Incident | TR: Y92-3.5G.B31.R/2164_0829_7912 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-25 13:45 | 2023-01-25 14:11 | 0 | 0 |  | Nathalie, l'option 7912 inclue le dégraissage/nettoyage, donc pas besoin d'option en plus. et pas de soucis pour le Q003 qui doit être demandé à la commande comme d'habitude et qui s'appliquera aux parties en contact. Best, Nicolas
INC278074 | Workgroup | Incident | BAKR | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-25 10:30 | 2023-01-25 11:15 | 0 | 0 |  | Nathalie, ci-dessous la configuration en MEX5: MEX5D30B22 Best, Nicolas
INC278064 | Workgroup | Incident | TR: Request for price and article no. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-25 09:09 | 2023-01-25 09:15 | 0 | 0 |  | Nathalie, c'est un thermo, ca devait etre stiko. Best, Nicolas
INC278045 | Workgroup | Incident | Customer Request MEP5 with ABR | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 16:30 | 2023-01-26 16:34 | 0 | 0 |  | Hi Sascha, it is possible, please send the price request to the sales dept: MEP5431B77/0192/0059/9750/0661 + 54000-66-AS-04-0L to manufacture, we will need the customer LOGO and according to KHW, the customer request a Q003. Best, Nicolas
INC278043 | Workgroup | Incident | BAIT - PAF - TSSE100-211.467.82T.5066/738A | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 16:03 | 2023-01-25 09:39 | 0 | 0 |  | Hi Francesco, this is not an SPDT (an SPDT means a Single pole will close on different circuits depending on its switching state). here you have more of 2 SPST. here are the informations on the drawing but I don't much different possible wiring: [Image] Best, Nicolas
INC278039 | Workgroup | Incident | WG: Differential gauge for nuclear application | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 15:33 | 2023-01-25 11:00 | 0 | 0 |  | Hi Carlo, here is the certificate about the Seismic validation of the MX by an external LAB accordin to the RCCE: Best, Nicolas
INC278034 | Workgroup | Incident | BAIT - PAF - TBI100-291.146.13T | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 15:12 | 2023-01-25 10:00 | 0 | 0 |  | Hi Francesco, yes, you'll receive the sliding connection like on the picture, it's the digit 9 of the code: Order code: TBI100-291.146.13T TBI old Code / SAP code: Nominal size: DN 100 100 Case AISI 304 (1.4301), stainless steel: AISI 304 (1.4301), stainless steel 2 Process connection: With sliding connection Sliding screw connection G1/2" female 9 Immersion tube outlet: Bottom 1 Indicator: EN 13190 (Standard) 1 Immersion tube diameter: 6 mm AISI 316Ti (1.4571), stainless steel 4 Immersion tube length L1: 250 mm 6 Temperature range: Single scale inC 0 ... 160C 13T Best, Nicolas
INC278025 | Workgroup | Incident | TR: PFMN-64.BC2R.A1551332001000+DAIL162L23+LRS1 (SAP 11168120) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 12:33 | 2023-02-02 14:13 | 1 | 0 |  | Nathalie, le montage est possible et on a un CCR2 de 101.48 pour le DAIL162L23 PFMN-64.BC2R.A1551332001000 CCR2: 309.75 LRS1 IC: 34.8 le code peut être rouvert sur demande. Best, Nicolas
INC278018 | Workgroup | Incident | Certificat EMC pour EF6 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-24 11:09 | 2023-01-24 11:53 | 0 | 0 |  | Bonjour Jacques, je trouve bien le même certificat que toi de 2018 par rapport a la norme NF EN 50121-3-2 (Novembre 2006). étant donné que l'EF6 est en phase out, nous n'avons pas prévu de renouveler les essais. par contre bien sur, les essais sur les PP10R seront fait selon les dernières normes en vigueur. Best, Nicolas
INC277960 | Workgroup | Incident | PBMN Low Pressure sur séparateur D820 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-23 12:48 | 2023-01-26 11:51 | 1 | 0 |  | Ok comme ci-dessous: PBMN-24B76RM11503201000_ D820-10L.42FG.B20_LRS1_0592_9010 (0.5% a 20C) Best, Nicolas
INC277875 | Workgroup | Incident | WG: AW: Alfalytik Anfrage 22008 Projekt | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-19 17:48 | 2023-01-25 15:18 | 0 | 1 |  | Hi Sasha, no problem for the AGF, here I did the first one AGFB-0216N.1W225.T (T: last 65mm from 26 to X) we need the diameter at the end of the cone but you can already send the request for price to your internal sales contact with X=15 which is the minimum (the price will not change with another value from 26 to 15mm) AGFB-0216N.1W225.T (T: last 65mm from 26 to 15) x3 AGFB-0216N.1W160.T (T: last 65mm from 26 to 15) x5 AGFB-0216N.1W225.T (T: last 65mm from 26 to 15) x3 last one is a standard threaded T8 or AGF with this, you need to ask for price and drawing from your sales contact and plan a validation of the drawings with the customer. Best, Nicolas
INC277874 | Workgroup | Incident | TR: 1935/2004 declaration of compliance. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-19 17:39 | 2023-01-19 17:59 | 0 | 0 |  | Nathalie, ci-joint au ticket les certificats: Best, Nicolas
INC277852 | Workgroup | Incident | Barflex 1000 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-19 12:00 | 2023-01-19 17:38 | 1 | 0 |  | Bonjour Cyrille, typiquement on peut proposer ce code, le 0.1%, fixe et sans joint sont obligatoire et un 1000bar absolu n'a pas de sens: 10270701 EA5AF0B41R Best, Nicolas
INC277836 | Workgroup | Incident | TR: BAIT ¦ Y92 with 150°C option | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-19 09:36 | 2023-01-20 16:27 | 1 | 0 |  | reponse de Antoine ci-dessous Oui cest réalisable, Le CCR2 est valable que pour 9 pièces (9 raccords en stock, ils nont pas tourné depuis 2013). CCR2(Y923.3K.B77.R/0410) = 65 Best, Nicolas
INC277795 | Workgroup | Incident | WG: Season´s Greetings 20497437_Flowrond | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-18 15:36 | 2023-01-19 17:33 | 1 | 0 |  | Hi Carlo, yes, we are not so much a manufacturer of steel/brass gauges. for the small diameter, up to 100mm, you can go with MIT. for the O2 gauges, I suggest MEX and for the 160mm MEX also. I think this is a request that will go also to chinese suppliers and I'm not sure we will be competitive but I can be wrong and with a mix of MIT and MEX we should not be too expensive for a european supplier ... Best, Nicolas
INC277668 | Workgroup | Incident | BSPT 1/8" - Inovyn Manufact. B. (21106377) | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-16 16:51 | 2023-01-17 15:40 | 0 | 0 |  | Hi Rudi, yes it is possible to manufacture this code: MEX3-6X0.B29/0751_0078_C0xx But, the fact that we have the drawing doesn't mean it's available, we have tone of drawings and some might not be available, that's the reason why we don't disclose too much the drawing. We can do this one but please ask Celin for the price of the adder with the number of units you'll order. Best, Nicolas
INC277652 | Workgroup | Incident | TR: Configuration MX7 ? - Engie Electrabel (50148724) - RFQ600058715 Pos.20 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-16 14:24 | 2023-01-16 17:28 | 0 | 0 |  | Celine, le code sera MX7-B30.X15H/9709/Dxxx (-0.5...0.5bar) (peut etre avec aussi la 9704) il faut aussi penser aux frais de creation du cadran. Best, Nicolas
INC277596 | Workgroup | Incident | BAIT ¦ Membrane Seal up to 1500 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-13 15:57 | 2023-01-16 09:16 | 0 | 0 |  | Hi Erald, I'm sorry, but the highest we can go will be 1000bar with a DT5. Best, Nicolas
INC277595 | Workgroup | Incident | TR: Help please | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-13 15:33 | 2023-01-13 15:46 | 0 | 0 |  | Nathalie, this code is for mounting of the MFE only (railway) see explaination document enclosed. Best, Nicolas
INC277579 | Workgroup | Incident | WG: Angebot 20496948 / 11301346-Matao-Anfrage Manometer Stickstoff | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-13 11:57 | 2023-01-13 12:00 | 0 | 0 |  | Hi Sascha, this length depends on the thread and a drawing can be sent separately, example below for G1/2: [Image] Best, Nicolas
INC277563 | Workgroup | Incident | TR: RFQ | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-13 09:06 | 2023-01-13 10:41 | 0 | 0 |  | Nathalie, Jerome a eu la meme demande il faut proposer en remplacement: Un Y91333B26R/2160 + raccord refroidisseur standard ou ASIPD Best, Nicolas
INC277520 | Workgroup | Incident | TR: DD de prix 2023 stp merci -- Offre de prix DEXIS | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-12 16:03 | 2023-01-12 16:27 | 0 | 0 |  | Celine, on peut faire et le montage est OK. pour le prix, tu peux prendre le CCR2 du 11238214 (32.37) et ajouter le 9941 avec l'IC du P2F. Best, Nicolas
INC277454 | Workgroup | Incident | TR: MJX7 w/ diaphragm | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-12 11:27 | 2023-01-12 15:13 | 0 | 0 |  | Nathalie, the mounting is possible and the MJX exists but with the old D295 code, it could be for 3", 4" or 5" pipes so one of the mounting belo, to be confirmed with the customer. 10290929 MJX7-461.X19/0679_9709_0693_D000 + D295-2N20.000 (No bottom part for replacement of an existing unit) D295-2N20.21K (for 3" pipe size) D295-2N20.21L (for 4" pipe size) D295-2N20.21M (for 5" pipe size) + Filling fluid Best, Nicolas
INC277323 | Workgroup | Incident | Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-11 11:33 | 2023-01-11 14:33 | 0 | 0 |  | Hi Kerstin, Here is what we can offer: ZPN--206SEM/2237_SETP I added the SETP as it's usually requested for nuclear and the option 2237 is standard with the SEM (it's the metallic cable gland) we don't have it in G3/4 but in G1/2 + 1/4NPT female as standard. but it should not be an issue as it's adapted to the nuclear Alstome standards otherwise an adaptor G1/2/G3/4 can be used. if they absolutely need a G3/4, please start an SPR request. Best, Nicolas
INC277314 | Workgroup | Incident | special range | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-11 11:03 | 2023-01-11 14:37 | 0 | 0 |  | Hi Sascha, we can offer those ranges: [Image] [Image] the dual scale mbar/psi need to be requested and created with a dial creation fee. Best, Nicolas
INC277308 | Workgroup | Incident | AW: Ask | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-11 10:18 | 2023-01-11 14:52 | 1 | 0 |  | Hi Carlo, if there is no movement, static 1000bar and 100C should be ok for the G1/2 connection and for the stem. if there is movement or if the customer needs confirmation, then needs to ask for a wake frequency calculation (stress calculation) that we can do here and should provide all the elements about the process (fluid, density, velocity, flow etc...). Best, Nicolas
INC277231 | Workgroup | Incident | WG: DRC 100/811.133/525/0677/1547/1359/1310 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-10 16:24 | 2023-01-10 17:51 | 0 | 0 |  | Hi Kerstin, with the obsorption of KHW, there is a few changes and the MEX5 that replace the DRxxx has slightly more volume so the min pressure for mounting with the diaphragm seal is a bit higher. you can offer the MEX5 + 1547 according to the datasheet of the 1547 (min 25bar with a MEX5 or 6bar with a MEX3 or MX35). Best, Nicolas
INC277228 | Workgroup | Incident | FW: DP INSTRUMENT. TECHNICAL QUERY | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-10 16:06 | 2023-01-10 17:23 | 0 | 0 |  | Hi Daniel, no real answer on this, for the standard gauges, it usually happens because during the travel there was changes in atmospheric pressure but here there is no impact. all I can advise before going to a return policy is to check it with the gauge vertical (it is how it's calibrated and how it should be used) and also do a few taps with the finger on the side of the case in case it was just a particle of dust gripping the movement. if after those standard precautions the gauge is still off-zero (more than the tolerance) you can go to Nathalie for a return procedure. Best, Nicolas
INC277225 | Workgroup | Incident | BAIT ¦ Precision Fluid control ¦ RICHIESTA DI OFFERTA RESTAG | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-10 16:00 | 2023-01-10 17:27 | 0 | 0 |  | Hi Erald, please find enclosed the datasheet of the E71TR, I guess it's now replaced by the TE2 from Baumer. Best, Nicolas
INC277136 | Workgroup | Incident | product datasheet ordering code A 06443.0010 (Nipple connection) in Brass | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-10 09:30 | 2023-01-10 14:08 | 1 | 0 |  | Hi Anil, I found out the complete datasheet where you extracted the information (enclosed to the ticket). but there is not more than what you have. Best, Nicolas
INC277127 | Workgroup | Incident | WG: AKP_INQ20230109_2130 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-10 08:57 | 2023-01-10 10:18 | 0 | 0 |  | Hi Carlo, for the titanium, we can't do it, but the welding community seems to clearly state Stainless with titanium welding should not be possible. for the second one, a D803 should fit, I'm expecting the answer on the M6 but please check with your customer for the face type and capillary. Best, Nicolas
INC277057 | Workgroup | Incident | Confirm Monel-K500 Bourdon tube & socket in 100mm dial size, model MMN5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-09 14:03 | 2023-01-09 14:14 | 0 | 0 |  | Hi Anil, our MMN is in Monel 400 please offer our standard with the deviation. we have also the option of the turret Plastic case with monel tube and socket (Monel 400) with the MPJ6 developping new tubes is a huge decision and we can't do it for less than long terms big markets. Best, Nicolas
INC276835 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-05 14:00 | 2023-01-05 14:41 | 0 | 0 |  | Answered directly after checking the need: [Image] Best, Nicolas
INC276792 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-05 09:03 | 2023-01-26 16:20 | 0 | 1 |  | Nathalie, c'est possible on a vu avec Sylvain: PBMN-24B12RA11402201000 / DANC02.2H6A.01/9010 (0.5% a 20C) Best, Nicolas
INC276765 | Workgroup | Incident | WG: Conflict Minerals Request 2022 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-04 17:15 | 2023-01-04 17:58 | 0 | 0 |  | Hi Sascha, with the thousands of components we buy, we are not able to trace back to where the raw extracted material comes from. we comply with the european rules in terms of suppliers and will not by ourself buy anything from conflict regions. we comply with Baumer ethic chart and also try to work local from our factories. it is not possible to fill this for all customer (we have quite a few requests for this type of documents). Best, Nicolas
INC276740 | Workgroup | Incident | TR: Request for price ordering code A 06443.0010 (Nipple connection) in Brass | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-04 15:03 | 2023-01-04 15:42 | 0 | 0 |  | Exactement, et félicitation d'avoir trouvé !! Best, Nicolas
INC276562 | Workgroup | Incident | E913.033.B79R/1545/1310/1359 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-03 13:00 | 2023-01-17 15:40 | 1 | 1 |  | Hello, the code is created: 11261649 Y91-3.2.3.B79.R/2151_1545_1310_1359_9010 Best, Nicolas
INC276538 | Workgroup | Incident | Enquiry for a replacement ¦ ZAPYTANIE OFERTOWE NESTLE | Archive > BHFR Technical Support | Middle |  |  | Closed | Kondras, Kamila [kkam] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-03 11:24 | 2023-01-03 14:13 | 0 | 0 |  | Hello Kamila, you're right, it seems like a good replacement, I added a longer cable (10m as in the PFSX) and the detailled SETP instructions. the main difference will be that there is no SIL2 approval. TED - 5. 3. 5. B77. R. 0609/SETP ADJUSTMENT OF THE SET POINT LIKE SPECIFICATION: SP1 0.050 SP2 0.050 tS1 0.0 tS2 0.0 UPdo1 Do UPdo2 Do H1 0.030 H2 0.030 tH1 0.0 tH2 0.0 OAut NO nCod YES Best, Nicolas
INC276526 | Workgroup | Incident | M317-2B-6-B16-B22 / Siti | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-03 10:18 | 2023-01-03 12:10 | 0 | 0 |  | Hi Bee Tin, this code exists and is open, so no problem: 11167730 M31-726.F16H Best, Nicolas
INC276483 | Workgroup | Incident | MPG6 with "0" / Theerasak | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-03 04:27 | 2023-01-04 15:21 | 0 | 0 |  | Dear Bee Tin, here is the standard ANSI dial for our new MPG6 redesign. so it should fit perfectly your customer's needs: [Image] Best, Nicolas
INC276469 | Workgroup | Incident | TR: inquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-02 14:57 | 2023-01-02 16:30 | 0 | 0 |  | Nathalie, oui, c'est du 1/2NPT des 2 coté, ça se monte. Best, Nicolas
INC276460 | Workgroup | Incident | TR: quote - MEX3-DX0.H410133C050 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2023-01-02 14:27 | 2023-01-02 16:33 | 0 | 0 |  | Nathalie, Ils doivent maintenant voir avec ESP mais de toute façon, ces connections tubulaires ne sont pour l'instant plus possible a ces hautes pressions, les codes ont étés bloqués. Best, Nicolas
INC276350 | Workgroup | Incident | Leaking D621 on HCL & FeCl3 - Taminco (21106806) - MEX5-D30.B22.D621-2L21.4DE.BF002 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-23 10:15 | 2022-12-23 11:34 | 0 | 0 |  | Rudi, considering the damages, we need to offer a stronger solution here is what we can offer and the changes details: MEX5-D30.B22.D621-2L72.4DE.BF002 (coupelle usinée 316L et revêtement tenu par les vis) diaphragm tantalum instead of stainless PTFE lining of the diaphragm stronger and larger top maintaining part in 316L machined (stronger and more chemically resistant). please offer this as a technical solution and ask for the temperature (it might be another factor that caused the leaks). Best, Nicolas
INC276342 | Workgroup | Incident | TR: question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-23 08:48 | 2022-12-23 10:06 | 0 | 0 |  | Nathalie, non, nos séparateurs et liquides de remplissages ne tiendront pas au dessus de 400C (avec du LRS9) Best, Nicolas
INC276318 | Workgroup | Incident | TR: Prix-IC ? - TED-5.6.9.B20.A - Ineos RFQ | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-22 15:30 | 2022-12-22 17:28 | 0 | 0 |  | Celine, oui c'est OK pour faire le TED-5.6.9.B20.A. Best, Nicolas
INC276313 | Workgroup | Incident | rfi | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-22 14:33 | 2022-12-22 17:08 | 0 | 0 |  | Hi Angel, I'll need more informations, this device can be diffferent measuring devices. I send you teams, I'm available to talk about it but we'll definitely need more information. Best, Nicolas
INC276293 | Workgroup | Incident | PNPP2/3252B - Demande de renseignements - Temporisation sur indicateur numérique type DX Display | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-22 11:42 | 2022-12-22 12:07 | 0 | 0 |  | Bonjour Véronique, ce n'est pas le sujet le plus simple vu que ce produit est spécifique, plus fabriqué et que la plupart des interlocuteurs ayant bossé dessus ne sont plus la ;) mais ce que je proposer c'est de caler un petit teams avec ce monsieur pour comprendre le réel problème (car pour moi les branchements dans la notice sont assez clair). je suis dispo cet aprem ou demain, essaye de voir si il a une 1/2h et mets moi une réunion. Best, Nicolas
INC276148 | Workgroup | Incident | DS-PLM-MAIA1-HDT-EDD-INSTR-103 DATASHEET MESURE PRESSION DUPLEX_.pdf | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-20 17:06 | 2023-01-12 14:48 | 0 | 1 |  | Véronique, J'ai vu Jérôme pour cette demande par hasard et il nous a bien débloqué. les montages sont validés par la datasheet du D918. les prix t'ont été envoyés par Jérôme en fin d'année. Nous conseillons également de proposer de l'HC a la place de l'inconel (pas possible a sourcer) et de proposer du 12m de capillaire conforme a la REQ plutôt que du 9m. Best, Nicolas
INC276142 | Workgroup | Incident | TR: PJ095454 - TRANSMETTEUR DE PRESSION E913 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-20 16:18 | 2022-12-20 16:33 | 0 | 0 |  | Bonjour Adrien, nous n'avons plus l'homologation marine ci-joint une lettre explicative pour le remplacement (voir si possible pour le client d'accepter). Best, Nicolas
INC276131 | Workgroup | Incident | TR: BAIT ¦ Y913 electrical connection for M12 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-20 15:30 | 2022-12-20 16:32 | 0 | 0 |  | Nathalie, les PIN 1 et 2 sont bien le plus et le moins, ca a changé quand on a changé les cartes, l'ancienne datasheet n'est plus a jour. Best, Nicolas
INC276103 | Workgroup | Incident | Stub connector info | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-20 11:57 | 2022-12-20 17:09 | 1 | 0 |  | Hi Jurgen, if the hole for the welding is not yet done, I suggest to go with the AEKT. otherwise behind the code N 05018 0005 is the reference below from schneider (apparently) and also, I thought there was an O ring but in fact no, it's a metal/metal tightness. [Image] You can ask celine for the price and tell her I'm aware of the subject if she needs anything Best, Nicolas
INC276079 | Workgroup | Incident | TR: Question to MFE | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-20 09:39 | 2022-12-21 11:28 | 1 | 0 |  | Nathalie, ci-dessous le code du raccord intermédiaire avec une vis frein M3 non percée (l'amortissement se fait par les filets). cette solution est existante et déjà en usage mais n'est faite que pour les gaz, si c'est un liquide il faudra créer un raccord intermédiaire avec une vis M3 percée. 11222607 G3/8 M/F avec vis frein M3 non percée. Best, Nicolas
INC275953 | Workgroup | Incident | WG: AKP_INQ_20221214_1124 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-19 09:54 | 2022-12-19 14:31 | 1 | 0 |  | Hi Carlo, we have drain and vent plugs in 1/2 and maybe 1/4NPT, but the one in the drawing is not a plug, there is a hole in it. we don't have it with a hole and I don't understand the purpose. Can you check this question with the customer ? Best, Nicolas
INC275865 | Workgroup | Incident | Mano UK et USA | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-16 11:39 | 2022-12-16 15:43 | 0 | 0 |  | Salut Hervé, on a pas l'UL (et ce n'est pas très important car c'est orienté électronique) par contre on respecte l'UKCA, à partir de janvier, un marquage sera ajouté sur les boites de produits et sur les papiers (facture je crois). et notre certificat CE standard contient la norme UKCA (qui n'est pas nommée comme ça). il n'y a pas d'option a ajouter pour être conforme. Best, Nicolas
INC275810 | Workgroup | Incident | TR: Y91-3.3.5.B81.R.DT2-2L20.0.230.02 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-15 17:27 | 2022-12-15 17:42 | 0 | 0 |  | Nathalie, Le montage est relativement OK, je suggère d'utiliser un joint de cellule standard comme ce ne sera pas en contact Y91-3.3.3.B81.R.DT2-2L20.0.230.02 Best, Nicolas
INC275734 | Workgroup | Incident | TED : option 9022 (Canal de pression dans raccord ø10 mm) | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-15 09:57 | 2022-12-15 11:16 | 0 | 0 |  | Bonjour Cyrille, c'est possible en G1/2, pas en 1/4NPT. Best, Nicolas
INC275699 | Workgroup | Incident | TBI100-212.164.70T with an TBL100 frontflange or Ringholder | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 17:12 | 2022-12-20 15:22 | 0 | 1 |  | Hi Sascha, yes, it is possible it's a TBI100-212.164.70T + TBL front 10330200. I send the cost to Nathalie so she can make you an offer. Best, Nicolas
INC275671 | Workgroup | Incident | TR: BAIT - Y91 UL | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 14:33 | 2022-12-14 14:36 | 0 | 0 |  | Nathalie, ci-joint notre certificat UL, le Y9 n'est pas dessus mais le PBSN, l'EF6 et le CTX le sont. Best, Nicolas
INC275655 | Workgroup | Incident | WG: EIL - Anfrage-Nr.: 01/20221214 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 13:03 | 2022-12-14 14:08 | 0 | 0 |  | Hi Sascha, the 1083 doesn't exist anymore, it's now the D055. What you need is a D055-2L1C to fit with your gauge. Please find enclose the datasheet: Best, Nicolas
INC275638 | Workgroup | Incident | 10342377 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 11:36 | 2022-12-14 14:30 | 0 | 0 |  | Bonjour Sascha, I'm sorry but it's an old special execution that we don't make anymore. in addition, for the customer, it is dangerous to have non-standard connections, because no one will replace them. He should standardize its connection process. otherwise it was a simple mano on a Diaphragm seal so after that we will have solutions on a standard connection. Best, Nicolas
INC275637 | Workgroup | Incident | WG: Anfrage Preis und Lieferzeit | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 11:30 | 2022-12-14 14:34 | 0 | 0 |  | Hi Sascha, I ran through the german communications and found this code that fits your drawing. the code is open, you can go ahead, the only thing I don't know is the filling fluid (LRS15) but it seems to exist and worst case scenario we have the LRS1 which is also glycerin with a lot of certifications. 11191762 MEX5-D31.K24/1020_LRS15_0106 Best, Nicolas
INC275615 | Workgroup | Incident | a very old gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 09:24 | 2022-12-16 12:02 | 1 | 0 |  | Dear Zhang Wei, I'm very sorry but this is too old, we changed system twice since 1998. I'm not able to find out what this device was. If they have more informations, with drawings for example, please forward it to me. Best, Nicolas
INC275612 | Workgroup | Incident | FW: TSS/TSF 63mm EN13190 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 09:12 | 2022-12-16 16:47 | 1 | 0 |  | Hello Zhang Wei, so I have a more detailled answer. this comes from old internal specification and it respect the norm as the norms says that if we want to set more restrictive usage reccomendations, it's possible and should be specified with the triangles. I also discussed with R&D to know if they want to keep this fr the future. Best, Nicolas
INC275586 | Workgroup | Incident | product inquiry of DMO100100E59AG600 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Rongwen [xro] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 06:54 | 2022-12-14 14:01 | 1 | 0 |  | Dear Mr, I'm really sorry but I don't find anything even in old documentation about a DMO100. do you have any additionnal infos or pictures ? Otherwise you'll need to redefine the product with the customer. Best, Nicolas
INC275585 | Workgroup | Incident | FW: TSS/TSF 63mm EN13190 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-14 04:42 | 2022-12-14 09:03 | 0 | 0 |  | Hi Zhang Wei, I answered you yesterday, you're right it seems to be a mistake on our end, we're sorry for this. Best, Nicolas
INC275565 | Workgroup | Incident | MEX5 with special adaptopr | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-13 16:24 | 2022-12-13 16:54 | 0 | 0 |  | Hi Sascha, it's not existing for now but if it's just an adaptor (without diaphragm) then it could be codified as an ARIN (specific). we need the qty to define a price and codify it but it's possible with the drawings you sent I think. Best, Nicolas
INC275438 | Workgroup | Incident | TR: Y91-3.2.3.H25.R/DS-M3A16-G/MOUNT-GLY ryan dorsey & ohlheiser | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-12 17:12 | 2022-12-13 09:20 | 0 | 0 |  | Nathalie, c'est un montage fait chez ESP d'un 11235279 (Y91-3.2.3.H25.R/9750) sur un de leurs séparateurs clamp: [Image] Best, Nicolas
INC275425 | Workgroup | Incident | BarFlex | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-12 16:06 | 2022-12-12 16:21 | 0 | 0 |  | Bonjour Cyrille, Voici le certificat CE du Barflex: Best, Nicolas
INC275418 | Workgroup | Incident | Certificats de la commande 102896412 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-12 15:21 | 2022-12-12 16:16 | 0 | 0 |  | Cyrille, Désolé, mais aucun document n'était demandé sur cette commande. il n'y a pas de certificat rattaché. Best, Nicolas
INC275353 | Workgroup | Incident | Glycerin | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-12 10:12 | 2022-12-13 11:46 | 1 | 0 |  | Probably Answered, no feedback. Best, Nicolas
INC275313 | Workgroup | Incident | moisture of MZ7 on store and on site | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-12 04:51 | 2022-12-12 15:56 | 1 | 0 |  | Dear Zhang Wei, moisture can appear in anything containing air with temperature changes. when the conditions are causing this, we advise filled gauges or we have the weep hole option, see slide below: [Image] Best, Nicolas
INC275297 | Workgroup | Incident | thermomètres à contacts TSFE | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-09 16:54 | 2022-12-12 15:59 | 1 | 0 |  | Bonjour Cyrille, ci-joint dans le ticket la documentation (un peu vieille mais complète) sur les contacts: Best, Nicolas
INC275201 | Workgroup | Incident | 3d drawing of MEX5-D30.D20.DANC01.2L2A.01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-09 04:51 | 2022-12-12 16:02 | 1 | 0 |  | Hi Xu Qiang, here is the step file: Best, Nicolas
INC275171 | Workgroup | Incident | PDR | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-08 15:42 | 2022-12-15 12:10 | 1 | 1 |  | Veronique, voici la notice complete des afficheurs DX avec les details sur les presses etoupes, connections et cablages. Best, Nicolas
INC275139 | Workgroup | Incident | compatibilité séparateur // Mano avec séparateur - Process 300°C | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-08 15:09 | 2022-12-12 17:17 | 1 | 0 |  | Code ok, pas de soucis. Best, Nicolas
INC274979 | Workgroup | Incident | Ref 11046581 bloquée IA - MEX 2 B21 H1C - 20492047 | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-07 09:30 | 2022-12-08 08:52 | 1 | 0 |  | Adrien, tu peux y aller, je ne vois pas de point bloquant, demande à la boite Material data de l'ouvrir. Best Nicolas
INC274948 | Workgroup | Incident | MEX3 With Range 0 to 9000 psi (H39) | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-07 05:12 | 2022-12-12 16:59 | 1 | 1 |  | Hi Bee Tin, We agree, you can go ahead with no creation cost as it's in our configurator. Best, Nicolas
INC274941 | Workgroup | Incident | DEMANDE DE PRIX URGENTE HONEYWELL | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-06 17:24 | 2022-12-08 11:21 | 1 | 0 |  | Veronique, c'est un projet pour proposer des séparateurs à Honeywell, il va falloir plusieurs étapes et des tests pour valider les montages, ça peut être très intéressant mais je ne peux pas gérer le projet pour toi. je te conseille de faire une revue avec Jérôme et surement Christophe Charpentier qui ont géré ça pour des transmetteurs Yoko et siemens. la première étape est que tu définisse les produits à partir des datasheets (même grossièrement car il va te manquer des choses) ensuite il faudra faire des pièces de tests mais je suppose qu'il y a de l'administratif entre les deux. Best, Nicolas
INC274859 | Workgroup | Incident | FW: TSS/TSF 63mm EN13190 | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-06 09:03 | 2022-12-13 16:23 | 0 | 1 |  | Dear Zhang Wei, you're right, there is an error on the limit and it's should be as you said according to the norm. If there is a need for a return, please follow thee procedure with Nathalie. Best, Nicolas
INC274804 | Workgroup | Incident | BAIT - Karl Mayer - Y92 | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-05 14:57 | 2023-01-03 17:05 | 0 | 1 |  | Hi Francesco, here are the IC prices: M12 --> Y9239FB16R/2164 + 0410 ICO = 148.32 DIN 43650 --> Y9239FB16R + 0410 ICO = 142.10 Best, Nicolas
INC274796 | Workgroup | Incident | TR: PJ094777 - DEMANDE DE PRIX ET DELAI | Archive > BHFR Technical Support | Middle |  |  | Closed | Roederer, Adrien [aroe] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-05 14:03 | 2022-12-05 15:49 | 0 | 0 |  | Bonjour Adrien, ci-dessous le produit qui devrait te correspondre, je pense que tu demandais une ancienne norme: Order code: D820-10L.42DE.B6000 D820 PP old Code / SAP code: EURO Mounting and capillary type: Direct mounting 1 0.00 Capillary in m: No capillary 0 0.00 Instrument connection: G 1/2 Female L 10.80 Standard Flange: EN 1092-1 4 0.00 Flange material: Stainless steel (1.4404) 2 0.00 PN/ DN: Flange Rating PN_16 DN25 DE 153.90 Flange face : Raised face RF (Type B1) EN 1092-1 B 0.00 Diaphragm material : Hastelloy C 276 (HC) DN25 6 21.50 Coating : Without 0 0.00 Oil Filling : no oil filling 00 0.00 OPTIONS no selection Purchasing-price in EURO 186.20 Best, Nicolas
INC274682 | Workgroup | Incident | FW: Kindly confirm the IC price old item ref 11128510 | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-02 15:00 | 2022-12-02 16:57 | 0 | 0 |  | Hi Anil, nothing is existing yet but you can request the creation of the code 11128510 + 0592+9010 Y91-3.2.3.B76.R/2151_1545_9007_1310_0592_9010 Best, Nicolas
INC274581 | Workgroup | Incident | WG: AW: AW: Enquiry - Nufarm - Q17112 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-01 21:09 | 2022-12-02 11:58 | 0 | 0 |  | Hi Carlo, you can go ahead with the PTFE line (code 2) with Tantalum diaphragm (it can be done even if we don't offer it in the datasheet). for the reason the rest is not possible, I'm investigating. Best, Nicolas
INC274471 | Workgroup | Incident | drawing of ARIN210H. | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-01 09:09 | 2022-12-01 12:06 | 0 | 0 |  | Hi Xu Qiang, please find hereafter the drawing, I underlined the values for your particular connections: [Image] Best, Nicolas
INC274467 | Workgroup | Incident | BAIT - WaterSystem - MEX/TBI | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-12-01 08:51 | 2022-12-01 11:45 | 0 | 0 |  | Hi Francesco, Please find enclosed (in the ticket) the list of certificates, I let you chack for the pricing and SAP things with your usual contact (Nathalie or Céline). Best, Nicolas
INC274368 | Workgroup | Incident | BAIT - Pietro Fiorentini - Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-30 09:36 | 2022-11-30 11:21 | 0 | 0 |  | Hi Francesco, it's not a faulty signal, seems more like a slight 0 drift, but the guy needs to be sure he's exactly at 0. Also, 0.05mbar drift is 0.3% drift so it's within the tolerance. Best, Nicolas
INC274351 | Workgroup | Incident | TR: Request for prcie and model check | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-30 08:57 | 2022-12-19 17:43 | 0 | 1 |  | Dear JK, I close the request for now, please tell me if you need further help on this when you hav news from the customer. Best, Nicolas
INC274249 | Workgroup | Incident | RP2E | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-29 11:27 | 2022-11-29 16:36 | 1 | 0 |  | Hi Jurgen, it's some RP2E and the range of the first one is 10bar (150psi) but the rest of the information are missing and I can't find sure matches in the SAP code. it needs to be define by the customer, missing elements are: type of switch process connection overpressure dead band or reset point (if not specified set at the min dead band). Best, Nicolas
INC274198 | Workgroup | Incident | drawing of ARIN2302 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-29 03:51 | 2022-12-01 11:34 | 0 | 0 |  | Transfered to another request.
INC274193 | Workgroup | Incident | TR: PI - Accessoires AORP // INC274117 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-28 16:42 | 2022-11-29 17:26 | 0 | 0 |  | Hervé, voici les conditions pour les montées et descentes, et le reset point est environs 25% sous le setp. [Image] Best, Nicolas
INC274165 | Workgroup | Incident | Filling Fluid | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-28 13:48 | 2022-11-28 14:42 | 0 | 0 |  | Hi Jurgen, the LRS9 is not for gauge filling as it's not clear. but if it is for high temperature dampening, we can offer the dashpot. it is a dampening system on the movement directly, and the temperature limits will be the same as for a dry gauge. Best, Nicolas
INC274117 | Workgroup | Incident | TR: PI - Accessoires AORP | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-28 10:45 | 2022-11-28 14:56 | 0 | 0 |  | Salut Hervé, avant d'aller plus loin il nous faudrait les conditions de test avec les vitesses de montées en pressions et les caractéristiques des cycles. si les tests sont cohérents avec les normes et que nos appareils montrent des faiblesses, on pourra faire un retour client pour expertise. Si besoin d'un teams j'essayerais d'être dispo. Best, Nicolas
INC274116 | Workgroup | Incident | TR: PRIX IC SEPARATEUR - 20490528 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-28 10:42 | 2022-11-28 14:37 | 0 | 0 |  | Résolu directement.
INC274096 | Workgroup | Incident | pression de surcharge PBSN 0...20 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-28 09:48 | 2022-11-28 14:53 | 1 | 0 |  | Bonjour Jacques, la cellule utilisée est la même que pour le 0...25bar (B26) donc la surpression admissible est de 50bar. Best, Nicolas
INC274070 | Workgroup | Incident | 3D drawing of Y9239GBxxR | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-27 08:09 | 2022-11-28 14:46 | 1 | 0 |  | Hi Xu Quiang, I don't have the one you want, but I have the sam with M12 electrical connection instead of DIN. so if they want it for the process connection, it will help them (G connection). Best, Nicolas
INC273941 | Workgroup | Incident | TR: Certificate for Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-24 15:57 | 2022-11-24 17:22 | 0 | 0 |  | Nathalie, je crois qu'on n'a plus rien d'applicable, après avec un joint EPDM, les matières en contacts devraient satisfaire aux exigences usuelles européennes mais je ne connais pas cette norme polonaise en particulier. si il y a un marché on pourrait revoir la question mais surement sur une compatibilité européenne, pas spécifique à la Pologne. nos matières en contact sur un Y913 sont de la céramique, du 316L et le joint au choix ( EPDM usuel pour l'eau potable). Best, Nicolas
INC273901 | Workgroup | Incident | TR: Item inquiry : M306 59-09658 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-24 11:45 | 2022-11-24 16:21 | 0 | 0 |  | Nathalie, c'est un produit Indien, il faut qu'ils voient directement avec eux. Best, Nicolas
INC273785 | Workgroup | Incident | Customer Request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-23 11:57 | 2022-11-23 12:14 | 0 | 0 |  | Yes, it's this code, you can go ahead: 10270530. it also seems they require a calibration cert, don't forget to add it. Best, Nicolas
INC273744 | Workgroup | Incident | Customer request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-23 09:33 | 2022-11-23 09:51 | 0 | 0 |  | Hi Sascha, you'll need some details (you should find it in the project technical definition), but it looks like a monoflange 3 ways manifold to me. something similar to below: [Image] still te sleeve material and the sizes of process connection (flanged) and threaded connectin to the instrument are unclear. Best, Nicolas
INC273735 | Workgroup | Incident | Customer request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-23 09:00 | 2022-11-23 09:27 | 0 | 0 |  | Hi Sascha, difficult to say if it's like a DP100 or a gauge on seal but as I don't see filling port, I would say you're right. but the problem is we are still facing the issue that over 9bar the diaphram is not withstanding the pressure. solutions are being tried in the lab but for now, please don't offer it. You can offer a solution with MEX5 + DAxx for now. Best, Nicolas
INC273433 | Workgroup | Incident | TR: SPR for APS / Tasmania Water / Bryn Estyn WTP Project | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-18 15:57 | 2022-11-21 09:40 | 0 | 0 |  | Nathalie, le 1er code n'est pas possible, le mini de montage est 10bar, pour le deuxieme pas de soucis. MEX5-461.D20.D821- 10N.2212.A20.01/0191_0751_0760_9941 NOK (PN/DN min 10bar) MEX5-461.D20.DT2- 2N20.0.260.01/0191_ 0751_0760_9941 OK Best, Nicolas
INC273434 | Workgroup | Incident | TR: OP Dinstec 04102022 20482135 ¦ DROG + DT2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-18 15:57 | 2022-11-21 10:03 | 0 | 0 |  | Salut Céline, J'ai trouvé le code et il est récent et en RL, tu peux y aller 11218011 Best, Nicolas
INC273329 | Workgroup | Incident | Old Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-18 09:15 | 2022-11-18 09:25 | 0 | 0 |  | Hi Daniel, yes you're right, we still have contact differentials but no 4...20mA output on differential. And the mCE is french for meter of water column. Best, Nicolas
INC273304 | Workgroup | Incident | Ex-certificate For RP2 E / Kong Rong | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-18 06:51 | 2022-11-18 09:30 | 0 | 0 |  | Hello Bee Tin, please find the certificates enclosed in the ticket: Best, Nicolas
INC273303 | Workgroup | Incident | TAX6S-115-400-6 /GTS | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-18 05:15 | 2022-11-18 09:36 | 0 | 0 |  | Hi Bee Tin, it was the former TBI, but I misst the information about the connection (I offered standard 1/2NPT) and the diameter of the stem (I think the 6 at the end of the code was for 6mm but not sure). Order code: TBI130-264.147.51T TBI old Code / SAP code: 10299876 Nominal size: DN 130 130 Case AISI 304 (1.4301), stainless steel: AISI 304 (1.4301), stainless steel 2 Process connection: With sliding connection Sliding screw connection 1/2"-14 NPT male 6 Immersion tube outlet: Center back, every angle 4 Indicator: EN 13190 (Standard) 1 Immersion tube diameter: 6 mm AISI 316Ti (1.4571), stainless steel 4 Immersion tube length L1: 400 mm 7 Temperature range: Single scale inC -30 ... 70C 51T Best, Nicolas
INC273265 | Workgroup | Incident | WG: Anfrage Manometer | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-17 14:09 | 2022-11-17 14:49 | 0 | 0 |  | Dear Colleague, this was done by Haenni a long time ago and we don't have this exact product anymore. but we can offer you the same functionnalities. below is what I suggest, the electrical connection is different and on the side and I suggest to use directly a /2 with a restriction screw instead of the adaptor they have (to be validated by the customer as I don't have details of a possible functionnality of it). Order code: MR5-4D30.B38/0771 MR old Code / SAP code: Nominal size: 100 mm 5 Functions of order: 1st contact breaks / 2nd contact makes (5021) 4 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection D Process connection: G1/2" 3 Oil filling: No oil filling 0 Pressure unit: Bar (EN standard) B Pressure range: 0 ... 400 bar 38 OPTIONS Restriction Screw: Stainless steel restriction screw d = 0.5 mm 0771 Best, Nicolas
INC273173 | Workgroup | Incident | INFOS PRODUITS | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-17 07:15 | 2022-11-17 12:20 | 0 | 0 |  | Bonjour Véronique, je vais pouvoir apporter des éléments de réponse, mais je sais qu'il y a une qualification qui a été faite avec EDF UTO et on a des fiches SEPTEN pour les produits donc ils devraient demander la qualif directement à UTO. Je suis un ingénieur chargé daffaires de la société Assystem qui travaille pour EDF UTO. Suite à lapprovisionnement de matériels BAUMER, nous aurions besoin dinformations supplémentaires pour installer les gammes suivantes : MDX / MX (MDX plus existant, voir MX) MEX MIX MJ ZDPHN ZDPN ZDPWN ZPHN ZPN ZT 3 (parties en contact Laiton) ZT 4 (parties en contact Laiton) ZTN 4 Les questions que nous avons sont les suivantes : - Les matériels de ces gammes passent-ils des essais hydrostatiques de fin de fabrication. Si oui pourriez-vous nous les transmettre ? oui, ils sont tous mis sous Pmax lors de la calibration (c'est en gamme et assuré par notre manuel qualité ISO) - Ces matériels sont en matière Acier Inox. Pouvez-vous nous confirmer spécifiquement que leur teneur en cobalt est bien inférieure à 0,2% et confirmer labsence dargent ou dantimoine. Vous pouvez demander un certificat matière 3.1 a la commande, sinon les matières exactes (normalisées) des éléments sont mentionnées dans les datasheets (316L pour les parties en contacts de votre liste ci-dessus (ou laiton)). - Vis-à-vis des procédures relatives au nettoyage et au contrôle propreté du matériel, en avez-vous et si oui pouvez-vous me les détailler ? Nous sommes en particulier intéressés par des éventuelles vérifications dabsence de souillure et de limitation de contamination ferritique superficielle à 1%. Sur les manomètres (codes commençant par M de la liste), nous pouvons inclure un nettoyage (en option). concernant la contamination ferritique, nos pièces sont usinées et nos soudures font l'objet de QMOS/DMOS connus de EDF Best, Nicolas
INC273172 | Workgroup | Incident | the meaning of 0201 option code | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-17 06:39 | 2022-11-17 09:18 | 0 | 0 |  | Hi Zhang Wei, the option code 0201 was only used in the past and was a "dummy" option number just to say there was options (we had limited amount of digit in SAP codes). that's the reason you find it in the column "old code". now all our codes have the second column type of code with details. on the product you tell me about, the code 0201 means option SETP. Best, Nicolas
INC273160 | Workgroup | Incident | connection BSP 1/8" | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-16 16:39 | 2022-11-17 11:47 | 0 | 0 |  | Hi Jurgen, you can confirm with the customer this is what he needs, if ok, you can ask for the price of the code MEX3-6X0.B29/0751_0078 with a connection according to 57206-11 to Nathalie or Celine. [Image] Best, Nicolas
INC273119 | Workgroup | Incident | Projet AIR LIQUIDE avec MEP5 // Corée | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-16 12:48 | 2022-11-16 17:42 | 1 | 0 |  | Salut Hervé, alors je ne suis pas expert en règlementation Coréenne, mais je vais poser la question a Baumer Corée. Dans la plupart des pays, les normes européennes ou américaines sont en générale reconnues comme acceptables. je traduirais ta question demain à BADK Best, Nicolas
INC273086 | Workgroup | Incident | TR: HPS2200730 - 3 pcs of Pressure Transmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-16 10:27 | 2022-11-17 14:08 | 1 | 0 |  | Nathalie, c'est Ok, mais le prix est à actualiser avec Antoine. Best, Nicolas
INC272979 | Workgroup | Incident | Special technical request | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-15 14:21 | 2022-11-15 15:20 | 0 | 0 |  | Hi Sascha, no problem fo doing scales in m3, but the end of the scale needs to be 400mbar. for example the first one will be a 0...9.3m3 with 7.3m3=0.315mbar so 9.3m3=400mbar. and check with Nathalie for the cost of dial creation. Best, Nicolas
INC272959 | Workgroup | Incident | TR: Leitenberger article | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-15 12:18 | 2022-11-15 15:26 | 0 | 0 |  | Salut Nathalie, pour moi c'est ok en terme de montage et de pression, et de toute façon c'est validé par Jorg. tu peux faire la création. Best, Nicolas
INC272953 | Workgroup | Incident | TR: HPS2200730 - 3 pcs of Pressure Transmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-15 11:51 | 2022-11-15 11:53 | 0 | 0 |  | Bonjour Nathalie, on ne fera pas d'étendue spécifique pour 3 pièces, ce sera du B81 (-1...15bar) best, Nicolas
INC272842 | Workgroup | Incident | TR: SDI 439/131.011 0...28'868 L SPEC | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-14 16:24 | 2022-11-18 15:56 | 0 | 1 |  | Clos pour absence de réponse client, à réouvrir au besoin. Best, Nicolas
INC272759 | Workgroup | Incident | MZ7-460.B15L/0751_0678 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-14 12:30 | 2022-11-14 14:31 | 0 | 0 |  | Hi Daniel, with this big difference between the static pressure and the measuring range, we need to use a demultiplied movement. and with this movement, we can't afford the adjustable pointers (0678 or 0679). the only solution would be to have less of a static and a bigger range, if that is possible, please tell me and we'll fin a solution to have maybe the 0678. otherwise I'ms sorry but I don't have a solution with an MZ7, we could check on the MFT7 which is built differently. The ticket is closed, but if you need me to investigate on one or the other solution, recontact me. Best, Nicolas
INC272521 | Workgroup | Incident | RE: AKP_INQ_20221110_0954 | Archive > BHFR Technical Support | Middle |  |  | Closed | Perrichon, Philippe [pper] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-10 11:15 | 2022-11-10 11:30 | 0 | 0 |  | Answered directly by Philippe.
INC272512 | Workgroup | Incident | WG: AKP_INQ_20221110_0954 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-10 10:42 | 2022-11-10 11:30 | 0 | 0 |  | Double
INC272494 | Workgroup | Incident | OT53185876 Ot53185877 821PG814 / 807 demande de prix mano capillaire | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-10 09:09 | 2022-11-29 10:19 | 0 | 1 |  | Bonjour Véronique, désolé que ça ait pris autant de temps, mais voici les brides de remplacement: [Hier 17:59] Pichon Sylvain D438-MD0.EN2.000000.000.000 + 10358048 MEX5-A61.B20/0816_0678_0751_0078 + LRS1/9? [Hier 17:59] Pichon Sylvain D438-MD0.3N2.000000.000.000 + 10280059 MEX5-A60.B19/0764 + LRS4 les capillaires sont en inox protection inox avec protection PVC blanc ATEX (le 2m passe en 3m standard). Best, Nicolas
INC272452 | Workgroup | Incident | Délai de livraison sur les CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-09 15:39 | 2022-11-09 16:04 | 0 | 0 |  | Vu au téléphone, 12/12 pour l'instant (si possible par 25). si besoin de pousser plus le delai me rappeler. Best, Nicolas
INC272423 | Workgroup | Incident | AW: MAH3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-09 13:06 | 2022-11-09 15:54 | 0 | 0 |  | Hi Carlo, we checked the BOM and it's Bronze everywhere. Best, Nicolas
INC272400 | Workgroup | Incident | need Article Code / Price / delivery time | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-09 11:00 | 2022-11-09 11:21 | 0 | 0 |  | Hi Sascha, request cancelled. Best, Nicolas
INC272336 | Workgroup | Incident | Enquiry For RP2E /Kong Rong, | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-09 04:09 | 2022-11-09 09:43 | 0 | 0 |  | Hi Bee Tin, here are the codes and setp (the codes might need to be created again with the setp included): 11078399 RP2-ER.646/SETP 70psi falling 11078398 RP2-ER.644/SETP 30psi falling 11078366 RP2-ER.665/9704/SETP 350psi rising Best, Nicolas
INC272327 | Workgroup | Incident | request article | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-08 17:30 | 2022-11-09 09:09 | 0 | 0 |  | Hi Sascha, I would say it's more similar to this code: 10351137 but the code needs to be created anyway. Best, Nicolas
INC272264 | Workgroup | Incident | ISO 17025 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-08 12:33 | 2022-11-09 09:20 | 1 | 0 |  | Bonjour Cyrille, nous ne sommes pas ISO17025, cette norme est reservée aux laboratoires. nous sommes par contre ISO9001 et ISO9100 pour notre fabrication et notre calibration (ce qui inclue les certificats d'etalonnages). Best, Nicolas
INC272144 | Workgroup | Incident | Honeywell Projet NOV | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-07 14:30 | 2022-11-07 15:37 | 0 | 0 |  | Salut Hervé, voici la nomenclature pour un exemple en 6m: [Image] Best, Nicolas
INC272062 | Workgroup | Incident | procedures of MX calibration | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-07 02:21 | 2022-11-07 15:22 | 1 | 0 |  | Hello Zhang Wei, We usually don't give our internal procedure, we are ensure the accuracy stated on the datasheet, and if a verification is requested by the customer it needs to be via a calibration certificate [Image] the MX and MZ have the exact same internal parts, only the position of the pressure connections are different so there is no accuracy differences possible. Best, Nicolas
INC271929 | Workgroup | Incident | D030 Compatible avec remplissage LRS4 ? | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-03 18:12 | 2022-11-04 11:30 | 0 | 0 |  | Salut Hervé, le LRS4 avec un D030 n'est pas possible à cause de la densité bien trop forte pour la membrane. on est par contre en train de rechercher la vrai raison pourquoi seuls les LRS2 et 7 sont autorisés (si j'ai une réponse je te tiens au courant). Best, Nicolas
INC271858 | Workgroup | Incident | FW: Thermowell Requirement | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-03 12:33 | 2022-11-14 10:04 | 0 | 0 |  | Hi Rudi, I close the ticket for now, tell me when you have some news. Best, Nicolas
INC271770 | Workgroup | Incident | FW: Throttle 0769 in MEM3 | Archive > BHFR Technical Support | Middle |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-02 15:57 | 2022-11-02 16:38 | 0 | 0 |  | Hello Kees, with the restrictor screw being made of monel and the few request we have for it, we needed to standardise to only one option for restrictor screws. the 0771 (0.5mm monel) is available. Best, Nicolas
INC271759 | Workgroup | Incident | LRS9 Fill Fluid (SAP 11080891) | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-02 15:12 | 2022-11-02 15:43 | 0 | 0 |  | Hi Daniel, no problem for ordering it then, but you need to go through Nathalie for pricing. Best, Nicolas
INC271709 | Workgroup | Incident | Pricing | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-02 10:51 | 2022-11-02 11:09 | 0 | 0 |  | Hi Sascha, I don't do pricing, you have to see with sales administration for this, I transfer your email. Best, Nicolas
INC271691 | Workgroup | Incident | TR: Bestellung Manometer | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-11-02 09:18 | 2022-11-02 10:39 | 0 | 0 |  | Salut Nathalie, Pour les manos avec spécificités, on préfère en générale proposer du MEX3. mais si il y a des quantités intéressantes, il serait peut être possible de le faire sur un MMD3 mais il faut une SPR. ce sera un processus avec approbation qui va prendre un peu de temps. Best, Nicolas
INC271554 | Workgroup | Incident | RFQ MJX5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-31 16:39 | 2022-11-02 10:12 | 1 | 0 |  | Answered via email, sorry Best, Nicolas
INC271397 | Workgroup | Incident | Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-28 15:15 | 2022-10-28 15:31 | 0 | 0 |  | En me le demandant ;) en attendant qu'il soit sur le site. je le mets a dispo dans les documents du ticket. Best, Nicolas
INC271313 | Workgroup | Incident | MEP5-463.B31.AORP0385/SETP.ARIN2606.DT1-2N60.0.630.09/0078_0678_0693_0763_9941/9030 / Neste | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-28 05:57 | 2022-10-28 09:20 | 0 | 0 |  | Hi Bee Tin, the mounting you asked for is OK, but there is no major benefit from the AORP as the DT1 will anyway only resist until 160 and the gauge up to 130, but it is possible. however, if you need more overpressure protection (and only if you do need it ), I can suggest this alternative: MEP5-463.B31.AORP0385/SETP (100bar).ARIN2606.DT8-2N60.0.630.09/0078_0678_0693_0763_9941/9030 then it's protected up to 400bar. you'll need to give the setp and the marking for the 9941 at the order. Best, Nicolas
INC271246 | Workgroup | Incident | RP2N | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-27 10:57 | 2022-10-27 12:10 | 0 | 0 |  | Hi Jurgen, the code: RP2NL646-0201 = RP2-NL.646/SETP 10296971 but the code is created without the setpoint defined, so it needs to be created according to the current code standards. concerning the Kxxx numbers, i doesn't come from our side, it might be customer numbers and if you need to have it somewhere on the product an option needs to be added (9941 to have it on a stanless steel tag plate for example). Best, Nicolas
INC271230 | Workgroup | Incident | BAIT - Hyter - MTTF Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-27 10:19 | 2022-10-27 12:12 | 0 | 0 |  | Hi Francesco, please find in the ticket the MTTF you asked for: Best, Nicolas
INC271196 | Workgroup | Incident | TR: Inquiry_k 202210241713 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-27 10:18 | 2022-10-27 12:14 | 0 | 0 |  | Nathalie, ils ne peuvent pas démonter nos produits comme ça, si il y a un problème, ils doivent le renvoyer en SAV. on ne fournira pas le joint seul pour un remplacement. Best, Nicolas
INC270832 | Workgroup | Incident | delivery time | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-24 17:06 | 2022-10-24 17:35 | 0 | 0 |  | Nathalie, can you please check the lead time for this product ? MMD3-D10.B33 the code is not existing. Best, Nicolas
INC270827 | Workgroup | Incident | open questions WEP5 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-24 16:39 | 2022-10-25 17:45 | 0 | 0 |  | Most probably the answers will be sufficient, tell me via teams if something is missing. I close the ticket, Best, Nicolas
INC270786 | Workgroup | Incident | Certificat CE Barflex EA4 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-24 12:00 | 2022-10-24 12:10 | 0 | 0 |  | Bonjour Jacques, ci-joint le certificat (dans le ticket). Best, Nicolas
INC270717 | Workgroup | Incident | WG: inquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-21 16:21 | 2022-10-24 15:59 | 0 | 0 |  | Sorry Carlo, As discussed, we have no solution to offer in differential with 4..20mA output. Best, Nicolas
INC270702 | Workgroup | Incident | TR: Delivery delay on 11230049, OC# 1976733. Urgent pls expedite | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-21 14:57 | 2022-10-21 16:29 | 0 | 0 |  | Nathalie, comme discuté, la solution de changement de membrane est trop longue et aleatoire (le fournisseur n'est pas sur de pouvoir faire) la meilleure solution pour proposer un remplacement equivalent en terme de fonction est un D820 full HC276. Sylvain est en train de chiffrer, il revient vers toi. tu peux informer le client en attendant qu'on essaye cette solution. Best, Nicolas
INC270679 | Workgroup | Incident | BOURDON SEDEME (Y943 025 B18R) & BOURDON SEDEME (Y943 025 B20R) | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-21 13:00 | 2022-10-21 13:56 | 0 | 0 |  | Hi Bee Tin, I'm sorry, the Y94 was a submersible pressure sensor we did for a very short time. it was replaced by the PSSN and PSMN from Baumer but it's phased out and there was no replacement. Best, Nicolas
INC270654 | Workgroup | Incident | Sealing Gasket 10305228 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-21 10:39 | 2022-10-21 11:21 | 0 | 0 |  | Sorry, answer send via email instead of closing
INC270646 | Workgroup | Incident | Demande d'offre pour un 2 manomètres différentiel | Archive > BHFR Technical Support | Middle |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-21 10:21 | 2022-10-21 11:01 | 0 | 0 |  | Hi Viola, those ranges are available at different conditions: the range width should be respected (with a 160mbar, you can do a -100...60mbar) it'll be codified as X + 09 two digits for the basic range (160mbar in our example) there will be a cost for creation in case of not existing Best, Nicolas
INC270589 | Workgroup | Incident | BAIT ¦ PFMN + DRD flange | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-20 15:21 | 2022-10-21 16:53 | 0 | 0 |  | Dear Erald, we checked with R&D and management and we will not be able to offer something on it. last request we thought it was possible but in fact the investments to be able to mount and test it would be too big. I'm sorry for the bad news, Best, Nicolas
INC270542 | Workgroup | Incident | Photo of MEX5, 0...10000 Psi / Kumpanart | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-20 11:03 | 2022-10-20 14:01 | 0 | 0 |  | Hi Bee Tin, here is the dial for a MEX5 0...10000psi: [Image] Best, Nicolas
INC270532 | Workgroup | Incident | Angebot 20484624 / Anfrage 1325 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-20 10:36 | 2022-10-20 11:18 | 0 | 0 |  | Dear Colleague, as far as I can see, we can offer a solution. MEX5 + AORP + ARIN2606 + DT1 (HC276) the Inconal is not available but HC is a standard replacement commonly used the AORP will block the pressure at max scale and the system can then withstand 160bar. Best, Nicolas
INC270522 | Workgroup | Incident | TR: URGENT - Confirmation technique - Membrane DP100 // SANOFI | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-20 09:54 | 2022-10-20 10:22 | 0 | 0 |  | Salut Hervé, on achète les membranes à ce Ra, donc on ne fait pas d'electropolissage. il y a un polissage méca de prévu si le Ra est trop haut sur le cordon de soudure. Best, Nicolas
INC270517 | Workgroup | Incident | replacement of MAH3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-20 09:36 | 2022-10-20 11:41 | 0 | 0 |  | Hello Xu Qiang, the MAH3 is phased out completely, and we didn't plan to replace it. please find enclosed the datasheet, but it's very special and will not be easy to find a replacement. Best, Nicolas
INC270440 | Workgroup | Incident | WG: RFQ 102061 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-19 14:21 | 2022-10-19 14:58 | 0 | 0 |  | Hi Carlo, your definition is good, the only thing I would change is the type of contact to a K (double gold contact) that is specifically design for low current and will allow the setp and reset point they need. as the mounting is not defined from them, what you offered seems good. and for the contacts where you said you don't know, you can say it's ok, we provide here double NC contacts (can be wired on NC or NO). Best, Nicolas
INC270394 | Workgroup | Incident | ref10301891 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-19 09:51 | 2022-10-19 10:47 | 0 | 0 |  | Celine, pas de soucis, tu peux y aller. Best, Nicolas
INC270269 | Workgroup | Incident | TR: Urgent. Over Range Protector AORP | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-18 08:48 | 2022-10-18 09:47 | 0 | 0 |  | Nathalie, oui c'est bon, tu peux lui préciser que c'est du Femelle/Femelle Best, Nicolas
INC270080 | Workgroup | Incident | AOPR Calibration certificate - ExxonMobil | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-14 14:24 | 2022-10-14 16:12 | 0 | 0 |  | Hi Rudi, if you order a 2.2 certificate for the AORP or the mounting, you'll have a certificate with the SETP on it. if it's not sufficient, come back to me. Best, Nicolas
INC270035 | Workgroup | Incident | TR: Solicitação de Cotação - 7003946687 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-14 10:18 | 2022-10-14 10:45 | 0 | 0 |  | Hi Vinicius, the MPE is stopped and replaced by the new MPG6 (PBT case), the phenolic is causing too much problems as it's too stiff and breakable. the main reason for the use of phenolic case is the conformity to U94 V0 (fire retardant and our new MPG6 complies also. A note will be issue shortly and also new datasheets and flyers. we also made other new improvements on the new MPG6. Best, Nicolas
INC270022 | Workgroup | Incident | RE: Oferta manometry. | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-14 09:30 | 2022-10-14 10:50 | 0 | 0 |  | Nathalie, the limit on a small gauge on a DADL DN25 is -1...3bar. we could launch some samples if there is qty but we can't go below -1...1.5 and it's even not sure and would need to go through an SPR. Best, Nicolas
INC269966 | Workgroup | Incident | indly confirm for pressure switch - suitability for using in the offshore-marine environment (FPSO environment) | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-13 12:33 | 2022-10-13 16:55 | 0 | 0 |  | Hi Anil, the things they look for for this type of application will be IP and materials you can find on the datasheet [Image] I can add a few details: the case (black) is painted by cataphorese with epoxy paint of a thickness of 20+/-4microns the cover (blue) is polyester paint with a 0.06 to 0.08mm thickness Best, Nicolas
INC269929 | Workgroup | Incident | BAIT - Neptune Venice - BAUMER BORDON GAUGE LEVEL | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-13 10:24 | 2022-10-24 15:58 | 1 | 1 |  | Hi Francesco, I close it as I don't know when we will have an answer from the customer, but tell me via teams if you got anything back and we will reactivate it. Best, Nicolas
INC269921 | Workgroup | Incident | PED certificate - ExxonMobil - 0020473411 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-13 09:48 | 2022-10-13 17:54 | 1 | 0 |  | Hi Rudi, here is the official answer on this from 2019: Réponse faite en 2019 par Philippe, que je confirme. Les accessoires sous pression types (ASIP / AKPL / Siphons/ ARIN )sont exclus de la directive sous pression 2014/68/UE car selon le fiche CLAP X059, ils ont de faibles PS, volumes et surtout un DN< 25 donc tombent sous les dispositions de larticle 4 3. Ces accessoires sous pression nont pas à satisfaire aux exigences essentielles mais seulement aux règles de lart. Donc pas dattestation PED pour ces accessoires sous pressions. Best, Nicolas
INC269886 | Workgroup | Incident | WG: akp_inq_20221012_1540_SzuuSTER | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-12 17:18 | 2022-10-12 17:31 | 0 | 0 |  | Hi Carlo, my comment is that the gauges are standard (as far as I can see, and the connectors are accessories. if we can get the information of the supplier of the connection, it's just the supply of those accessories with our gauges. do you think you can have access to this information ? Best, Nicolas
INC269876 | Workgroup | Incident | Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-12 16:06 | 2022-10-12 17:27 | 0 | 0 |  | Hi Kerstin, what had you doubting ? the gauges are full standard, the seals exist (codes below) and the mounting are not near the limits, you can go ahead: MEX5-D30.F22 + D820-10L.2217.G2002 (10267254) MEX5-D30.F26 + D820-10L.2227.G2002 (10267263) MEX5-D30.F29 + D824-10L.2227.G7002 (11105272) Best, Nicolas
INC269848 | Workgroup | Incident | RP2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-12 14:30 | 2022-10-12 15:15 | 0 | 0 |  | Hi Carlo, the switch at 0.4bar falling is not an issue, but it means the max deadband will be 250mbar and so the reset point will be 0.650mbar rising. it means that when the pressure will drop below 0.4bar, the switch will close. and when the pressure is back above 0.65bar, the switch will reopen. to make something closer to what you ask for, I suggest: RP2EL851 setp: 0.45bar falling resetpoint: 1.4bar rising deadband: 0.95bar the specific connection on this will be crazy expensive, I suggest to use an ARIN if necessary. Best, Nicolas
INC269822 | Workgroup | Incident | ASIPD1L3 Maximum Pressure & Temperature Limits / Neste | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-12 11:51 | 2022-10-12 14:29 | 0 | 0 |  | Hi Bee Tin, Max pressure and température are on the datasheet, and the T max is even a digit in the code you choose. 400bar and 220C max for the one you choose. And no problem for the mounting. Best, Nicolas
INC269752 | Workgroup | Incident | capsule gauge 0....25mbar | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-11 17:45 | 2022-10-17 17:33 | 0 | 1 |  | Hello Sascha, we can do the code DK 63/211.111/0..25 MBAR (10320388) but not with G type connection (it would need to be checked via an SPR). Best, Nicolas
INC269751 | Workgroup | Incident | adaptateur 8126-936 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-11 17:24 | 2022-10-12 14:02 | 0 | 0 |  | Bonjour Cyrille, a priori la question est déjà répondue via Hervé Courdier, c'est un adaptateur DN51 PN30 du Danemark. Best, Nicolas
INC269511 | Workgroup | Incident | Request for MX35+Y913+1084 | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-07 17:33 | 2022-11-04 09:47 | 0 | 1 |  | Hi Jorg, we can finally close the ticket and here is the code and IC: 11259721 IC 454.22 Best, Nicolas
INC269416 | Workgroup | Incident | TR: BAIT ¦ ED522 alternative | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-07 10:09 | 2022-10-12 11:21 | 0 | 0 |  | Hello Erald, without feedback on this, I can't do anything more, I close the ticket. come back to me if you get a picture please. Best, Nicolas
INC269350 | Workgroup | Incident | Noord Natie (21107396) - Weight of MEX5-D61.B79.ABR42GEB0N/0751_0078_9030 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-06 16:09 | 2022-10-07 09:23 | 0 | 0 |  | Hi Rudi, you have the weight in the article code normatlly, see hereafter: [Image] Best, Nicolas
INC269202 | Workgroup | Incident | PG for Hydrogen | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-05 17:39 | 2022-10-12 11:20 | 0 | 1 |  | Hey Carlo, I suggest a quick meeting with Jorg on this subject, I can't validate it by myself. Best, Nicolas
INC268945 | Workgroup | Incident | DR080 X 0-60 BAR X ¼” NPT BOTTOM x 40 units | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-04 12:18 | 2022-10-04 15:15 | 0 | 0 |  | Hi Bee Tin, yes it is but we will need a complete codification, in the ticket I add the datasheet. And I supposed it's DRO80 and not DR080. Best, Nicolas
INC268920 | Workgroup | Incident | MCF7-P30.R60/0693_0078_0651_9941+SAMFSA-G4G4-A+ASIPD1L3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-04 11:00 | 2022-10-04 11:36 | 0 | 0 |  | Hi Bee Tin, in terms of mounting, connections matching etc..., it works. other solution would be with a standard AMFD and ARIN but I'm not sure it'll be much better in terms of price and it's one more accessory. the only thing is for the 2" pipe mounting as your whole mounting is fixed so make sure it fit's on the tube of the customer. Best, Nicolas
INC268851 | Workgroup | Incident | TR: Product confirmation/ PGS_ Y913 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-04 09:00 | 2022-10-04 09:53 | 0 | 0 |  | Bonjour Nathalie, nous n'avons plus les certifications FM/CSA (non renouvelées), donc les options 0839 et 9936 ne sont plus disponibles. ci-joint un courrier qu'on a fait pour proposer aux clients une alternative. le reste du code est OK. Best, Nicolas
INC268772 | Workgroup | Incident | TR: Zamówienie INL/ZMZ/2022/243 wg oferty nr DF22042001 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-10-03 14:33 | 2022-10-03 15:19 | 0 | 0 |  | Nathalie, they have the key inside the electrical connection on the side of the gauge. if they don't have it, they can check in the box and otherwise, we can ship a couple. Best, Nicolas
INC268582 | Workgroup | Incident | réglages sur les TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-29 11:54 | 2022-09-29 17:03 | 0 | 0 |  | Bonjour Cyrille, dans le TED, c'est un microcontrôleur donc une fois que les settings sont fait ils n'ont pas de dates d'expiration, même en cas de coupure de courant etc... la seule façon que les paramètres changent est une action directe de changement via les boutons en façade mais ton client peut mettre un code pour en limiter l'accès. Best, Nicolas
INC268545 | Workgroup | Incident | MX7-P60.B18K/0751_0765 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-29 08:09 | 2022-09-29 13:58 | 0 | 0 |  | Hello Bee Tin, please find hereafter the dial layout of the product you asked for: [Image] Best, Nicolas
INC268517 | Workgroup | Incident | need article code | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-28 16:12 | 2022-09-28 17:08 | 0 | 0 |  | Hi, I would suggest to offer: Y91333B24R + DT2-2L20.0.2L0.01 The DT2 = D211 and we don't mount anymore PBSN, but usually, the Y9 will make a good replacement. Best, Nicolas
INC268419 | Workgroup | Incident | étalonnage des capteurs de pression | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-27 17:48 | 2022-09-28 10:37 | 0 | 0 |  | Bonjour Cyrille, le Q005 c'est 5 points a la montée et 4 a la descente. Best, Nicolas.
INC268418 | Workgroup | Incident | the declaration of conformity -CE-ATEX for Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-27 17:42 | 2022-09-28 11:00 | 0 | 0 |  | Hi Xu Qiang, this is an interesting question, and here is the answer from our quality dept: norms evolutions are not immediately applied by the european commission, it can take up to several years. versions mentionned in our conformity declaration are still the ones applied in the official journal of the directive 2014/30/UE ( CEM). we revise it periodically and evolutions will be applied when necessary. Best, Nicolas
INC268415 | Workgroup | Incident | barflex | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-27 17:00 | 2022-09-28 10:26 | 0 | 0 |  | Hi Angel, you can go ahead, we closed it on the website, as the datasheets is partially outdated (IrDA etc.) but you can order this one: EA5-AF0.H21R/0587 (0587 = menu in english). Best, Nicolas.
INC268372 | Workgroup | Incident | SAP 11133571 - Status DO - Electrabel RFQ-585978 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-27 12:06 | 2022-09-27 14:05 | 0 | 0 |  | Hi Rudi, it seems like a test code with nothing inside, I created you a new one: MIX7-B30.B77 11259129 Best, Nicolas
INC268304 | Workgroup | Incident | low range -2.5/2.5mbar code 110 in DP switch model RDN-4A^.110/Q003/0407/SETP | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-27 06:30 | 2022-09-27 11:54 | 0 | 0 |  | Hi Anil, we can do this but the setp is to be confirmed and the Q003 is separated from the code: RDN-4AE.110/0407/SETP Best, Nicolas
INC268215 | Workgroup | Incident | TR: 2022-09-22 - Sun Chemical [45.016.22] - [Zapytanie techniczno-ofertowe] Manometry (EG22091608) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-26 12:18 | 2022-09-26 14:06 | 0 | 0 |  | Nathalie, voici les montages validés, j'ai enlevé les certificats qui ne feront pas partie du code et ajouté les LRS Pour le premier montage, on ne peut pas mettre de vis frein, mais la transmission via la membrane et le fluide de transmission ont aussi ce rôle. 1. MEX5-431.B22/0078_0751_0760 + D820-10L.42GG.B2001 15 pcs2. MFT5-403.B20N.32/0078 + 2*D820-BEL.42DGB2001 Probably possible but needs to be confirm by supplier of MFT as it'll not be a D820 but local seals supply. or MX7-431.B20H/0751_0078 + 2*D820-BEL.42DGB2001 Best, Nicolas
INC268186 | Workgroup | Incident | Enquiry For ETTN / Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-26 10:24 | 2022-09-26 10:53 | 0 | 0 |  | Dear Bee Tin, generally speaking, the datasheet is not on the website for 2 main reasons, the YTTN and the ETTN7 (modbus) are not available anymore. for the rest, you can follow the datasheet, but don't hesitate to ask if you're not sure or the code is not existing or not released. Best, Nicolas
INC268146 | Workgroup | Incident | Y91 application for steam up to 300 deg C | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-26 05:15 | 2022-09-26 10:56 | 0 | 0 |  | Hi Xu Qiang, yes, you have a table on the ASIP datasheet and for a schedule 40 and 300C, the max pressure is 100bar, so no problem. Best, Nicolas
INC268144 | Workgroup | Incident | FW: MCF+DT3, -70+70 inWC/-174+174 mbar / Kumpanart | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-26 03:36 | 2022-09-26 10:59 | 0 | 0 |  | Hi Bee Tin, yes, this is possible except it'll not be S3 (our MCF doesn't have a blow out back). and there will be a dial design to do and validate, but other wise the mounting is possible with those pressures. Best, Nicolas
INC268135 | Workgroup | Incident | Pbsn + separateur DT2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-23 16:30 | 2022-09-26 11:30 | 0 | 0 |  | Bonjour Jacques, c'est pas que ce n'est pas possible mais on a pas vraiment de données sur le ZPX1 et on ne monte pas de PBSN sur séparateur normalement. je conseille donc soit: Y91-3.3.3.B16.R/2164 + ASIPD + DT2-2L51.0.F30.01/9010 (probably 2% at 20C) or PBMN-24B16RA11403201000 + ASIPD + DT2-2L51.0.F30.01/0592/9010 (probably 1% at 20C) Best, Nicolas
INC268107 | Workgroup | Incident | FW: Customer request: product | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-23 13:18 | 2022-09-28 17:37 | 0 | 1 |  | Hi Jurgen, please find enclosed (in the ticket), the barflex instructions. the 4s press needs to be done at atmospheric pressure 1min at least after it was switched ON. also it would be good to have the complete code and check if it's not an absolute version. Otherwise after the checks if it still doesn't work and it's still under warranty, I suggest they send it back . Best, Nicolas
INC268055 | Workgroup | Incident | MEX5 With Range 0 to 760 mmHg | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-23 05:51 | 2022-09-23 10:12 | 0 | 0 |  | Hi Bee Tin, I was surprised, it exist in cmHg but not in mmHg. Best, Nicolas
INC267999 | Workgroup | Incident | BARFLEX | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-22 13:42 | 2022-09-22 14:57 | 0 | 0 |  | Hi Rudi, we don't do the IrDa tech anymore, anyway the componenets doesn't exist anymore. so there will be no options 0585 and 0588. the rest would be possible, but B88 doesn't exist, I guess it is B38 EA4-AC3.B38R/0587/0024 Best, Nicolas
INC267998 | Workgroup | Incident | MFT7 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-22 13:39 | 2022-09-22 14:59 | 0 | 0 |  | Hi Jurgen, Yes, the code is released and it's just a specific dial, you can go ahead with it: 11046137 MFT7-213.X10.Q82.100/9709_D546 [Image] Best, Nicolas
INC267904 | Workgroup | Incident | STEP file | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-21 15:39 | 2022-09-21 15:56 | 0 | 0 |  | Hi Jurgen, I don't have the fitted with complete but I have the MEX3 and the DANC. you can find the files in the ticket. Best, Nicolas
INC267886 | Workgroup | Incident | WG: AKP_INQ20220921_1302 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-21 13:18 | 2022-09-27 09:58 | 0 | 0 |  | Hi Carlo, I would suggest this mounting, but i'm also investigating on one with a TED: PFMN-64.BC4R.A1153.33200.4003 + 1627 + LRS1 Best, Nicolas
INC267882 | Workgroup | Incident | Barflex EA4 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-21 11:51 | 2022-09-21 16:16 | 1 | 0 |  | Hi Bee Tin, No :( you can go down to +/-0.1% of the full scale with a EA5A, which means +/-0.4bar of 400bar: [Image] 10270723 EA5-AP0.B38R Best, Nicolas
INC267855 | Workgroup | Incident | differencial pressure gauge spcial scale | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-21 09:12 | 2022-09-28 16:15 | 0 | 1 |  | Dear Zhang Wei, sorry for the delayed answer, we can offer you the 2 ranges below but the accuracy will change to 3 or 4% (to be determined via tests). lEM -10.6 bar MX7-A63.X16H lEM -0.90.1 MX7-A63.X15H Best, Nicolas
INC267813 | Workgroup | Incident | WG: AW: Inquiry thermowell | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-20 16:12 | 2022-09-20 17:12 | 0 | 0 |  | Hi Carlo, yes we can, it's the supplier of thermowells who'll do the tests and provide the reports. when ordering, you need to add a Q1201 with the comments "external + internal hydrostatic test and reports" Best, Nicolas
INC267811 | Workgroup | Incident | WG: Cert | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-20 15:51 | 2022-09-20 17:15 | 0 | 0 |  | Hi Carlo, Obviously we don't have a certificate for that but I can give you the list of the products we use to clean but it's basically Alcohol/water. try with this, otherwise we have to involve quality. Best, Nicolas
INC267777 | Workgroup | Incident | Louwers Mech. Parts (55153226) - PO 702200059 - Baumer OC 1906875 - RP2-EU.654 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-20 10:48 | 2022-09-20 14:49 | 0 | 0 |  | Hi Rudi, first thing first, no setting was required, if they want something they can ask for the SETP option and we set it here according to their instructions. But I see that we usually set it mid range with a minimum deadband as standard and it's not really the case on the first one. We have setting instructions for the RP2, they have it in the boxes and I see that they have what is needed as they tested it. it's quite simple to set, there is 2 screws, the good way to do it is to first reduce the hysteresis, then change the setp to what they need and finish by adjusting the hysteresis. Best, Nicolas
INC267773 | Workgroup | Incident | RPPN7 / Jack | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-20 10:42 | 2022-09-20 14:18 | 0 | 0 |  | Hi Bee Tin, please find enclosed the datasheet of the PX. I would suggest to go more standard with the instrumensts below, but the set and reset point must be confirmed, I'm not clear with the wording they use sor the "Hysteresis" (I guess they talk about the hysteresis). Order code: RP2-NL.866/SETP RP2 old Code / SAP code: Type of switch: RP2N N Type of microswitch: Microswitches RP2N RP2Y 1 standard changeover switch L Hydraulic connection: 1/4" NPT female 8 Pressure range (bar): Static pressure up to 100 bar, peak 200 bar 0 ... 100 bar 66 OPTIONS Accuracy / Setting: Set point adjustment SETP Order code: RP2-NP.866/SETP RP2 old Code / SAP code: Type of switch: RP2N N Type of microswitch: Microswitches RP2N RP2Y 1 hermetically ultra sensitive changeover switch P Hydraulic connection: 1/4" NPT female 8 Pressure range (bar): Static pressure up to 100 bar, peak 200 bar 0 ... 100 bar 66 OPTIONS Accuracy / Setting: Set point adjustment SETP Best, Nicolas
INC267765 | Workgroup | Incident | RE: AIOT: ขอราคา Baumer 5 Set. | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-20 09:24 | 2022-09-20 09:29 | 0 | 0 |  | Hi Bee Tin, I thought I answered, I'm very sorry. this seems to be an MJX7 back connection, but it's one from before the traceability with this label was fully operational. And I can't define much without seing the front. Best, Nicolas
INC267739 | Workgroup | Incident | FW: MDR documents MEX5 Latex free REACH & RoHS. | Archive > BHFR Technical Support | Middle |  |  | Closed | Adel, Haroon [adh] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-19 16:33 | 2022-09-20 14:42 | 0 | 0 |  | Hi Haroon, here is the document (in the ticket). Best, Nicolas
INC267712 | Workgroup | Incident | SAP codes ? - Fluxys | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-19 14:21 | 2022-09-19 15:58 | 0 | 0 |  | Hi Rudi, the codes are created, Josette, will validate it and send you the codes. The codes in DO were test codes, it'll be cancelled shortly. Best, Nicolas
INC267633 | Workgroup | Incident | TR: spez. MEX3 for Bobst | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-19 08:27 | 2022-09-21 14:25 | 0 | 0 |  | Nathalie, Gabriel, Sachant que nous ne trouvons dans nos archives qu'une seule commande sur ces dernières années, que la commande n'est que pour une pièce, que c'est hors standard (Bourdon and EN). Et que cela nécessite une nouvelle procédure et un outillage spécifique. Nous ne pouvons pas vous proposer cette option spécifique pour l'instant. Bien sur, nous en profitons pour conseiller au client de passer a un montage selon les normes en vigueur, cela lui permettra de commander des pièces standard chez la plupart des fournisseur et lui fera des économies en terme de prix des appareils. si le besoin s'avère plus important que prévu en terme de quantité, n'hésitez pas a nous recontacter. Best, Nicolas
INC267572 | Workgroup | Incident | EA4-AP3.B38R/0030, 40 units | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-16 10:00 | 2022-09-16 10:09 | 0 | 0 |  | Hi Bee Tin, I don't give offers, prices, lead times etc... but I can confirm you the code you asked for is possible EA4-AP3.B38R/0030. I don't see any seal with the code "2"in our products sorry. Best, Nicolas
INC267571 | Workgroup | Incident | New MEP5 -case 304 versus case 316 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-16 09:57 | 2022-09-16 10:12 | 0 | 0 |  | Hi Rudi, for now you can go ahead with the MEP5 304SS, there will be an official communication before it stops. and for that the new datasheets, communications etc... needs to be done and it's not done for now. but please start warning your customer that in the future we will standardize the 316. Best, Nicolas
INC267570 | Workgroup | Incident | FW: RFQ material DLT614148 THERMOMETER "BOURDON" DIA 63 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-16 09:42 | 2022-09-16 10:19 | 0 | 0 |  | Hi Rudi, there was a problem attaching it to the configurator, you can see an option 9003 with nothing behind. it was a very old code sold to wartsila once in 2010. I would suggest to not offer a special one again, but to try to define something more standard. Best, Nicolas
INC267560 | Workgroup | Incident | FW: Bourdon Haenni Pulsar - Epcor - Schiphol-Rijk NL | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-16 08:51 | 2022-09-16 09:21 | 0 | 0 |  | Hi Rudi, we don't do those since more than 10 years, and it was trade. I don't have anyone that have the competences on this since long. but what I can suggest is some of the usual solutions: put the machine off , unplug the RTD and replug it and restart (it should restart the link between the 2) change the RTD for one you're sure is working. if the 2 solutions doesn't work, there is a problem with the oven itself and they probably need to find a new one. Best, Nicolas
INC267394 | Workgroup | Incident | TR: inquiry MEX/MCX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-14 12:30 | 2022-09-14 14:13 | 0 | 0 |  | Nathalie, oui c'est OK pour tous (on vient de créer un MCX5 N66 sur D631 qui a la même membrane). tu feras attention, les 2 premiers sont les mêmes, il peut y avoir une erreur. Best, Nicolas
INC267343 | Workgroup | Incident | Updated 3A & FDA certificates for DAxx / Wandee | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-14 08:12 | 2022-09-14 10:20 | 0 | 0 |  | Dear Bee Tin, here are the FDA and 3A certificates but be careful, the 3A is not only for the DAxx, but for the full assembly with the instrument and the FDA concerns the filling fluids. Best, Nicolas
INC267285 | Workgroup | Incident | MEX3 back connection with wingnut / CSL | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-13 12:06 | 2022-09-13 14:23 | 0 | 0 |  | Hi Bee Tin, yes, you can find the MEX3 type G in the datasheet, configurator, P2F etc... If they checked the datasheet of the crimped MEX3, it's not available, but on the MEX3 with bezel ring, no problem. Best, Nicolas
INC267282 | Workgroup | Incident | LRS10 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-13 11:36 | 2022-09-13 12:22 | 0 | 0 |  | Hi Bee Tin, here is the MSDS and the fluid list, it is FDA compliant. Best, Nicolas
INC267163 | Workgroup | Incident | TBI080-252.160.13T/9003_0085 For ZPT4 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-12 12:21 | 2022-09-12 14:51 | 0 | 0 |  | Hi Bee Tin, Yes, what you propose will fit. When you have a sliding connection, don't hesitate to take 10mm more on he TBI length. Best, Nicolas
INC267105 | Workgroup | Incident | FW: AIOT: ขอราคา Baumer 5 Set. | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-12 08:39 | 2022-09-12 15:36 | 1 | 0 |  | Hi Bee Tin, this label was done before our traceability system was fully operational. Can you ask the customer for more information (order n, SAP code, ordering codes) or picture ? then I can help you finding out what to offer. Best, Nicolas
INC267082 | Workgroup | Incident | BACA ¦ Baumer Accumulator Pressure Transducer | Archive > BHFR Technical Support | Middle |  |  | Closed | Schmid, Sascha [schsa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-12 07:39 | 2022-09-12 14:45 | 1 | 0 |  | Hi Paul, the code you ask for is the 10301726 but it was closed because we don't renew the FM/CSA certifications (see explanation letter) but the code below is the same without the FM/CSA and you can ask for this code to be reopen, I don't see any other issue. 11117668 Y91-3.5.7.H34.R/2158_2151_2154_2169 Best, Nicolas
INC267066 | Workgroup | Incident | TR: Hjælp US parts | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-09 16:06 | 2022-09-09 17:02 | 0 | 0 |  | Nathalie, On a vérifié avec Nico et on a aucun fournisseur américain sur les composants du MEX5 demandé. Best, Nicolas
INC267010 | Workgroup | Incident | TR: Requirement of Hygienic Pressure gauge for Abbott Singapore project | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-09 10:03 | 2022-09-09 10:20 | 0 | 0 |  | Nathalie, ci-dessous les options disponibles au config' avec le montage: [Image] Best, Nicolas
INC267006 | Workgroup | Incident | TR: Kindly confirm the price - Model DP100-213-13-B-545 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-09 09:30 | 2022-09-09 10:22 | 0 | 0 |  | on a eu quelques limitations sur les membranes, mais tu peux utiliser ce code (on a limité la surpression a 9bar dans tous les cas). [Image] Best, Nicolas
INC266990 | Workgroup | Incident | FW: Verpakking levering | Archive > BHFR Technical Support | Middle |  |  | Closed | Adel, Haroon [adh] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-09 08:12 | 2022-09-09 10:15 | 0 | 0 |  | Hi Haroon, as you noticed, on the shipment boxes, we have the obligation to have certain information legally. but on the product and product boxes, there is an optioon for customer private labelling (only customer's name). you'll need the options 9912 and 9710, D000 is for the special dial and the rest is internal but needs to be defined with sales administration and you and your customer. hereafter an example for a customer I was managing before: [Image] Best, Nicolas
INC266987 | Workgroup | Incident | Contact Gauge / Steve | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-09 07:21 | 2022-09-09 10:01 | 0 | 0 |  | Hi Bee Tin, we can't replace the complete block, we could replace the parts of the cable gland (holding the cable) but with the look of the gauge, I would suggest to replace it. Best, Nicolas
INC266985 | Workgroup | Incident | 10270773- TESTEUR PORTABLE P/T° | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-08 23:42 | 2022-09-20 09:23 | 0 | 1 |  | Hi Daniel, I had the answer about this unit, and as it's linked to the E5 Zodiac, this unit is also stopped. Also we thought all of that was centralised in france. I close the ticket with this answer but if you have any feedback to get me, please send it to my personnal email. Best, Nicolas
INC266912 | Workgroup | Incident | type code is different | Archive > BHFR Technical Support | Middle |  |  | Closed | Zhang, Wei [zhw] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-08 11:30 | 2022-09-13 16:17 | 0 | 0 |  | Dear Zhang Wei, we confirmed and delivered the follower for the old article-number. Design and function are still the same. Only the article number changed. Best, Nicolas
INC266771 | Workgroup | Incident | TR: Material 10298527/ Jack | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-07 12:18 | 2022-09-07 14:33 | 0 | 0 |  | Nathalie, oui, c'est bien le code qu'il a donné, c'est le 3eme code ci-dessous: [Image] et on a du stock Best, Nicolas
INC266702 | Workgroup | Incident | WG: Inquiry thermowell | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-07 08:48 | 2022-09-07 09:23 | 0 | 0 |  | Hi Carlo, according to the text inside the AGWR and the reason it's blocked, it's not possible under 80mm. you have to check the AGFR-0260N2F.060 10264274 Best, Nicolas
INC266676 | Workgroup | Incident | TBI100-2N2.162.13T/0752 / Christine | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-07 03:42 | 2022-09-07 09:32 | 0 | 0 |  | Hi Bee Tin, As this is for a price, I transfer this to Nathalie, Nathalie, pou le prix tu peux prendre comme référence ce code (identique mais en 60mm): 11256513 TBI100-2N2.160.13T/9003_0060_9414_0752 Best, Nicolas
INC266668 | Workgroup | Incident | TR: Dde information serrage manomètre // YD | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-06 16:51 | 2022-09-07 11:15 | 0 | 0 |  | Salut Hervé, les serrages qu'on applique sont: G1/2 joint Alu : 80Nm environ G1/4 joint Alu: 50NM environ Best, Nicolas
INC266667 | Workgroup | Incident | customer logo on pressure sensors engraved | Archive > BHFR Technical Support | Middle |  |  | Closed | Adel, Haroon [adh] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-06 16:48 | 2022-09-06 17:42 | 0 | 0 |  | Hi Haroon, yes it is possible, it is the option 9750 if I recall correctly "Specific LOGO without Bourdon LOGO" but a test will need to be done (to validate the result is OK for the customer) and there might be a minimum order (100units/order seems ok) you can check the price with Celine. Best, Nicolas
INC266648 | Workgroup | Incident | TR: PRICE | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-06 15:03 | 2022-09-07 11:12 | 0 | 0 |  | Nathalie, pas possible en MEX5 mais possible en MX35: MX35-D31.B20.DAVA23.2L2A.01/0751_9911_9912 Best, Nicolas
INC266579 | Workgroup | Incident | FW: Question about MFE42 | Archive > BHFR Technical Support | Middle |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-06 11:33 | 2022-09-06 12:11 | 0 | 0 |  | Hi Giorgio, this is the technical support mailbox, you need to send the request to the "material database" adress: FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> Best, Nicolas
INC266554 | Workgroup | Incident | TR: AKP_INQ20220831_1818 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-06 09:15 | 2022-09-06 10:29 | 0 | 0 |  | Nathalie, oui c'est OK et ce sera du 0764 selon le config. Best, Nicolas
INC266325 | Workgroup | Incident | FW: ETTN | Archive > BHFR Technical Support | Middle |  |  | Closed | Chavan, Sudhir [chsu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-05 07:09 | 2022-09-05 09:47 | 0 | 0 |  | Hi Sudhir, Yes it is still available and please find enclosed the latest datasheet. in order to give you a price, your sales administration contact will need you to define the designation of the product you need. Best, Nicolas
INC266322 | Workgroup | Incident | pressure gauge and diaphragm seal confirmation | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-05 06:36 | 2022-09-07 11:08 | 0 | 0 |  | Hello Xu Quiang, both codes are OK technically. Best, Nicolas
INC266307 | Workgroup | Incident | TR: Termostatos PEDIDO: 42176817 y 42115309 ¦ OP Gamesa 08032022 20448826 ¦ GP399128 ¦ RT-NAA.403.120C3 // Document CE pour le thermostat | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-02 16:15 | 2022-09-07 11:18 | 0 | 0 |  | Salut Hervé, c'est dans les mains de la suisse pour approbation, on ne l'aura pas tout de suite je suis désolé. Best, Nicolas
INC266304 | Workgroup | Incident | Défaillance affichage TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-02 15:51 | 2022-09-02 16:36 | 0 | 0 |  | Bonjour Jacques, Vu les circonstance et la période, je pencherais pour des problèmes liés a l'installation en extérieur avec les conditions extrêmes qu'a subit l'inde cette année. mais il faut absolument faire une NQ avec Celine pour qu'on puisse faire une expertise et vous donner une solution concrète. Best, Nicolas
INC266292 | Workgroup | Incident | FW: French Manual | Archive > BHFR Technical Support | Middle |  |  | Closed | Viret, Shane [shv] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-02 15:15 | 2022-09-02 15:50 | 0 | 0 |  | Hi Shane, here it is, it's not perfect but I think it can help. just so you know, Google translate can directly translate a PDF, and as you can see, the result is not bad. Best, Nicolas
INC266255 | Workgroup | Incident | FW: Our inquiry / prijsaanvraag 22255117 | Archive > BHFR Technical Support | Middle |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-02 12:06 | 2022-09-02 14:20 | 0 | 0 |  | Hi Kees, not an easy question as none of our product is Lloyd now (the Y9 was) and we don't have HC276 wetted parts anymore. but I would still suggest to go with the Y92 flush (enclosed a paper about why we didn't renewed the Marine approval) and if the Hastelloy is really needed, a diaphragm seal. for now it'll be a DT2, but I know we worked on having the D04 in Hastelloy too. Option 1: Y9233GB20R/0809 Option 2: Y91333B20R/0809 + DT22L6002(or 6 or coated)3001 if the piezo is important to the customer, there is the PBMN option (also possible with seal). Best, Nicolas
INC266176 | Workgroup | Incident | RE: pressure transmitter marine applications | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 17:42 | 2022-09-01 18:05 | 0 | 0 |  | Answered.
INC266177 | Workgroup | Incident | TR: TED6 - ATEX Zone 2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 17:42 | 2022-09-02 09:37 | 0 | 0 |  | Salut Nathalie, Non du coup on a rien de vraiment équivalent. au plus proche on pourrait avoir le PFMN, on peut toujours leurs envoyer le lien pour la datasheet au cas ou mais c'est quand même très différent dans la forme même si globalement ca pourrait être une alternative. Best, Nicolas,
INC266174 | Workgroup | Incident | TR: BAIT ¦ Customer Euro 6 ¦ #2 adattatori 1/4 NPT F --> 1" 150 RF | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 17:15 | 2022-09-01 18:29 | 0 | 0 |  | Pas de soucis, sauf pour la face surélevée, ce sera un code G pour du 1" 150. ABR2214G08 par contre il faut demander le prix aux achats. Best, Nicolas
INC266165 | Workgroup | Incident | TED5 hygenic 0...6 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 16:27 | 2022-09-01 16:55 | 0 | 0 |  | Bonjour Jacques, la mesure ne sera pas assurée mais sur une dépression (de l'ordre de 10 mbar absolu) le système ne sera pas endommagé. la valeur que je donne est juste indicative. si ils savent a combien ils vont tirer au vide je pourrais répondre plus précisément. Best, Nicolas
INC266164 | Workgroup | Incident | RE: pressure transmitter marine applications | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 16:24 | 2022-09-01 16:25 | 0 | 0 |  | Double request
INC266162 | Workgroup | Incident | pressure transmitter marine applications | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 16:21 | 2022-09-01 16:26 | 0 | 0 |  | Angel, I didn't noticed you used the ticketing system, I should have answered you here. Answered via email with the technical note about FM/CSA. Best, Nicolas
INC266010 | Workgroup | Incident | Enquiry For MEC5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-09-01 05:33 | 2022-09-01 10:11 | 0 | 0 |  | Hi Bee Tin, the MEC5 is an old product and is now replaced by the MR5 (similar exept the front face design and electrical connection location). you can offer the products below and I enclosed the old datasheet of the MEC5: 1) MR5-4D30-B20 - 1 unit 2) MR5-4D30-B22 - 1 unit 3) MR5-4D30-B24 - 1 unit Best, Nicolas
INC265986 | Workgroup | Incident | find the right product | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-31 15:48 | 2022-08-31 16:35 | 0 | 0 |  | Hi Sasha, here is the most probable equivalents: MDX7 A60D15D31W316AT MX7-A60.D15L (would be 100bar static but then you need an MFT7) . MDJ7AN0D15D34LNG MFT7-A00.D15Q.N2 . MDJ7N0D15X-0201 (10290621) MFT7-A00.D15N.N2/0759 (option availability and static pressure to check) . MDX7 A60 D15 D31-W-316-AT MX7-A60.D15L (would be 100bar static but then you need an MFT7) Best, Nicolas
INC265918 | Workgroup | Incident | BAIT - Danella | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-31 09:45 | 2022-08-31 10:31 | 0 | 0 |  | Hello Francesco, D63 is not a type of gauge and the OEM number is not in the system. you'll need to contact the customer to define the product. Best, Nicolas
INC265880 | Workgroup | Incident | TR: Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-30 16:48 | 2022-08-31 09:14 | 0 | 0 |  | Yes, the mounting is OK. Best, Nicolas
INC265879 | Workgroup | Incident | TR: inquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-30 16:36 | 2022-08-31 14:16 | 1 | 0 |  | Nadia, peux tu corriger ce code pour le rendre plus clair stp ? ce code : 10326773 T 8911-II-LS=250 SCHUTZHU. K. devrait contenir ces informations: T8911.0203 T 8911: thermowell with threaded connection internal diameter 14mm 02: G3/4 process connection stainless steel 1.4571 03: 250mm datasheet ci-jointe Best, Nicolas
INC265854 | Workgroup | Incident | PO 2777 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-30 14:36 | 2022-08-30 15:58 | 0 | 0 |  | Celine, tu peux utiliser ce code: 11235609 Best, Nicolas
INC265716 | Workgroup | Incident | BAIT - ANDRITZ HYDRO - M363 | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 15:33 | 2022-08-30 09:44 | 0 | 0 |  | Dear Francesco, this looks like an old version of an MR5, before we stopped this kind of windows. I would suggest to replace by the modern version but I miss some éléments like the connection, contacts opening/closing or options: MR5XBX0S25 Best, Nicolas
INC265675 | Workgroup | Incident | TR: 10325462 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 12:00 | 2022-09-02 14:54 | 0 | 1 |  | Nathalie, On peut faire ce code 10325462 (confirmation Christophe) mais il faut demander une consultation d'usinage avec le nombre de pièces aux achats. Best, Nicolas
INC265651 | Workgroup | Incident | Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 10:12 | 2022-08-29 14:14 | 0 | 0 |  | Bonjour Cyrille, voici le MTTF (dans le ticket). Best, Nicolas
INC265649 | Workgroup | Incident | FW: request 3A & FDA certificate for DAxx from Flowmaster | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 10:00 | 2022-08-29 11:47 | 0 | 0 |  | Dear Bee Tin, here is 2022 certificate: Best, Nicolas
INC265625 | Workgroup | Incident | Gaskets between Pressure gauge and Siphons | Archive > BHFR Technical Support | Middle |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 09:09 | 2022-08-29 11:45 | 0 | 0 |  | Hi Solveig, if it's G type connections, then the aluminium gasket is provided with the gauges. for NPT, there is no need for gasket, it's usually mounted with PTFE taping. Best, Nicolas
INC265623 | Workgroup | Incident | TR: SAP code | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-29 09:09 | 2022-08-29 16:48 | 0 | 0 |  | Nathalie, la limite pour faire du 0.7% est 0...1.6bar, with a 1bar cell we cannot do it anymore. Best, Nicolas
INC265566 | Workgroup | Incident | TR: DIN 43650 connector | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-26 16:27 | 2022-08-29 16:43 | 1 | 0 |  | Nathalie, here is what the 2297 option looks like: [Image] but I don't think they need to replace the metallic part and the code for just the plastic one is 10303707. Best, Nicolas
INC265522 | Workgroup | Incident | MAR5 Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-26 10:03 | 2022-09-09 09:36 | 1 | 1 |  | Dear Bee Tin, on top of the answer I gave you already, I have news about the scale. we will maintain the bar Bourdon (A) scales, so no change on this. Best, Nicolas
INC265026 | Workgroup | Incident | Request Anfrage 709515JH | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-19 14:27 | 2022-08-31 14:04 | 1 | 1 |  | Technically answered, details of product definition to be verified with customer. Best, Nicolas
INC264314 | Workgroup | Incident | WG: Anfrage Druckschalter / Projekt 343156 | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-10 16:12 | 2022-08-29 12:16 | 1 | 1 |  | Answered via email. Best, Nicolas
INC264070 | Workgroup | Incident | We You create the following pressure transmitter to a order Y913 Intrinsically Safe Pressure TransmitterVS: Y913 GZ G1/4 LCIE 02 ATEX 6133X for Air liquide. | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-08 13:57 | 2022-08-29 10:32 | 1 | 1 |  | Dear Colleague, I'm not sure I understand your question. but yes we do this kind of Y913, we would suggest this code with a Viton seal (instead of the EPDM in the code you had) 11134215 Y91-3.2.9.B35.R/0765 you can then get price from your usual sales contact. Best, Nicolas
INC263910 | Workgroup | Incident | Y91-2.5.9.B16.A_22200-000486A/ Mark | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-05 03:27 | 2022-08-05 09:18 | 0 | 0 |  | Hi Bee Tin, I don't give price and lead time but we first need to clarify this, the yellow part makes no sense to me, you know where it comes from ? Y91-2.5.9.B16.A_22200-000486A Best, Nicolas
INC263782 | Workgroup | Incident | pressure of TSF stem | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-03 17:15 | 2022-08-03 17:26 | 0 | 0 |  | Hello Xu Qiang, concerning the resistance of thermometers, it relies a lot more on geometry, type of fluid, velocity than just pressure. Usually, to confirm a thermometer fits an application, we use what is called stress calculation, we can do that but there is a cost. also, usually our thermometers are used in thermowells because it avoid this issue (stress calculation can also be done if needed even if a thermowell is much more solid) and it makes maintenance easier. you can see with Jean Marie Moreau for stress calculation (back in 2 weeks) and you can offer an option with thermowells. Best, Nicolas.
INC263702 | Workgroup | Incident | replacements for TED6-3.9.B72.R/0609. type today and SAP no for this to order. for Greenland ODN@nukissiorfiit.gl | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-03 09:24 | 2022-08-03 11:25 | 0 | 0 |  | Hello Jens, TED6 doesn't have an 8pin output, so the cable 10m M12 for TED6 will be the 0606. TED639B72R/0606 Best, Nicolas
INC263657 | Workgroup | Incident | CE certificate of CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-03 05:42 | 2022-08-03 11:27 | 0 | 0 |  | Hi Xu Qiang, here is the CTX certificate (in the ticket): Best, Nicolas
INC263540 | Workgroup | Incident | TR: Material creation : 11014577 / ZDPN--211SGM6/SETP | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-02 11:24 | 2022-08-02 11:49 | 0 | 0 |  | Celine, je ne vois pas de raison technique, il y a un commentaire de 2013 qui dit qu'il a été bloqué car il faut réviser le prix. sur la dernière commande il avait été annulé et remplacé pour le SETP (demandé finalement sans SETP). [Image] Best, Nicolas
INC263539 | Workgroup | Incident | Demande offre | Archive > BHFR Technical Support | Middle |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-02 11:24 | 2022-08-02 11:54 | 0 | 0 |  | Hi Giorgio, I would suggest to replace this by a standard MEX3 (datasheet enclosed for redefinition). concerning the ring, I don't find it in our system and I don't know what it is. if it's the G1/4 Aluminium O'ring, we don't provide it. Best, Nicolas
INC263534 | Workgroup | Incident | Enquiry For EA5-AP0.B38A/0030 / Moore | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-02 11:03 | 2022-08-02 11:06 | 0 | 0 |  | Dear Bee Tin, please send this enquiry to celine martin at the sales administration. Best, Nicolas
INC263526 | Workgroup | Incident | Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-02 10:39 | 2022-08-02 11:05 | 0 | 0 |  | Hi Giorgio, this is from the old Linkprogrammer which is replaced by the flexprogrammer. it'll allow the same programming, please find hereafter the link. FlexProgrammer 9701 ¦ Accessories process sensors ¦ Baumer France "https://www.baumer.com/fr/en/product-overview/accessories/accessories-process-sensors/interfaces/flexprogrammer-9701/p/38174" Best, Nicolas
INC263501 | Workgroup | Incident | TR: demmande de ce pour rpn2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-02 09:18 | 2022-08-02 09:52 | 0 | 0 |  | Désolé, le certificat CE du RP2N est en cours et il ne contiendra que la directive basse tension. Best, Nicolas
INC263467 | Workgroup | Incident | manomètre pour terminal méthanier | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-01 16:36 | 2022-08-03 14:38 | 0 | 0 |  | Bonjour Cyrille, je suis désolé, mais nous n'avons pas d'idée précise du comportement d'un capillaire (ou ASIPD) à ces températures. le froid est un des axes de développement de nos produits et connaissances actuellement mais je n'ai pas de réponse a apporter dans un futur proche. Désolé, Best, Nicolas
INC263456 | Workgroup | Incident | VS: 0-100 bar tryktransmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-01 15:18 | 2022-08-01 16:03 | 0 | 0 |  | Answered via another ticket number.
INC263446 | Workgroup | Incident | BAIT - Fluid Processing - Thermometer | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-01 14:15 | 2022-08-02 09:38 | 0 | 0 |  | Hi Francesco, you can offer this, I'll check for the certificates this afternoon. Order code: TBI100-2N2.161.13T/0752 TBI old Code / SAP code: Nominal size: DN 100 100 Case AISI 304 (1.4301), stainless steel: AISI 304 (1.4301), stainless steel 2 Process connection: DN38 ISO (10318762) N Immersion tube outlet: Center back 2 Indicator: EN 13190 (Standard) 1 Immersion tube diameter: 8 mm AISI 316Ti (1.4571), stainless steel 6 Immersion tube length L1: 60 mm 1 Temperature range: Single scale inC 0 ... 160C 13T IC after discount: 172 Best, Nicolas
INC263390 | Workgroup | Incident | drawing of T 8911-II-LS=250 SCHUTZHU. K. | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-01 07:18 | 2022-08-02 09:14 | 0 | 0 |  | Hi Xu qiang; please find the datasheet in the ticket, your designation is T 8911-0203 Best, Nicolas
INC263389 | Workgroup | Incident | Material no. 11017380 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-08-01 06:36 | 2022-08-01 09:32 | 0 | 0 |  | Dear Bee Tin, option 0808 is not available anymore as we didn't renewed it. please find enclosed the communication for customers. Best, Nicolas
INC263370 | Workgroup | Incident | EA4A -1000+1600MBAR R + PILE ZODIAC - SAP 10270571 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-29 15:21 | 2022-07-29 15:27 | 0 | 0 |  | Hi Daniel, sure, here it is: Best, Nicolas
INC263356 | Workgroup | Incident | technical inquiry for ZPN-151SCM/9704(-100~100mar) | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-29 13:36 | 2022-07-29 15:25 | 0 | 0 |  | Hi Xu Qiang, I had to check with R&D for that and the best compromise we can offer you is a ZPHN104SCM. -8...100mbar (but with 10bar allowable overpressure so the 200mbar is not an issue) setp at 2mbar falling Best, Nicolas
INC263129 | Workgroup | Incident | we need the variant that fits with this see PIC TED-6.3.9.B24.R/0606 | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 13:33 | 2022-07-27 14:55 | 0 | 0 |  | Dear Andersen, here is the code from the order 1045012/20 11169325 TED-6.3.9.B24.R/0606 Best, Nicolas
INC263128 | Workgroup | Incident | Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Viola, Giorgio [gvi] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 13:06 | 2022-07-27 14:23 | 0 | 0 |  | Hi Giorgio, it is possible, here is an example code that is the same price: [Image] you can ask for the code creation. Best, Nicolas
INC263107 | Workgroup | Incident | DP100 hygenic electro polishing | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 11:18 | 2022-07-27 14:07 | 0 | 0 |  | Hi Xu Qiang, here is both codes, and option 0194 is now mandatory to limit the max pressure to 9 bar: Ra: 0.8 is the standard. [Image] [Image] Best, Nicolas
INC263106 | Workgroup | Incident | taux de ferrite | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 11:15 | 2022-07-29 16:10 | 0 | 0 |  | Celine, après avoir discuté avec Paul, on a pas de moyen de contrôle du taux de ferrite dans les soudure. pour info, la ferrite se forme en grande partie lors de la fusion de l'inox sur les soudures. nos soudures de membranes (en contact avec le fluides) sont faites automatiquement au laser et sont donc très bien au niveau technique. par contre on ne peut pas contrôler le taux de ferrite et il faut savoir que ce serait très compliqué et aléatoire a contrôler dans un cordon de soudure. Best, Nicolas
INC263057 | Workgroup | Incident | SAP Code | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 08:18 | 2022-07-27 14:24 | 0 | 0 |  | Hi Jurgen, I think you can directly request code creation to: FUNC BHFR Material Data <func-frven-mat-dat@baumernet.org> Best, Nicolas
INC263054 | Workgroup | Incident | FW: Janssen DPC100 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 08:15 | 2022-07-27 14:21 | 1 | 0 |  | Hi Jurgen, The customer request a replacement for : 11141026 DP100-213.13B.545/0655_9711_0194 The Dashpot seems to be an issue in production as long as the new diaphragm type is not in place. So we can provide it only without the dashpot (0655: dashpot and 9711: inscription mechanical dampening). DP100-213.13B.545/0194
INC263012 | Workgroup | Incident | MEX5-D32.D76/0693_0191_0720_0751_0760_9941 /Moore | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-27 06:48 | 2022-07-27 10:36 | 0 | 0 |  | Hi Bee Tin, here is the table: [Image] Best, Nicolas
INC262888 | Workgroup | Incident | TED-6.3.9.B24.R/0606 Is there a reason why this is blocked for Quotation etc. for Nukissiorfiit ole nielsen | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-25 15:36 | 2022-07-25 16:03 | 0 | 0 |  | Hi Jens, you can ask for reopening as I don't see any issue in the code. the text says it was closed because it wasn't solved since very long. Best, Nicolas
INC262826 | Workgroup | Incident | Type request | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-25 09:33 | 2022-07-25 10:15 | 0 | 0 |  | Hi Jurgen, the MDT7 is the old name of the MT7, the code bellow is the best fit except the static pressure is 10bar instead of 16bar. Order code: MT7-P62.N10H MT old Code / SAP code: Nominal size: 150 mm 7 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection with mounting bracket for 2"pipes P Process connection: 1/2" NPT 6 Oil filling: BH2, glycerin, high viscosity 2 Pressure unit: mbar (EN standard) N Pressure range: 0 ... 250 mbar 10 Static pressure: 10 bar H Best, Nicolas
INC262788 | Workgroup | Incident | vaccum resistance of Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-23 05:45 | 2022-07-25 10:11 | 0 | 0 |  | Hi Bee Tin, it can withstand a bit of vacuum, but it greatly depends on the intensity of it. for a rough vacuum it is OK. Best, Nicolas
INC262774 | Workgroup | Incident | Alternative product for 11102905 and 11118482 SAP code ARLA Slagelse 4200 | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-22 14:48 | 2022-07-25 10:08 | 1 | 0 |  | Hello, it's just old codifications, here are the standard to replace it: Order code: MEX5-D31.B20.DANC01.2L2A.01 MEX old Code / SAP code: Dimension ( Diameter): 100 mm 5 Type of mounting & material sensing element: Housing and bezel ring 1.4301 stainless steel Bottom connection D Internal connection (gauge/seal): G1/2" 3 Oil Filling: BH1, glycerin/water, low viscosity 1 Pressure unit: Bar B Pressure range: Pressure range for Bar 0 ... 6 Bar 20 Seal type: DA series DA Seal according to following standard: Clamp NF E 29521, ISO 2852 DN 38 NC01 Top Housing material Stainless Steel: Stainless steel 2 Top housing thread: G1/2" female L Diaphragm material: Stainless steel (316 L) 2 Surface finish: Standard finish (Ra 0.8) A Oil filling for seal: LRS1 (-15 ... 150C) vaseline 01 Order code: MEX5-D31.B22.DANC01.2L2A.01/0751 MEX old Code / SAP code: Dimension ( Diameter): 100 mm 5 Type of mounting & material sensing element: Housing and bezel ring 1.4301 stainless steel Bottom connection D Internal connection (gauge/seal): G1/2" 3 Oil Filling: BH1, glycerin/water, low viscosity 1 Pressure unit: Bar B Pressure range: Pressure range for Bar 0 ... 10 Bar 22 Seal type: DA series DA Seal according to following standard: Clamp NF E 29521, ISO 2852 DN 38 NC01 Top Housing material Stainless Steel: Stainless steel 2 Top housing thread: G1/2" female L Diaphragm material: Stainless steel (316 L) 2 Surface finish: Standard finish (Ra 0.8) A Oil filling for seal: LRS1 (-15 ... 150C) vaseline 01 OPTIONS Window: Safety glass, triplex 0751 Best, Nicolas
INC262679 | Workgroup | Incident | BAIT - Tonello - RMA: CL148973.1 - Y91 | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-21 12:39 | 2022-07-25 09:55 | 1 | 0 |  | Dear Francesco, I can't help on this, I have no power on what is invoiced on claims, please check this out with Nathalie or Celine. Best, Nicolas
INC262660 | Workgroup | Incident | Enquiry For PBSN22B31RA1440521 / Michael | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-21 09:00 | 2022-07-25 09:52 | 1 | 0 |  | Hello Bee Tin, yes you can order it, here is the code: 11233370 Best, Nicolas
INC262635 | Workgroup | Incident | Item specs | Archive > BHFR Technical Support | Middle |  |  | Closed | Latouche Cabrera, Jorge [ljo] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-20 18:27 | 2022-07-27 11:06 | 1 | 1 |  | Jorge, please find enclosed the latest drawing, if it's not the right drawing you're looking for, please tell me.
INC262556 | Workgroup | Incident | FW: demande de fichier CAO | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Lecomte, David [leda] | 2022-07-20 10:36 | 2022-08-29 12:09 | 1 | 1 |  | David sent the file. Best, Nicolas
INC262458 | Workgroup | Incident | DP100 SAPcode 11141026 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 15:51 | 2022-07-27 14:23 | 0 | 1 |  | Answered on other ticket. Best, Nicolas
INC262447 | Workgroup | Incident | VS: Datasheet and user manual for Y7 pressure transmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Stojanovic, Zoran [stoz] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 14:48 | 2022-07-19 15:30 | 0 | 0 |  | Hello Zoran, Sure, here is the datasheet (obviously not existing anymore, it is replaced by the Y913). Best, Nicolas
INC262441 | Workgroup | Incident | FW: Janssen DPC100 | Archive > BHFR Technical Support | Middle |  |  | Closed | De Mezel, Jürgen [deju] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 13:54 | 2022-07-19 14:08 | 0 | 0 |  | Hi Jurgen, the construction is different as a standard gauge, and not all options are possible, for example the dampening screws are not possible. but this code is possible, the code can be reopened [Image] Best, Nicolas
INC262420 | Workgroup | Incident | MEX3 & MEX5,G threaded connection with surface roughness ≤ 0.8µm Ra mech polish for wetted part / For Pharma Application /Christine | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 12:15 | 2022-07-19 14:11 | 0 | 0 |  | Dear Bee Tin, for pharmaceutical application, and especially if there is an Ra requirement, they will not accept a G connection and the Ra is not controllable (as the fluid is in contact with the internal of the connection and tube). No manufacturer can do that there must be a diaphragm. please check with them. Best, Nicolas
INC262379 | Workgroup | Incident | TR: Item material | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 09:48 | 2022-07-19 14:59 | 1 | 0 |  | Céline, Voici la matières: EQ-13-17-REP 8 [Image] Best, Nicolas
INC262377 | Workgroup | Incident | VS: 0-100 bar tryktransmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Ingversen, Morten Werge [inm] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 09:33 | 2022-07-19 11:41 | 0 | 0 |  | Thanks for sending it to the tech support email box, I answer already via Nathalie for this one. Best, Nicolas
INC262376 | Workgroup | Incident | TR: Question | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 09:30 | 2022-07-19 11:59 | 0 | 0 |  | Nathalie, voici les réponses: Clamp DN25: diaphragm size: 19mm Clamp DN38: diaphragm size: 32mm material in contact: 316L Filling fluid: LRS1 (FDA conform) Best, Nicolas
INC262374 | Workgroup | Incident | replacement of ZTN402CHN270 | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-19 09:27 | 2022-07-19 12:03 | 0 | 0 |  | Hi Xu Qiang, you're perfectly right, the CHN was phased out and our recomandation is to replace by a CHM. enclosed the datasheet so your customer can check if it fits the need (differences were really minimal). Best, Nicolas
INC262294 | Workgroup | Incident | TR: 0-100 bar tryktransmitter | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-18 11:33 | 2022-07-18 14:25 | 0 | 0 |  | Nathalie, One connection looks like an homogenizer connection (the diaphragm part) but ours use fixing flange and the one they have a big female thread. for 2 units, no new connection can be designed so we have no direct replacement. but you can present them the diaphragm seal we have for this type of application which is more standard. Best, Nicolas
INC262290 | Workgroup | Incident | FW: Our inquiry / prijsaanvraag 22245914 | Archive > BHFR Technical Support | Middle |  |  | Closed | van Bladel, Kees [vank] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-18 11:21 | 2022-07-18 14:35 | 0 | 0 |  | Hi Kees, here is the complete order details of this product [Image] Best, Nicolas
INC262272 | Workgroup | Incident | Thermomètre bimétallique | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-18 09:27 | 2022-07-18 11:02 | 0 | 0 |  | Bonjour Cyrille, pour l'instant on a rien trouvé, on a fait la demande pour un plan comme la datasheet n'existe pas. Désolé de ne pas pouvoir aider plus. Best, Nicolas
INC262201 | Workgroup | Incident | Survitec Items | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-15 09:33 | 2022-07-18 09:25 | 1 | 0 |  | Hi Daniel, I had the same request from France a couple days ago, and I'm sorry but I don't have a doc/datasheet to provide you. please make sure they update with iour complete code (especially SAP code ). Best, Nicolas
INC262136 | Workgroup | Incident | Degreasing Service for O2 Application MEX5-D30.D20.DANC01-2L2A.01/0693_0720_0751_0760_0764_9941 / CSL | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-14 11:12 | 2022-07-18 09:41 | 1 | 0 |  | Hello Bee Tin, this will not be possible with the LRS1 filling, you must use LRS4. then it is possible but not sure if it'll be the 0765 or 0764 option (O2 application EN compliant or not), I would need to check with quality. Best, Nicolas
INC262065 | Workgroup | Incident | RE: RE: ED70 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pollard, Jeff [jpo] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-13 14:42 | 2022-07-13 16:10 | 0 | 0 |  | Hi Jeff, here is my suggestion: [Image] Best, Nicolas
INC262031 | Workgroup | Incident | TR: Anfrage | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-13 11:21 | 2022-07-13 14:14 | 0 | 0 |  | Hi Thomas, I think it could be an old design of D295, please check witht the customer for the dimensions Best, Nicolas
INC261934 | Workgroup | Incident | TSS/TSF | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-12 15:06 | 2022-07-19 09:04 | 0 | 1 |  | Vu par email: bloquage des codes sans connection par precaution. Sujet a redisctuter avec Jorg prochainement. Best, Nicolas
INC261906 | Workgroup | Incident | TR: BAIT ¦ Oxygen application | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-12 11:42 | 2022-07-12 16:11 | 0 | 0 |  | Bonjour Nathalie, répondu par teams, le 0765 est OK sur le Y91. pour le prix, je demande a Antoine de te le transmettre. best, Nicolas
INC261804 | Workgroup | Incident | WG: Baumer gauge service life info | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-11 13:48 | 2022-07-11 15:57 | 0 | 0 |  | Hi Carlo, we didn't perform specific MTTF/MTBF tests as it's included in the EN837-1 and our gauges comply. there is a lot of different tests performed for all aspects, but the key one for you might be the endurance test with the cycles below. the gauge should still be conform to all characteristics after the test (accuracy, linearity, functionnality...) [Image] Best, Nicolas
INC261655 | Workgroup | Incident | MFE425040JB3-02 / 6346 73 016-03- SAP 10287907 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-08 10:42 | 2022-07-08 14:48 | 0 | 0 |  | Hi Daniel, I'm sorry but we don't have an FAI on this code. but we can do include one in the next order if Knorr-Bremse requires it. for the price, I let you check with Nathalie Best, Nicolas
INC261538 | Workgroup | Incident | OPTION 0745 // Compatibilité sur MEX5 ? | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-07 11:09 | 2022-07-07 11:47 | 0 | 0 |  | Hervé, oui c'est OK, j'ai un code ouvert et ca passe au config, par contre j'ai essayé avec un 6 bar, attention il peut y avoir des restrictions en basse pression. Best, Nicolas
INC261467 | Workgroup | Incident | TR: Customer request: product | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-06 15:51 | 2022-07-06 16:19 | 0 | 0 |  | Hi Celine, yes, in the picture of the TED6, you can see "bar R" which stands for bar relative and if you ordered an absolute (A in the last digit of the code instead of R) you would read bar A on the front of the product. Best, Nicolas
INC261431 | Workgroup | Incident | RE: refroidisseur ZPX1-007 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-06 13:57 | 2022-07-06 14:07 | 0 | 0 |  | Vue, erreur d'envoi.
INC261399 | Workgroup | Incident | TR: MEX5 with a DA seal | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-06 11:57 | 2022-07-06 12:10 | 0 | 0 |  | Hi Jim, you can look at the seal's guide for this, and you'll see that with a DADL05 (DN50) you can mount a MEX3 at 0...1bar. MEX35 would be also a solution. you can also find it on the datasheet. Best, Nicolas
INC261385 | Workgroup | Incident | TR: Request for prcie and model check | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-06 10:18 | 2022-07-18 15:58 | 0 | 1 |  | Suite a ton Email, je clôture le ticket, si ils reviennent avec des info on pourra répondre sans problèmes. Best, Nicolas
INC261378 | Workgroup | Incident | TR: Enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-06 09:45 | 2022-07-06 11:01 | 0 | 0 |  | Bonjour Jacques, je ne pense pas que ce soit un code de chez nous, je suis remonté jusqu'au premier catalogue bourdon et je ne trouve rien, de plus a aucune periode le code d'un appareil de mesure de pression n'a commencé par un Z. Best, Nicolas
INC261297 | Workgroup | Incident | TR: SPR for GUE / Tuanku Jaafar Power Plant | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-05 15:09 | 2022-07-06 15:18 | 0 | 0 |  | Nathalie, l'option 0722 est bien possible sur les MEX3, cela va ajouter une butée caoutchouc à 0 et supprimer une ou deux graduations après le 0 selon les échelles. Best, Nicolas
INC261292 | Workgroup | Incident | TR: Question about product | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-05 14:51 | 2022-07-05 16:21 | 0 | 0 |  | Nathalie, c'est un produit BTIN, ils doivent s'adresser directement a eux. Best , Nicolas
INC261268 | Workgroup | Incident | MR5 1A30 B20 4710 SETP+ AKPL 4 F5 3 02 1521 01 | Archive > BHFR Technical Support | Middle |  |  | Closed | Faulkner, Daniel [fad] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-05 12:15 | 2022-07-05 13:59 | 0 | 0 |  | Hi Daniel, yes it is possible, and the code you did is right, just think about indicating the setp you need when asking for the code: MR51A30B20/4710/SETP + AKPL4F5302 + 1521 + 01 (SETP = 3bar falling) Best, Nicolas
INC261165 | Workgroup | Incident | TR: Manometer til centrifuge | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-04 16:21 | 2022-07-04 16:54 | 0 | 0 |  | Nathalie, C'est OK pour le montage ci-dessous: MX35-D22.B22/0751 + DADL062H2A01 Best, Nicolas
INC261143 | Workgroup | Incident | WG: AW: AW: AW: Inq on Bourdon no 433 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-04 13:48 | 2022-07-04 15:21 | 0 | 0 |  | Hi Carlo, As said on the phone, your definition of the product is good, it's probably a misunderstanding from te customer. Best, Nicolas
INC261123 | Workgroup | Incident | BAIT - A1R - Flange | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-04 11:36 | 2022-07-06 09:01 | 0 | 0 |  | Discussed with Francesco, request seems closed.
INC261106 | Workgroup | Incident | TR: Baumer - Auftragsbestätigung: 1967742 zu 3359172 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-04 10:21 | 2022-07-04 15:29 | 1 | 0 |  | Nathalie, the black DIN 43650 connector is delivered with the device. Best, Nicolas
INC261087 | Workgroup | Incident | WG: RICHIESTA DI OFFERTA COMM.F21387 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-04 09:27 | 2022-07-04 16:05 | 1 | 0 |  | Hi Carlo, We are currently working on updating reach/rohs for TBI (september) for now the only thing we can provide is this group note: Best, Nicolas
INC261029 | Workgroup | Incident | TR: rfq DT2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-01 17:33 | 2022-07-25 09:45 | 1 | 1 |  | Celine, voici les CCR2: DT2-2L20.0.2X0.00 X= G3/4 (1pcs) 167,25 DT2-2L20.0.2Y0.00 Y= G 1 (3pcs) 107,45 Best, Nicolas
INC260969 | Workgroup | Incident | WG: Salt Power | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-07-01 11:18 | 2022-07-01 16:52 | 1 | 0 |  | Hi Carlo, here you go, the very limit for te switch is 10bar, I would suggest to switch at 15 (falling) for safety. it's fine for mount with the DT1 you defined, only the max pressure is limited to 160bar. Order code: RP2-NL.366/SETP RP2 old Code / SAP code: Type of switch: RP2N N Type of microswitch: Microswitches RP2N RP2Y 1 standard changeover switch L Hydraulic connection: G1/2" male 3 Pressure range (bar): Static pressure up to 100 bar, peak 200 bar 0 ... 100 bar 66 OPTIONS Accuracy / Setting: Set point adjustment (15bar falling) SETP Best, Nicolas
INC260924 | Workgroup | Incident | DP100 hygenic | Archive > BHFR Technical Support | Middle |  |  | Closed | Xu, Qiang [qxu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-30 16:30 | 2022-06-30 16:47 | 0 | 0 |  | Dear Xu Qiang, yes it is still available, the request to add it to the website with an updated datasheet was also done (but it's long). once defined, you can ask Nathalie for the price. Best, Nicolas
INC260919 | Workgroup | Incident | BAIT - Microswitches - request | Archive > BHFR Technical Support | Middle |  |  | Closed | Fochi, Francesco [fof] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-30 15:39 | 2022-06-30 15:42 | 0 | 0 |  | Hi Francesco, here is the code they have: Old SAP: MCX5D60N09-0202 New: MCX5-D60.N09/0751 10272905 Best, Nicolas
INC260837 | Workgroup | Incident | RFQ for the Pressure Gauges for IOCL AVU project | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-30 06:48 | 2022-07-12 15:47 | 0 | 1 |  | Hi Anil, after discussing and trying to find a solution, the mounting is too complex and we can't provide this. Best, Nicolas
INC260790 | Workgroup | Incident | TR: MEP5 pricing | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-29 12:21 | 2022-06-30 14:30 | 0 | 0 |  | Nathalie, il faut qu'il nous disent quel raccord ils veulent ci-dessous, et ensuite faire une SPR: [Image] [Image][Image] Best, Nicolas
INC260736 | Workgroup | Incident | TR: Request for Quote NEWZ220130R01 - thorm | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-29 09:06 | 2022-06-29 11:59 | 0 | 0 |  | Nathalie, ci-dessous le code du produit par Annie: 10323207 Best, Nicolas
INC260706 | Workgroup | Incident | BAIT ¦ Customer INTECH ¦ Tech questions | Archive > BHFR Technical Support | Middle |  |  | Closed | Dema, Erald [dee] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-28 16:39 | 2022-07-07 17:08 | 0 | 1 |  | Celine, nous proposons ce montage: RP2NR343/setp (250mbar falling) + DT3-2L62.0.C60.01 (CCR2 sans LRS: 358.7) Best, Nicolas
INC260611 | Workgroup | Incident | MVE-7.D60.B15, Minimum Scale Per Division | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-28 08:54 | 2022-06-28 09:57 | 0 | 0 |  | Hello Bee Tin, here is the drawing of the dial: [Image] Best, Nicolas
INC260608 | Workgroup | Incident | PMI Test For MEP5-D62.D22.D820-10N.2217.G6002/0719/0763/9941 / Avant | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-28 08:45 | 2022-06-28 09:14 | 0 | 0 |  | Hi Bee Tin, here is the code and IC price (per shot, here I guess you need one for the diaphragm but your customer might also ask for one on the body of the diaphragm seal). please note the lead time will be around 1week more than standard. MEP5-D62.D22.D820-10N.2217.G6002/0719/0763/9941/QPMI QPMI: 100.5 IC / shot Best, Nicolas
INC260568 | Workgroup | Incident | WG: DMS - RFQ ENQ12084/1 20467398_EVO | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-27 16:12 | 2022-06-28 13:53 | 0 | 0 |  | First draft answer done, a new ticket will be created in case of need for further investigation.
INC260549 | Workgroup | Incident | Impact test 27J.... // Sclumberger // GC | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-27 14:54 | 2022-06-27 16:02 | 0 | 0 |  | Hervé, sur nos Q003, dans le certificat du fournisseur de matiere, beaucoup de tests et resultats sont presentés dont le test de Charpy selon l'ISO148-1, ca devrait repondre a leur demande. tu peux le voir ci-dessous (ne pas diffuser, c'est le certificat d'un client). [Image] Best, Nicolas
INC260505 | Workgroup | Incident | Certificat de conformité CE | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-27 11:09 | 2022-06-27 11:46 | 0 | 0 |  | Bonjour Jacques, maintenant ils sont séparés par produits, je t'ai mis les 3 principaux pour l'elec. tu me dira si c'est d'un autre dont tu as besoin. Best, Nicolas
INC260489 | Workgroup | Incident | MCF5-430.N10/0693_0651_9941+ ARIN260L+AORPB501/SETP+ARIN2306 / Neste | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-27 10:12 | 2022-06-27 11:01 | 0 | 0 |  | Hello Bee Tin, there is a couple things I can't see on a picture (materials of case, movement or customer connection) but according to what you defined, here is what I advise : I would also advise to ceck the need for the stainless steel movement as the price adder is important: Order code: MCF5-460.N10/0693_0651_9941 MCF old Code / SAP code: Nominal size: 100 mm 5 Type of mounting: Case and bezel ring AISI 316L (1.4404), stainless steel Bottom connection 4 Process connection: 1/2" NPT 6 Oil filling: No oil filling 0 Pressure unit: bar (EN standard) B Pressure range: 0 ... 0,25bar 10 OPTIONS Movement: Stainless steel movement 0651 Dial: "G" added to the pressure unit 0693 Identification: Stainless steel label and wire 9941 Order code: AORPB501/SETP AORP old Code / SAP code: Design: for low pressure up to 16 bar B Pressure set-point in bar: 0.1 ... 0.4 bar 501 OPTIONS Accuracy / Setting: Set point adjustment SETP Order code: ARIN2606 ARIN old Code / SAP code: Material: Stainless Steel (1.4404) 2 Connector instrument side / Process side: Connector instrument side 1/2 NPT M M1/2 NPT / 1/2 NPT 606 Best, Nicolas
INC260403 | Workgroup | Incident | MEX5-D32.E19/0678_0751_0760 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-24 12:27 | 2022-06-24 13:57 | 0 | 0 |  | Dear Bee Tin, Here is the drawing: [Image] Best, Nicolas
INC260367 | Workgroup | Incident | AW: akp_inq_20220623_1715_pmontin | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-24 09:18 | 2022-06-24 09:45 | 0 | 0 |  | Hi Carlo, yes, you need 2 intermediated ARIN to connecto to the AMFD. Best, Nicolas
INC260355 | Workgroup | Incident | MR5-4D30F20+ASIPD1L3+DANC01-2L2A01 / Trevor | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-24 07:45 | 2022-07-01 11:32 | 0 | 1 |  | Dear colleague; the mounting below is OK, with a max accuracy of 4% at the limit (150C) MG5-4D30 F20 +ASIPD1L3+DANC01-2L2A.0 will work also without the accessory. Best, Nicolas
INC259965 | Workgroup | Incident | Enquiry Nuclear | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-20 15:09 | 2022-06-20 16:26 | 0 | 0 |  | Hi Kerstin, here are the corresponding codes: ZT402CHM TD2K8M 0-45 Grad = ZT402CHM280 (connection is to check as there is nothing in your code I choose "without connection") ZT403CHM TD2K2M 40-120 Grad = ZT403CHM220 (connection is to check as there is nothing in your code I choose "without connection") also on both the capillary will be: Copper capillary and bulb w/ stainless steel sheath (which match the TD2 from the RT series) Best, Nicolas
INC259964 | Workgroup | Incident | Enquiry Rohrdruckmittler 1690 DS | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-20 15:06 | 2022-06-20 16:00 | 0 | 0 |  | Hi Nathalie, the mounting below was validated by Jorg, and on the website you can find the details of the 1693, LRS9 and PBMN. here are the details of the remaining parts: PBMN-24B26RA0440320100 +0402+1693+LRS9+9007/0150 0402: HIGH-TEMPERATURE-ADAPTER FOR MEDIUM TEMPERATURE UP TO 300 DEGREES C 9007/0150: TEMPERATURE OF MEDIUM 150 DEGREES C Best, Nicolas
INC259925 | Workgroup | Incident | TR: IC price and deliver time | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-20 11:39 | 2022-06-20 15:43 | 0 | 0 |  | Nathalie, only the diaphragm of the DT3 (the biggest standard we have would be suitable). you can also find it on the D435 or the D820 (DN100 or 80) Also if the hygienic connection is important, the PFMH could be the best solution. usually, the PFMH is more the type of instruments used on hygienic applications, more than a heavy duty switch with high cutting power. Best, Nicolas
INC259922 | Workgroup | Incident | TR: MV gauges | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-20 11:27 | 2022-06-20 15:29 | 0 | 0 |  | Hi Nathalie, As MV are test gauges, the environment it's used in is mostly indoors but otherwise, you have the material, the IP, the temperatures etc. on the datasheet as usual: [Image] [Image] Best, Nicolas
INC259759 | Workgroup | Incident | Demande de faisabilité TBI client JST | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-16 17:45 | 2022-06-21 12:07 | 0 | 0 |  | Salut Hervé, La norme EN pour les TBI ne demande ni test en brouillard salin ni test en vibration mais j'ai quand même quelques informations pour toi. selon Stefan Wanner, des tests en brouillard salin ont été fait pendant la phase de développement des TBI qui l'on passé avec succès. étant donné que les boitiers de TBI sont dans la même matière que pour nos MEX5, on peut aussi en déduire qu'ils tiennent au brouillard salin comme nos manomètres (respect de l'EN837) l'option boitier 316L permettra d'accroitre encore cette résistance L'IP est donné sur la fiche technique. les tests en vibrations seront fait selon l'EN837 bientôt, les premiers tests sont concluant (on pourra surement bientôt donner un rapport (Octobre)). Best, Nicolas
INC259707 | Workgroup | Incident | TR: PO#2763.pdf | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-16 10:18 | 2022-06-16 14:15 | 0 | 0 |  | Salut Céline, Aucun soucis a faire, je pense que c'est juste la façon de codifier le produit avec 0...120C au lieu de 20T qui pose problème: [Image] Best, Nicolas
INC259694 | Workgroup | Incident | TR: Tube stub 1000 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-16 08:30 | 2022-06-16 09:03 | 0 | 0 |  | Hi Nathalie, right now it's not possible, we don't have test benches to go over 100bar with tube connection. Also, even with the nut and ferrule mounted I think it's too high for the customer to use. Best, Nicolas
INC259651 | Workgroup | Incident | VS: Pressure Transmitter - in high pressure system (tetra pak) | Archive > BHFR Technical Support | Middle |  |  | Closed | Sejersen, Solveig [sds] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-15 14:33 | 2022-06-15 15:08 | 0 | 0 |  | Hi Solveig, the pressure a clamp can withstand is linked to the collar and the standard one stop at 40bar. but with a special high pressure collar, you could go higher. I know that we offered a couple and I'm sure we could replace this but I don't know yet how quality will want to handle this. we have a couple MEX5 60bar with DANC open, so it'll be OK. Please send your code creation request to the database or the price request to Nathalie, and I'll see with them how to handle it. Best, Nicolas
INC259626 | Workgroup | Incident | Subsea Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-15 12:00 | 2022-06-15 15:51 | 0 | 0 |  | Dear Bee Tin, this kind of gauges exist but is being phased out, what I suggest is you ask your customer if the standard would fit (10271985 MAH3-G50.B27/0764), the only difference is it's a simple bar scale (please validate the product with them, enclosed the datasheet). And you should also ask if they want to order more as the product will be stopped to have stock (last buy order). Then once they agree on the standard version and the qty, plase send an SPR to Nathalie and we will be able to decide to do it or not and confirm you a price. Best, Nicolas
INC259615 | Workgroup | Incident | PMI Certificate | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-15 11:18 | 2022-06-15 11:33 | 0 | 0 |  | Hi Carlo, this price comes from the time we had to send the units to a lab for testing. now we have the machine to do it on site, we can test and deliver a report (on most parts). when you need it please ask for the price to Nathalie with the list of parts that needs to be tested. Best, Nicolas
INC259522 | Workgroup | Incident | LEAD 30126084 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-14 14:15 | 2022-06-14 14:34 | 0 | 0 |  | Salut Hervé, MEP5 600bar: avec RAS1: 720bar avec RAS2: 900bar Best, Nicolas
INC259467 | Workgroup | Incident | TR: Price request for MX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-14 09:06 | 2022-06-14 10:02 | 0 | 0 |  | Answered via email (14/06/2022)
INC259444 | Workgroup | Incident | TR: rfq avi06444 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-13 17:18 | 2022-06-14 12:08 | 0 | 0 |  | Vu par téléphone, voir avec le client pour les détails de la connexion (Parker ou Hamlet par exemple peuvent nous fournir des connections EC-79) si 600bar tube, alors SPR, et possible uniquement avec l montage de la bague et de l'écrou. si le manomètre doit être qualifié, alors voir R&D/qualité pour une demande de qualification. Best, Nicolas
INC259420 | Workgroup | Incident | WG: Anfrage 97880WT22 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-13 14:24 | 2022-06-14 09:58 | 1 | 0 |  | Dear Colleagues, this will not be for me, I transfer the request to sales administration. Best, Nicolas
INC259403 | Workgroup | Incident | GSK-Pharma - Véronique Etien | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-13 13:24 | 2022-06-14 09:56 | 1 | 0 |  | Hi Rudi, the Mounting is OK, I transfer to Nathalie for the price: 11105441 DRCG_101.211.133.055/3272 + 10268029 DANC02-2L2B + LRS1 Best, Nicolas
INC259368 | Workgroup | Incident | TR: BAIT - TECHNOALPIN - CTX 0 - 20 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-13 10:00 | 2022-06-27 15:15 | 0 | 1 |  | Hi Francesco, I understand for the connectors, so I close the request here. they shold keep the references for validation even if for next year. concerning the 1000units, we requested a lead time this morning and I'll go and check on the line this afternoon. Best, Nicolas
INC259332 | Workgroup | Incident | TR: Technical support for Bourdon Pressure- Gauge MIX7D60B20, D04009 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 16:57 | 2022-06-13 11:08 | 0 | 0 |  | Bonjour Nathalie, c'est du Meca donc pour ESP. Best, Nicolas
INC259323 | Workgroup | Incident | TR: Price request for MX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 15:24 | 2022-06-10 15:52 | 0 | 0 |  | Nathalie, oui, le "19" dans le code du D438 est pour 3" 150, ce n'est pas standard donc pas au catalogue. Mais c'est bien possible. Best, Nicolas
INC259272 | Workgroup | Incident | MCF5-460.B10/0693/0651/9941+AORPB501/SETP+ARIN2306/ Neste Singapore | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 09:57 | 2022-06-10 15:29 | 0 | 0 |  | Dear Bee Tin, it would not be possible to have NACE capsules according to quality dept. potentially we could consider having a NACE certificate for the connection but it would require a specific treatment of the order and finding the supplier. Best, Nicolas
INC259261 | Workgroup | Incident | TR: TBI cal 5 points charge | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 09:00 | 2022-06-10 10:42 | 0 | 0 |  | Nathalie, Voici la limite en temperature pour les TBIs: [Image] Best, Nicolas
INC259259 | Workgroup | Incident | TR: INQUIRY OF Level Gauge HAENNI 439(‪8391062‬) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 08:57 | 2022-06-10 10:55 | 0 | 0 |  | Nathalie, c'est celui -ci: 10323207 Best, Nicolas
INC259241 | Workgroup | Incident | M1256D30B72 BOURDON HAENNI | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-10 03:54 | 2022-06-10 10:49 | 0 | 0 |  | Hi Bee Tin, the M12 is replaced by the MR5 MR56D30B72 Best, Nicolas
INC259226 | Workgroup | Incident | test étanchéité Mano sur séparateur // cuenot | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 16:30 | 2022-06-09 17:50 | 0 | 0 |  | Hervé, oui, c'est OK, on testera sur le produit fini l'etancheité client (enceinte interne du separateur en contact avec le fluide process). a 10-7 comme demandé (validation avec Christophe Charpentier OK) Best, Nicolas
INC259218 | Workgroup | Incident | Product Code? | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 15:48 | 2022-06-14 10:04 | 0 | 0 |  | Hello Sasha, here is the code of the product you need: 10320212 DRD 100/111.911/0..10 BAR Best, Nicolas
INC259214 | Workgroup | Incident | TR: Bourdon-Haenni Flex Program | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 15:27 | 2022-06-09 16:38 | 0 | 0 |  | Dear Nathalie, the software and updates can be found on the page of the flexprogrammer: FlexProgrammer 9701 ¦ Accessories process sensors ¦ Baumer Switzerland "https://www.baumer.com/ch/en/product-overview/accessories/accessories-process-sensors/interfaces/flexprogrammer-9701/p/38174" it's currently working on my computer under Windaows 10 Best, Nicolas
INC259215 | Workgroup | Incident | TR: RFQ 46103 | Archive > BHFR Technical Support | Middle |  |  | Closed | Birrien, Nicolas [bin] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 15:27 | 2022-06-09 16:41 | 0 | 0 |  | Nathalie, tu peux envoyer a ESP, c'est plus du meca. Best, Nicolas
INC259212 | Workgroup | Incident | ARID | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 15:18 | 2022-06-09 17:12 | 0 | 0 |  | Hervé, en faisant le parallèle avec un autre produit avec les mêmes embouts et soudures, le MAX sera de 400bar. avec un seul raccord, on pourrait prendre un DRCG100 qui a une sortie 4...20mA Best, Nicolas
INC259154 | Workgroup | Incident | short description, Seal media temperature about 235°C, type 1693 with pressure transmitter 4-20 mA signal | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-09 11:05 | 2022-06-10 14:31 | 0 | 0 |  | local solution provided by Jorg
INC259025 | Workgroup | Incident | Temp. Application from Bilfinger | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-08 14:21 | 2022-06-09 17:54 | 0 | 0 |  | Sascha, we don't have this weird mounting but I think we can offer better: Order code: TBI100-2N2.161.13T/0752 TBI old Code / SAP code: Nominal size: DN 100 100 Case AISI 304 (1.4301), stainless steel: AISI 304 (1.4301), stainless steel 2 Process connection: DN38 ISO (10318762) N Immersion tube outlet: Center back 2 Indicator: EN 13190 (Standard) 1 Immersion tube diameter: 8 mm AISI 316Ti (1.4571), stainless steel 6 Immersion tube length L1: 60 mm 1 Temperature range: Single scale inC 0 ... 160C 13T OPTIONS Window: Window acrylglass (polymethacrylate/plexi) 0752 This keeps all the technical features needed by the customer in a simple, less expensive and easily replacable solution. the Clamp connection will be welded on the TBI tube for the best tightness. I send the cost to Nathalie, she'll follow up with you for the price. Best, Nicolas
INC258915 | Workgroup | Incident | Cannot find product | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-07 16:48 | 2022-06-07 17:07 | 0 | 0 |  | Hi Sascha, it's the Old codification, here is what you need (but the code needs to be created): Order code: RPPN-7W6.201/SETP RPPN PP old Code / SAP code: EURO Sensing element and diaphragm: Type: PPN - Medium range -1 ... 600 bar 7 185.00 Microswitch: Microswitch 2 hermetically changeover switches W 186.70 Hydraulic connection: 1/2" NPT male 6 0.00 Pressure range: Pressure range (bar) -1 ... 2.5 bar 201 95.50 OPTIONS Accuracy / Setting: Set point adjustment SETP 16.30 Purchasing-price in EURO 483.50 Best, Nicolas
INC258902 | Workgroup | Incident | TR: Enquiry about 3 components | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-07 15:24 | 2022-06-07 15:40 | 0 | 0 |  | Nathalie, C'est ok , ci dessous le code: TED633B26R + D82110L12CEB2101 Best, Nicolas
INC258873 | Workgroup | Incident | TR: BAIT ¦ UL certification on BHFR products | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-07 12:21 | 2022-06-07 14:42 | 0 | 0 |  | Nathalie, ci-joint notre certificat de conformité UL, il y a la liste de nos produits certifiés: Pxxx CTx CPx EF6 TED Best, Nicolas
INC258867 | Workgroup | Incident | WG: R100479-THERMOMETER WITH THERMOWELL | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-07 11:33 | 2022-06-07 14:23 | 0 | 0 |  | Answered on the phone, we never get feedback from KOC and they only work with BTIN. this type of request will be received by many distributors and the general specs are not included (and we know there is specific design) we will let the Indians deal directly on this. Best, Nicolas
INC258826 | Workgroup | Incident | TR: MCX5/ Kumpanart | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-07 09:12 | 2022-06-07 10:49 | 0 | 0 |  | Nathalie, on va faire l'offre en kPa / mmH2O. par contre si la creation est de toute facon a payer, ils vont peut etre insister pour les mmAq (seul le nom est different, sinon ce sont des mmH2O). Best, Nicolas.
INC258709 | Workgroup | Incident | TR: Price request for MX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-03 12:03 | 2022-06-07 16:42 | 0 | 0 |  | Nathalie, voici quoi proposer: MX7-P60.D18J + AKPLB06607 + D438L100N20219G0L0A066 (66 = 2 trous 1/4NPT avec bouchons). le CCR2 du D438 est: 242.74 Best, Nicolas
INC258698 | Workgroup | Incident | TR: AL E&C - Manomètres différentiel pour service oxygène | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-03 11:00 | 2022-06-03 11:09 | 0 | 0 |  | Salut Hervé, j'avais relancé fantinelli hier, j'en profite aujourd'hui pour leurs faire une relance en poussant avec le commentaire du client. je ferme cette demande car j'ai continué sur l'original et je préfère ne pas faire de doublons. Best, Nicolas
INC258692 | Workgroup | Incident | N° 20462691 / ACMAR / TED-6.3.7.B76.R/0592_9010.D055-2L16-2L140.09 / poste 10 - Your Support Ticket was closed | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-03 10:33 | 2022-06-09 16:14 | 0 | 1 |  | Hervé, la demande est refusée par la R&D car cela peut poser potentiellement un problème en température. TED-6.3G.3.B76.R/0410_9010 TED-6.3G.3.B76.R/9010 le minimum pour du G1/2 flush sera du -1...9bar. Best, Nicolas
INC258673 | Workgroup | Incident | TR: max pointer | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-03 09:00 | 2022-06-03 16:23 | 0 | 0 |  | Nathalie, Vu avec Jim, pour le MEX5 ce sera un MEX5D61X1E/9712 (-30"...500psi / F (ammonia)) CCR2: 25 Best, Nicolas
INC258617 | Workgroup | Incident | TR: product confirmation | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-02 15:00 | 2022-06-02 15:40 | 0 | 0 |  | Nathalie, voici notre réponse: Those products are very old and we can't trace it anymore (2 changes of system since 1998). but it seems to be simple gauges for the most part and you can find enclosed our datasheet for similar products (MEX3/5) and the link for our website below. the accessory I see in a couple picture could be a pressure dampener or limitor and for this we also have solutions (AORP ARA). option on the gauge itself could also be a possibility for resisting higher pressure (RAS options) or pulsations (restrictor screw). Please come back to us with more details about your need and we will be happy to help you find the right solution. Best, Nicolas
INC258443 | Workgroup | Incident | TR: rfq TBHI | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-01 12:18 | 2022-06-02 10:19 | 1 | 0 |  | Céline, Oui, Charles a demandé a S. Wanner et Jorg de l'enlever sur le site et il corrige le config. Céline, Il y a une erreur sur le site web, la plage 800C est uniquement dispo sur les TSF. ces code ne sont donc pas possibles. Désolé Best, Nicolas
INC258432 | Workgroup | Incident | RE: MX7+AKPL_D425 | Archive > BHFR Technical Support | Middle |  |  | Closed | Tiong, Bee Tin [tib] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-01 11:18 | 2022-06-01 16:41 | 1 | 0 |  | Dear Bee Tin, the technical solution is as below, but it needs to go through the SPR process as the bottom flange needs to be design and bought for 1 unit, the intermediate flushing ring will be special and the compatibility of the whole mounting needs to be validated by R&D. considering this, please send your SPR to Nathalie for approval by management. Best, Nicolas
INC258417 | Workgroup | Incident | MEX3+DT2 - LRS1 - BH1 - Janssen Pharma (21106313) | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-06-01 10:12 | 2022-06-01 12:10 | 0 | 0 |  | Hi Rudi, here are the qty of filling we have in your mounting: BH1: 63g LRS1: 15g best, Nicolas
INC258240 | Workgroup | Incident | VS: Price an delivery MT7-833.X16J.AKPL.DT3.AKPL.ARIN.XMFD.KIT | Archive > BHFR Technical Support | Middle |  |  | Closed | Andersen, Jens Grønbech [ajen] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-30 11:03 | 2022-06-01 11:19 | 1 | 0 |  | Hi Jens, here is the code and the details, for the price, please ask Nathalie Billault, it's not up to date. 11198166 MT7-833.X16J.AKPL.DT3.AKPL.ARIN.XMFD.KIT MT7 DIFFERENTIAL PRESSURE PARTITION OF SAFETY GAUGE BELLOWS ALL STAINL. STEEL DIFFERENTIAL PRESSURE, DN=150MM FILLABLE CASE 8 BOTTOM CONNECTION, PIPE CLAMP 2" P, CASE AND BEZEL IN 316L 3 CONNECTION G 1/2 3 LIQUID FILLED BH3 X16 SENSING ELEMENT 0+1,6 BAR J STATIC PRESSURE 0+16 BAR 9712 PRINTER DIAL I1097 1ere scale 0+100% 2ème scale 0+3000 m3/h square root 0 m3/h = 0% = 0 mbar 0+3000 m3/h = 100% = 1650 mbar 9711 = MADE IN FRANCE 9711 TECHNICAL INSCRIPTION MADE IN FRANCE ON DIAL 0678 MICRO ADJUSTABLE POINTER 0078 ATEX 94/9/CE CONFORMITY FOR USE CE II 2 GDc T6 - I M2c TRIPLEX SAFETY GLASS 9941 STAINLESS LABEL AND WIRE AKPL CAPILLARYB STAINLESS STEEL DIA 2.5 + STAINLESS STEEL PROTECTION F5 ORIENTABLE SCREW G 1/2 3 CONNECTION G 1/2" 01 LENGTH 1M DT THREADED DIAPHRAGM SEAL 3 LOW PRESSURE RANGE (D3XX) 2 TOP HOUSING INOX 316L (SS) L G 1/2 FEMALE 6 HASTELLOY C 276 (HC) DIAPHRAGM 0 DIAPHRAGM WITHOUT COATING 0 WITHOUT CLEANING RING 6 BOTTOM HOUSING HASTELLOY C 276 (HC) 6 THREAD 1/2 NPT 0 NO COATING 01 LRS1 (FOR TEMP.-15..150C)VASELINE 1981 BOLTS AND NUTS IN SUPER DUPLEX 1.4410 AKPL CAPILLARY B STAINLESS STEEL DIA 2.5 + STAINLESS STEEL PROTECTION 0N TAPPING 1/2 NPT 3 CONNECTION G 1/2 01 LENGTH 1M ARIN THREAD ADAPTOR 2 STAINLESS STEEL 316L 6 CONNECTION 1/2 NPT 0L TAPPING G 1/2 XMFD:R5AAHA-N4N4-A001 TYPE: MANIFOLD CODE/DRAWING: R5AAHA-N4N4-A001 MATERIAL: ALLOY C276 TYPE: 5 WAYS INSTRUMENT CONNECTION: 1/2NPT FEMALE PROCESS CONNECTION: 1/2NPT FEMALE isolate valves lockable with handwheel équalize valve lockable with handwheel test/drain: anti-tamper with handwheel gland nut and lock nut in Alloy C-276 XKIT:AKM-RPS KIT AKM-RPS FOR MANIFOLD EXECUTION:MOUTING ON A 2"PIPE 1 MOUTING BRACKET + 1 CALIPER ATTACHMENT Best, Nicolas
INC257929 | Workgroup | Incident | TR: BAIT ¦ Mi2 ¦ TED flush 4 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-23 16:57 | 2022-05-24 09:00 | 0 | 0 |  | Nathalie, On peut faire, ce sera créé hors config. TED-6.3F.3.B19.R Best, Nicolas
INC257890 | Workgroup | Incident | demande yoko | Archive > BHFR Technical Support | Middle |  |  | Closed | Thieffry, Véronique [tve] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-23 14:21 | 2022-05-23 14:59 | 0 | 0 |  | Veronique, il faut revenir vers eux avec ces demandes: Le code complet de l'EJ Yoko Confirmer la demande de LRS1 FDA (pas dans la REQ) Mesure Absolue ? PN/DN au lieux de class and rating ? (EN au lieux de ANSI) Matière HC276 (pas HC22) Sinon voici un exemple de codif qui pourrait peut être convenir et les datasheets qui peuvent les intéresser. D912L100X9602L24G0A000 La solution devra necessairement passer par notre service R&D pour validation du montage. Best, Nicolas
INC257872 | Workgroup | Incident | Pricing | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-23 12:39 | 2022-05-23 15:24 | 0 | 0 |  | Hi Sascha, here is what I would suggest: PBMN-24B24RA11403201000 + D820-10L.42DH.A20.09 the accuracy must still be confirmed by R&D. Please confirm if the mounting is OK for you ? Best, Nicolas
INC257784 | Workgroup | Incident | Double scale ??? | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-20 15:03 | 2022-05-20 15:41 | 0 | 0 |  | Hi Kees, it doesn't exist in SAP (a bit too old I think) but here is how to codify it: (Photo 11) What scale is this? MEX5-D61.X26/9712 (kg/cm / MPa black) (Photo 12) What scale is this? MEX5-D61.X22/9712 (kg/cm / MPa black) the Logo will change to Bourdon and a dial will need to be validated. the 11 looks like a 1/2G connection to me. Best, Nicolas
INC257776 | Workgroup | Incident | WG: MEX5 & MIX7 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-20 14:33 | 2022-05-20 15:31 | 0 | 0 |  | Hi Carlo, I have these information and we do slightly different: Case/window gasket: PVC/Nitrile grade 4 Dial: Duralumin white primed Paint for the dial: LUS 12 (for MIMAKI printers) Best, Nicolas
INC257716 | Workgroup | Incident | Product SAP#10346564 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-20 10:33 | 2022-06-01 16:22 | 0 | 1 |  | Hi Sasha, after reviewing your request and as the mounting was extremely special, we would like to offer you a direct replacement solution that should be less expensive and easier for the future replacement: DROG 101 211 133 535/9710/9700 + AKPL2F5301 + DANC012L2A/9007/0080 The différences are: Not S3 304 Housing -1...5bar Not ATEX No span adjustment 1m cappilary Mounting will be simpler with only 1 instrument, capillary and Clamp connection. Please come back with customer's feedback. Best, Nicolas
INC257668 | Workgroup | Incident | TR: BAIT ¦ TECHNOALPIN ¦ CTX 0 - 20 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-19 17:12 | 2022-06-02 11:11 | 0 | 1 |  | Nathalie, Voici le document. Best, Nicolas
INC257667 | Workgroup | Incident | TR: BAIT ¦ TERMOGAMMA Impianti e Servizi Srl / Concorezzo (MB) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-19 17:12 | 2022-05-19 17:56 | 0 | 0 |  | Vue avec Nathalie, répondu par Carlo.
INC257605 | Workgroup | Incident | TR: BAIT ¦ TECHNOALPIN ¦ CTX 0 - 20 bar | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-19 10:42 | 2022-05-19 15:00 | 0 | 0 |  | Hi Nathalie, it will not be added to the datasheet. this is a specific product we did for the customer application (outside of the EN837-1 standard ranges). the datasheet is public and should contain only the standard possibilities we want to show. Best, Nicolas
INC257595 | Workgroup | Incident | DT3 Flasque inférieur PTFE | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-19 10:06 | 2022-05-23 16:03 | 0 | 0 |  | Hervé, je viens de faire la verif finale avec CCH et il demande un changement sur le 160mbar, pas possible Halar, on passe en coating PTFE: MCX5-D60.N09/9941+DT3-2N21.0.F60.01 +Q003 +Q008 Echelle 0...160mBar / Membrane revêtue PTFE MEX5-D60.B20/9941+DT2-2N24.0.F60.01 +Q003 +Q008 Echelle 06Bar / Membrane revêtue Halar MEX5-D60.B77/9941+DT2-2N24.0.F60.01 +Q003 +Q008 Echelle -15Bar / Membrane revêtue Halar Best, Nicolas
INC257550 | Workgroup | Incident | TR: SATHERM Anfrage Nr. 202205573 - SPR for Y933 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-18 17:36 | 2022-05-19 11:10 | 0 | 0 |  | Nathalie, la SPR est refusée. Désolé. Best, Nicolas
INC257524 | Workgroup | Incident | Special dials | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-18 13:51 | 2022-05-19 10:40 | 1 | 0 |  | Taken care of by database
INC257509 | Workgroup | Incident | WG: MEP5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-18 11:21 | 2022-05-23 09:11 | 1 | 0 |  | Hi Carlo, I can confirm you the new IC for the price of the connection C014, it's 17. Best, Nicolas
INC257475 | Workgroup | Incident | Barflex // air liquide | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-18 09:00 | 2022-05-19 10:10 | 1 | 0 |  | Salut Hervé, pour cette demande, ce n'est pas possible de répondre. je n'ai rien dans les tables de compatibilités et le NO est instable et réactif avec l'air et l'eau. selon les conditions de température et d'hygrométrie, la compatibilité sera potentiellement différente. après en règle générale, l'EPDM et le Viton sont plus résistants, tu peux commencer par proposer ca. et comme d'habitude, c'est le client qui est responsable de son process. best, Nicolas
INC257374 | Workgroup | Incident | TR: Fw: Anfrage Nr. 202205573 DRINGEND | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-17 09:18 | 2022-05-17 10:06 | 0 | 0 |  | Nathalie, Comme discuté, étant donné que l'on bosse sur ce sujet qui est assez brulant en ce moment, je suggère de faire une SPR a envoyer a Franck Ferarese pour avoir son avis. Best, Nicolas
INC257330 | Workgroup | Incident | need Article Code and Price | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-16 18:03 | 2022-05-17 09:09 | 0 | 0 |  | Dear Sascha, I received the same one from singapore, but with a seceond picture (attached), here is what I answered: There was some changes in the designation and a mistake on the product name and the microswitch type N is no longer available. But its a good thing you came back, because the switching point was also too low for the range so its better to redefine the product. I suggest to replace with the one below (to be confirmed with customer): Order code: RDN-8A3.224 RDN old Code / SAP code: Sensing element and diaphragm: Type: DPN - medium pressure 0.05 ... 30 bar 8 Microswitch: Microswitch 1 standard changeover switch A Hydraulic connection: G1/2" 3 Pressure range: Pressure range (bar) 0.15 ... 1 bar 224 They can then have a switch where they need and the min setting for the hysteresis will be 0.22bar. Example: 0.48bar falling 0.7bar rising Best, Nicolas
INC257300 | Workgroup | Incident | Scale 0,005 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-16 14:06 | 2022-05-17 15:04 | 0 | 0 |  | Hi Sasha, graduation every 5mbar is impossible, it would be unreadable. here is what we can do with the Bar Bourdon (MEX5XXXA74) [Image] Best, Nicolas
INC257293 | Workgroup | Incident | TR: highest medium temperature for MEX5-A60.E22/0655 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-16 13:30 | 2022-05-17 09:01 | 1 | 0 |  | Nathalie, -40...200C, pas de changement avec le dashpot. Best, Nicolas
INC257269 | Workgroup | Incident | TR: Gasket on PBSN#11147422 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-16 10:51 | 2022-05-17 10:26 | 1 | 0 |  | Nathalie, je leurs conseille d'utiliser le code 11164372 tel que sur la spec à jour ci-jointe. ainsi les 2 joints (de cellule et de raccord) seront en EPDM. Best, Nicolas
INC257181 | Workgroup | Incident | TR: Fw: Anfrage Nr. 202205573 DRINGEND | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-13 13:51 | 2022-05-13 14:21 | 0 | 0 |  | Nathalie, je peux prendre ces informations sur les photos et commande: 04bar R 420mA DN38 clamp E733 (Y933). DIN43650 Gold plated diaphragm le doute est sur la connexion, pour moi c'est un clamp DN38 mais le code 38 sur la commande peut aussi indiquer un SMS FL DN38 + nut mais ca ne correspondrait pas a la photo. je proposerais un Y93318/38B19R/gold plated mais le produit est fermé et la demande devra passer par un SPR. Best, Nicolas
INC257126 | Workgroup | Incident | TR: MEX35 + D055 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-13 09:09 | 2022-05-13 09:19 | 0 | 0 |  | Nathalie, le MX35 + D055 en 1"NPT est possible (confirmer qu'on ajoute pas de contrainte dans le code complet). pas de 0...4bar par contre. [Image] Best, Nicolas
INC257090 | Workgroup | Incident | TR: BAIT ¦ Temperature switch info | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-12 16:30 | 2022-05-13 09:16 | 0 | 0 |  | Nathalie, On a pas de manual reset en antidéflagrant, je suis désolé mais on a rien pour remplacer. Best, Nicolas
INC256957 | Workgroup | Incident | TR: rf vendor list | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-11 15:24 | 2022-05-11 16:01 | 0 | 0 |  | Celine, the RF type Angel needs got different digit depending on the class, so here is what sylvain is correcting: MEP5-P31.S31/Q001_Q003_Q008_9941_0678/D820-AAL.2237.R2001 MEP5-P31.S31/Q001_Q003_Q008_9941_0678/D820-AAL.2237.R2001 MEP5-P31.S31/Q001_Q003_Q008_9941_0678/D820-AAL.2237.R2001 MEP5-P31.S34/Q001_Q003_Q008_0678_9941/D820-A9L.2247.R2101 MEP5-P31.S34/Q001_Q003_Q008_0678_9941/D820-A9L.2247.R2101 MEP5-P31.S29/Q001_Q003_Q008_9941_0678/D820-AAL.2227.G2001 MEP5-P31.S29/Q001_Q003_Q008_9941_0678/D820-AAL.2227.G2001 MEP5-P31.S25/Q001_Q003_Q008_0678_9941/D820-AAL.2217.G2001 Best, Nicolas
INC256888 | Workgroup | Incident | TR: UL certs on Y913s | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-11 08:36 | 2022-05-11 09:21 | 0 | 0 |  | Hi Nathalie, we don't have the UL on the Y91, but we have it on the PBSN (see certificate enclosed). Best, Nicolas
INC256808 | Workgroup | Incident | TR: MEINDO - 0060K00000eeSbHQAU | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 12:00 | 2022-05-10 15:28 | 0 | 0 |  | Hi Nathalie, for this mounting with the option 9030, it is not point welding but full welding (which is the right choice) then, each connection will be welded, so 3 weldings (in green below) MPE6-D62.H25/9941+AORP0383/SETP+ARIN2606+D418-210.0N2.025.4Q0.M0A000 best, Nicolas
INC256765 | Workgroup | Incident | TR: Inquiry_k 202205061526 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 09:36 | 2022-05-10 11:11 | 0 | 0 |  | Bonjour Nathalie, Nous n'avons pas prévu de redemander l'ATEX sur le Barflex, ca demanderait un changement de conception. il n'y a pas encore de remplacement de prévu on peut juste proposer la version non-ATEX. Best, Nicolas
INC256763 | Workgroup | Incident | TR: SPR zu Anfrage Drucksensor, unser Projekt [O85] - 30122989 - update for 0...10V version | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 09:30 | 2022-05-10 11:06 | 0 | 0 |  | Nathalie, Voici les 2 CCR2: 10265705 CTX-2.B.3.B22.A 19.6 10265867 CTX-3.B.3.B22.A 19.29 Best, Nicolas
INC256758 | Workgroup | Incident | TR: rfi about pressure gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 09:18 | 2022-05-10 10:43 | 0 | 0 |  | Hi Angel, The gauge can be calibrated horizontally on request with the option 0217. concerning the other mountings, the calibration will always be done with the connection at the bottom, but we guarantee the accuracy will be respected as long as the gauge is vertical. Best, Nicolas
INC256747 | Workgroup | Incident | TR: MEINDO - 0060K00000eeSbHQAU | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 08:39 | 2022-05-10 10:25 | 0 | 0 |  | Nathalie, c'est possible, par contre en class 1500, il faut passer en forgé: MPE6-D62.H25/9941+AORP0383/SETP+ARIN2606+D418-210.0N2.025.4Q0.M0A000 Best, Nicolas
INC256738 | Workgroup | Incident | WG: AKP_INQ20220509_2138 - MERRID Project | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-10 08:24 | 2022-05-10 13:54 | 0 | 0 |  | Hi Carlo, Safety glass is not really an issue. but ATEX is as it states temperature levels, we have to respect the datasheet on this (-40 process; -20 ambiant). Best, Nicolas
INC256703 | Workgroup | Incident | TR: stock DANC | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-09 15:33 | 2022-05-10 09:57 | 0 | 0 |  | Celine, des codes existent mais on a pas de stock. Aussi le montage Clamp en NPT est vivement déconseillé, on en a fait extrêmement ces dernières années, et au besoin il faudra valider le montage avec la R&D Best, Nicolas
INC256701 | Workgroup | Incident | Turndown CTX-3.B.3.B15.0 en 500 mbar | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-09 15:18 | 2022-05-10 09:51 | 0 | 0 |  | Bonjour Jacques, Nous ne pourrons pas faire un CTX 500mbar en l'état actuel mais tu peux passer en Y913 sans problème. en cas de très grandes qté, il faudrait faire un SPR et compter du temps pour mener a bien ce projet si il était accepté. Best, Nicolas
INC256663 | Workgroup | Incident | Seals with diaphragm inconel | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-09 13:06 | 2022-05-10 11:44 | 1 | 0 |  | Hi Jorg, the inconel sheets is an issue since very long for Jerome. Christophe is saying there is no problem for doing it technically. the problem is that if we need to do it, we need to buy a big qty of material, which is really expensive. so we need to validate an SPR if you have some qty. Best, Nicolas
INC256597 | Workgroup | Incident | TR: E933 complete type | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-09 09:09 | 2022-05-13 09:20 | 1 | 0 |  | Bonjour, Ouverture du code 11084957 Y93318B77R approuvée R&D et approuvée Franck Ferarese. Best, Nicolas
INC256505 | Workgroup | Incident | Séparateur sur TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Bourgeon, Cyrille [bcy] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-06 13:45 | 2022-05-06 16:50 | 0 | 0 |  | Bonjour, après un premier coup d'il, ce sera vraiment spécifique, pouvez vous faire une SPR svp ? Best, Nicolas
INC256497 | Workgroup | Incident | TR: Séparateur sur TED | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-06 12:57 | 2022-05-06 13:55 | 0 | 0 |  | Doublon demande de Cyrille Bourgeon.
INC256332 | Workgroup | Incident | TR: rf vendor list | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-05 09:06 | 2022-05-06 09:14 | 1 | 0 |  | Céline, c'est ok pour 2200psi sur un MEX3 ou 5. la sérigraphie sera a créer. Best, Nicolas
INC256329 | Workgroup | Incident | millipore // Matière du bouchon rouge | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-05 08:51 | 2022-05-05 10:56 | 0 | 0 |  | Salut Hervé, voici les infos sur le Bouchon (matiere PE-LD) [Image] Best, Nicolas
INC256240 | Workgroup | Incident | MTTFD pour PBSN et CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-04 13:03 | 2022-05-04 15:15 | 0 | 0 |  | Bonjour Jacques, Ci-joint le MTTF du PBSN et du CTX. Best, Nicolas
INC256152 | Workgroup | Incident | WG: Transformator Temperatur Messgerät | Archive > BHFR Technical Support | Middle |  |  | Closed | Strüfing, Jörg [jst] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-03 21:03 | 2022-05-04 15:12 | 0 | 0 |  | Information sent.
INC256146 | Workgroup | Incident | Ziemann Complaint MX35 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-03 18:42 | 2022-05-04 09:38 | 0 | 0 |  | Hi Sascha, I don't have all the information on this, but I really think the best option is to send your request to Nathalie and Paul Jeauneau, the know about the claim and Paul will now about the conclusions, tests that were performed etc... Best, Nicolas
INC256080 | Workgroup | Incident | alternative DRO100 SAP Code 10260473 for Dürr 0.....100 mbar | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-05-03 12:18 | 2022-05-06 09:16 | 1 | 0 |  | Hi Sascha, we can do all below with an MCX5 except the filling. the option numbers used might be slightly different. in 100mbar, anyway the issue would be the same on Dxx and MCX Best, Nicolas
INC255833 | Workgroup | Incident | Cadran de SAP code 11046137 - Electrabel (50148724) - RFQ 6000583518 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-29 11:30 | 2022-06-28 15:41 | 0 | 1 |  | Old request.
INC255756 | Workgroup | Incident | TR: DPCE | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-28 15:24 | 2022-04-28 16:34 | 0 | 0 |  | Oui, bien préciser par contre que c'est un diamètre 100, il y a des infos contradictoires dans leur demande. Best, Nicolas
INC255718 | Workgroup | Incident | WG: RE: Quotation Request 182448-1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-28 13:15 | 2022-04-28 14:28 | 0 | 0 |  | Hi Carlo, I cannot give pricing, but I asked for the consultation to purchasing. they'll answer to Nathalie and myself and Nathalie will give you a price. Best, Nicolas
INC255648 | Workgroup | Incident | TR: Prix-IC ? - Inovyn Manufact. B. (55181808) - RFQ 6500028740-0292-IHL - MEX5-D30.B74.DT3-2L22.0.230.01/0751_0078 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-27 17:27 | 2022-04-27 17:59 | 0 | 0 |  | Oui, pas de liner PTFE adhésif en vide et pression, il ne réagi pas bien au vide. on peut proposer du Halar (code 4) ou un coating PTFE mais plus fin (20microns) code 1. Best, Nicolas
INC255591 | Workgroup | Incident | TR: SPN S.A.: PO 4567525581 Vendor 0100007758 - 1946337 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-27 12:27 | 2022-04-27 14:01 | 0 | 0 |  | Nathalie, ci-dessous le code: 10296214 Best, Nicolas
INC255467 | Workgroup | Incident | Ineos Aromatics (21108018) - RFQ 6000320935 - HansaFlex GMM 100-400 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-26 13:24 | 2022-04-26 15:31 | 0 | 0 |  | Hi Rudi, of course, it seems to be G1/2 bottom as standard and the 400 seems to be for the range (400bar) as the MIT5 and MMD5 don't go that high, I suggest a MEX5. If you have qty I'm sure we can get an MIT5 400bar from our supplier but as it's special I think it's worth it only if you have more than 100units. best, Nicolas
INC255339 | Workgroup | Incident | DROE100/813.... alternativ MG5 with DO55 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-25 15:00 | 2022-04-28 16:45 | 1 | 0 |  | Hi Sascha, your solution below is validated by R&D: MG5_7F3K.B22_DO55_2L.16 you can go ahead with the P2F price, don't forget the filling. Best, Nicolas
INC255328 | Workgroup | Incident | 11253403 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-25 13:42 | 2022-04-26 09:31 | 1 | 0 |  | Hi Sascha, we do not have the 1935 compliance certificate on MG5. and this product is not meant for use on food processes with a standard connection. Best, Nicolas
INC255292 | Workgroup | Incident | TR: SPR zu Anfrage Drucksensor, unser Projekt [O85] - 30122989 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-25 10:21 | 2022-04-29 15:02 | 1 | 0 |  | Nathalie, voici le prix par Antoine: CCR2 : CTX3B3B22A/9710/C01 = 23.15 best, Nicolas
INC255280 | Workgroup | Incident | Declaration of conformity | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-25 09:45 | 2022-04-26 18:01 | 1 | 0 |  | Hi Sascha, I should have started with this. it's the declaration of conformity to standard EN/ATEX norms: Best, Nicolas
INC255271 | Workgroup | Incident | TR: Material inquiry : Y922 3G B29 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-25 08:57 | 2022-04-26 09:36 | 1 | 0 |  | Nathalie, les codes sont possibles, il faut demander l'ouverture au PM (Franck Ferrarese), ca ne devrait pas poser de problème. Best, Nicolas
INC255210 | Workgroup | Incident | price request | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-22 15:33 | 2022-04-22 16:11 | 0 | 0 |  | Hi Sascha, it's ok technically, but for price and SAP code, you better ask your sales contact (Nathalie I guess). Best, Nicolas
INC255209 | Workgroup | Incident | DANC with option 9330 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-22 15:30 | 2022-05-19 09:09 | 0 | 1 |  | Answered via Nathalie 18/05/2022
INC255176 | Workgroup | Incident | TR: SPR of CTX for LS Tech | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-22 11:15 | 2022-04-26 17:25 | 0 | 0 |  | Nathalie, réponse de la R&D, le CCR2 est de 35 (basé sur 100 pièces). CTX-3.5.3.B25.2 (0...20bar) Best, Nicolas
INC255164 | Workgroup | Incident | WG: Declaration of Conformity / AUREX | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-22 09:57 | 2022-04-27 12:08 | 0 | 1 |  | Hi Carlo, As Discussed, on RP2E/Y the ATEX include temperature limitations that can't be change. But please keep in mind we could still do a statement on gauges and TBI (non-ATEX versions). Best, Nicolas
INC255149 | Workgroup | Incident | TR: INQUIRY OF Level Gauge HAENNI 439(‪8391062‬) | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-22 08:42 | 2022-04-27 14:05 | 0 | 0 |  | Nathalie, Sans reponse du client, je clos le ticket. tu me diras si ils reviennent vers toi avec le code interne au boitier. Best, Nicolas
INC255123 | Workgroup | Incident | WG: Material Rohrfedermanometer 10808523 | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-21 16:30 | 2022-04-22 10:11 | 0 | 0 |  | Bonjour Sascha, I checked on our side and we delivered according to our datasheet and order text (you have the matrial certificate enclose) but th customer is spaking about a PBMH in his request... Datasheet: [Image] Order: [Image] Best, Nicolas
INC255109 | Workgroup | Incident | TBHI 100 Every Angle | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-21 15:18 | 2022-04-21 15:44 | 0 | 0 |  | Hey Carlo, here is an .stp TBHI plan. Best, Nicolas
INC255039 | Workgroup | Incident | Faisabilité TED-6.11.3.B76.R/0410 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-21 10:57 | 2022-04-21 12:08 | 0 | 0 |  | Bonjour Jacques, nous n'avons pas la connexion sur le coté sur le TED (4710) le boitier est fait d'une seule pièce en INOX. il faudrait concevoir un nouveau boitier complet, donc on n'étudiera cette possibilité qu'en cas de grande quantité, sinon le boitier deviendra extrêmement cher. Best, Nicolas
INC254882 | Workgroup | Incident | RE: Fiches descriptives 10263474 et 11251564 | Archive > BHFR Technical Support | Middle |  |  | Closed | Apaloo, Jacques [apj] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-19 17:48 | 2022-04-20 14:39 | 0 | 0 |  | Bonjour Jacques, Voici un comparatif des caractéristiques principales. Apres, il va être compliqué d'aller plus loin sachant que nous avons fait des test différents selon la demande du client à l'époque. Characteristic EF6 E913 Comments Range 06bar 06bar other available for both Overpressure 12bar 9bar Burst pressure 18bar 18bar Global accuracy +/- 0.5% FS +/- 0.4% FS Temperature compensation -1060C -2570C Thermal drift +/-0.02% / C +/-0.04% / C output 420mA 420mA power supply 1232Vdc 1240Vdc Process connection G1/4 A special G1/4 A special Compatibles Material in contact 316L / Ceramic / NBR / FVQM Z2CND 17 12 / A12 03/ NBR Electrical connection Embase HE302 6 cts 10B6P Embase HE 302 6 contacts taille20 disposition 10 -6 Operating temperature -4085C -2570C Best, Nicolas
INC254881 | Workgroup | Incident | RE: arkema | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-19 17:06 | 2022-04-21 10:12 | 0 | 0 |  | Passé en Codif.
INC254594 | Workgroup | Incident | Need detail technical specification against instrument sr. no. 160198001 model TSF | Archive > BHFR Technical Support | Middle |  |  | Closed | Vijayvargia, Anil [via] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-13 13:30 | 2022-04-13 15:15 | 0 | 0 |  | Hi Anil, here is the code, it was ordered by BAUK in 2010. 11044454 TSF160-231.168.14T.50/735J Best, Nicolas
INC254549 | Workgroup | Incident | TR: BAIT - Y923FB18R - Danfoss - MBS4510 060G2423 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-13 09:27 | 2022-04-13 11:04 | 0 | 0 |  | Ouverture du code approuvée par Franck. 10301874 Best, Nicolas
INC254515 | Workgroup | Incident | Verordnung (EG)1935/2004, (EG) 2023/2006 und eine Unbedenklichkeitsbescheinigung | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 17:33 | 2022-04-14 10:28 | 0 | 0 |  | Hi Sascha, I had to check with quality for this. we don't have those certificates but: we have 3A for MEX+DANC (see enclosed) concerning the 2023/2006, I saw other suppliers ensure the traceability via their ISO9001 quality insurance. best, Nicolas
INC254448 | Workgroup | Incident | TR: [EXTERNAL] BGPI - 6220151914 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 11:15 | 2022-04-15 16:29 | 0 | 0 |  | Hervé, c'est OK pour ce séparateur, j'ai transmis le cout a Céline qui pourra faire le prix. D418-2100.N.20.654O0M.0M066.09 Top housing material: Stainless steel (1.4404) - 2 Mounting : # Direct mounting - 1 Outlet position: #Axial outlet - 0 Capillary length: Without capillary direct mounting - 0 Instrument connection: 1/2 NPT female for pressure gauge - N Diaphragm material: #Stainless steel 1,4435 - 2 Diapraghm coating: #Without - 0 Flange standard: EN1759-1 - 6 PN / DN: Class 1500 / 1"/DN 25 -54 SELF (Simple Emboitement Large Femelle) - O Flange face finish: Standard finish - 0 Flange material: Forged NACE compliant 1.4404Stainless steel - M Bottom housing coating : Without - 0 Bolts: ISO (metric) - M Heat Tracing circuit: Without - 0 Drain whole (bottom whole): #1/4 NPT avec bouchon - 6 Vent whole (top whole): #1/4 NPT avec bouchon - 6 Oil Filling: sans remplissage - 00 Best, Nicolas
INC254441 | Workgroup | Incident | TR: DDP URGENT AVF - N/demande DP2204-0447 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 10:57 | 2022-04-26 15:08 | 0 | 1 |  | Céline, ci-joint l'offre de KHW pour ce produit. Best, Nicolas
INC254424 | Workgroup | Incident | TR: Demande de prix 040422AG00 (FY21393) // yokogawa | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 09:51 | 2022-04-12 17:11 | 0 | 0 |  | Doublon
INC254422 | Workgroup | Incident | TR: BAIT - Renova - Y92 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 09:33 | 2022-04-13 11:03 | 0 | 0 |  | L'ouverture du code est approuvée par Franck. Y9233FB16R Best, Nicolas
INC254413 | Workgroup | Incident | TR: RFQ MEX5 with Max Drag Pointer | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-12 08:36 | 2022-04-12 09:11 | 0 | 0 |  | Hi Nathalie, here is the code, if it's a Mix glycerin +water (type BH1) it'll be fine. but pure glycerine if really viscous (type BH2) might be too much for the pointer to drag, you'll loose in accuracy. Order code: MEX5-D51.H19/0059 MEX old Code / SAP code: Nominal size: 100 mm 5 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection D Process connection: 1/4" NPT 5 Oil filling: BH1, glycerin/water, low viscosity 1 Pressure unit: psi (ANSI standard) H Pressure range: 0 ... 60 psi 19 OPTIONS Window: Polycarbonate with max. pointer, zero reset by fixed button P>1.6 bar 0059 Best, Nicolas
INC254383 | Workgroup | Incident | TR: BAIT - Technoalpin - SPR CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-11 17:18 | 2022-04-11 17:43 | 0 | 0 |  | Bonjour Nathalie, hereafter the codification: Y913X5B26R/2160 (the cable output is 1.5m instead of 1m on the CTX) Best, Nicolas
INC254380 | Workgroup | Incident | MEX5 with SPM900 | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-11 16:57 | 2022-04-11 17:22 | 0 | 0 |  | Hi Sascha, you can go ahead and offer the code below, only the filling is different (if it's a problem, we will need an SPR as we never did LRS1 filling on a D030) 11152070 MEX5-D31.B20.D030.LL.02 I suggest you make an SPR if there is a bit of volume. Best, Nicolas
INC254369 | Workgroup | Incident | TR: BAIT - Renova - Y92 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-11 15:57 | 2022-04-11 17:10 | 0 | 0 |  | Hello, With a Y92, you could possibly offer a 1.6bar range with a G1" connection. but the code needs to be reopen as still officially phased out. you'll also have solutions with the PBMN flush (even in G1/2) [Image]
INC254348 | Workgroup | Incident | TR: demande de devis manomètre réf MAR5BX1A31MNA13+D002-26144171 LRS1 // 11126310 | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-11 14:57 | 2022-04-13 12:17 | 0 | 0 |  | Bonjour Hervé, On vient de me confirmer l'obsolescence de ce produit. A priori on avait proposé il y a des années de faire un dernier lot mais il n'y a pas eu de suite. Best, Nicolas
INC254331 | Workgroup | Incident | TR: BAIT - Renova - Y92 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-11 13:48 | 2022-04-11 14:46 | 0 | 0 |  | Dear Colleagues, the diaphragm on the flush connection is directly linked to the size of the connection. so on a 1/2", the diaphragm is very small and cannot be sensitive enough to measure low pressure. the minimum pressure for connection 3G existing is 6bar (B20). According to the datasheet we could go as low as 4bar (to be verified). Best, Nicolas
INC254231 | Workgroup | Incident | resistance of diaphragm seals with nitric acid 98% | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-08 14:15 | 2022-04-08 15:16 | 0 | 0 |  | Hi Sasha, I'm sorry but we don't do "Compatibility statement" as only the user can know completely the process (temperature, pressure, exact chemical composition and external parameters). we can advise thanks to the different compatibility tables we have (below a confirmation that in general, the PTFE is resistant to the Nitric Acid). [Image] If the customer needs to validate by testing it is absolutely not an issue for us. Best, Nicolas
INC254199 | Workgroup | Incident | TR: Demande de prix 040422AG00 (FY21393) | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-08 11:27 | 2022-05-06 15:42 | 1 | 1 |  | CCR2 envoyé a ADV pour réponse a HCO
INC254166 | Workgroup | Incident | TR: BAIT - Technoalpin - SPR CTX | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-08 09:18 | 2022-04-11 16:17 | 0 | 0 |  | Bonjour Nathalie, Apres debrief avec la R&D pour clarifier un peu tout ca, voici nos conclusion. le CTX n'est pas une option car on ne peut pas faire un trou plus grand dans le raccord. Le Y913 est possible en G1/4 avec un trou de 6mm jusqu'à la cellule (la section la plus petite du Danfoss est aussi de 6mm). Le code du raccord sera un X (G1/4 DIN, norme a preciser avec un trou de 6mm). Best, Nicolas
INC254010 | Workgroup | Incident | TR: Average delivery time | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-07 10:18 | 2022-04-07 16:22 | 0 | 0 |  | Celine, on en a 5 en stock et le CCR2 unitaire est de 133 Par contre c'est la piece EQ-13-17-2 Best, Nicolas
INC253945 | Workgroup | Incident | RFQ THERMOWELL | Archive > BHFR Technical Support | Middle |  |  | Closed | Hernandez, Angel [hera] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-06 19:09 | 2022-04-07 09:11 | 0 | 0 |  | Angel, No problem with those codes, Céline will ask for the prices, only thing is the length is done with 5mm steps, I corrected the codes: AGFQ model X 075mm Extention (Eg) 2 Stainless steel 1.4404 5 1500 psi 6 1" 1/2 (DN 40) N 1/2 NPT 2 8 mm R 22 mm .XXX 380 mm Plunger length Pg T Tapered Option Full penetration welding AGFQ model X 075mm Extension (Eg) 2 Stainless steel 1.4404 5 1500 psi 6 1" 1/2 (DN 40) N 1/2 NPT 2 8 mm R 22 mm .XXX 305 mm Plunger length Pg T Tapered Option Full penetration welding AGFQ model X 100 mm Extension (Eg) 2 Stainless steel 1.4404 5 1500 psi 6 1" 1/2 (DN 40) N 1/2 NPT 2 8 mm R 22 mm .XXX 355 mm Plunger length Pg T Tapered Option Full penetration welding Best, Nicolas
INC253939 | Workgroup | Incident | TR: AL E&C - Manomètres différentiel pour service oxygène | Archive > BHFR Technical Support | Middle |  |  | Closed | Courdier, Hervé [hco] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-06 17:30 | 2022-06-08 16:24 | 0 | 1 |  | Clos pour cette demande du certificat du joint, le reste de la procédure d'homologation sera suivie sur le mail: TR: #INC258698 TR: AL E&C - Manomètres différentiel pour service oxygène - Your Support Ticket was closed Best, Nicolas
INC253842 | Workgroup | Incident | TR: RFQ PBMN DT3 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-06 09:12 | 2022-04-13 14:05 | 0 | 1 |  | Ticket closed because no feedback from the customer.
INC253745 | Workgroup | Incident | TR: Remplissage Dow corning | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-05 09:39 | 2022-04-05 09:53 | 0 | 0 |  | Bonjour Céline, Jacques, Dow Chemicals fait de nombreux fluides. de plus, si le client nous impose ses contraintes techniques, il ne peut pas nous imposer de "Marque". Il nous faut les contraintes de montage (type de séparateur), de température et spécifiques (approbations) pour définir un produit stp. Best, Nicolas
INC253684 | Workgroup | Incident | Screwed Thermowell - EVAC BV - WSS10314-04 | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-04 16:09 | 2022-04-13 14:51 | 0 | 1 |  | Hi Rudi, Our supplier can do it, and we received the offer, I transfer to Celine, she'll send you the offer. Best, Nicolas
INC253630 | Workgroup | Incident | TR: devis pour mano type MFE41 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-04 12:27 | 2022-04-05 13:57 | 0 | 0 |  | Hervé, je clôture le ticket, la demande est en cours a la R&D et database. Best, Nicolas
INC253624 | Workgroup | Incident | WEP5 - Datasheets, Information, ... | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-04 12:00 | 2022-04-04 12:14 | 0 | 0 |  | Hi Rudi, Yes the WEP5 is currently being open to the public. This doesn't change anything on the handling with TOTAL as this is a specific agreement. The datasheet will be released soon and I'll send it to you as soon as I have it. Best, Nicolas
INC253623 | Workgroup | Incident | TR: PO#2730 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-04 11:51 | 2022-04-04 14:04 | 0 | 0 |  | Anne, there is nothing blocking technically but it seems the price need to be reviewed before ordering. you can check it with stephane I guess. the CCR2 is 105.5 in 2022 and the IC P2F is 205.4 so I don't see any issue, but maybe there is something I don't see. Best, Nicolas
INC253494 | Workgroup | Incident | AW: Purchase Order for BOURDON - L2022031001 from Linhorn | Archive > BHFR Technical Support | Middle |  |  | Closed |  | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-04-01 12:57 | 2022-04-01 13:58 | 0 | 0 |  | Hello Carla, Don't worry, the 1.4404 on the dial is the material of the tube, not of the case. What Josette offered you is fine considering what we can see on the picture. MVX7-B30.B12 Best, Nicolas
INC253293 | Workgroup | Incident | WG: WIKA 232.50.100 and TG54.100 substitute needed | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-30 15:33 | 2022-03-31 16:10 | 0 | 0 |  | Hi Carlo, as discussed, here are the current limits, and we will work globally on the "cold issue". Ambiant in use: -20...+70C Storage: -40...+70C Best, Nicolas
INC253191 | Workgroup | Incident | TR: MEX5 + ARP-X06N - Wasif | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-30 10:36 | 2022-03-30 14:44 | 0 | 0 |  | Anne, the customer is going on complicated things when it couls be easier. No need for a F/F adaptor if they take a female 1/2NPT welded connector (I would order it through a maintenance webshop or so) there is no gasket needed for NPT connection (the one they show is a crushable washer for Gas connection I think). Best, Nicolas
INC253190 | Workgroup | Incident | TR: BAIT ¦ Manometers with gas and dust approvals | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-30 10:36 | 2022-03-30 15:00 | 0 | 0 |  | Anne, We don't have ATEX for dust atmosphere for the Gauges with contacts. for this type of protection, they can go to the pure switches (RP2Y for example) or the PFMN with a display and contacts. the IP protection is the protection against water and dust projection/immersion but nothing to do with explosive atmosphere. Best, Nicolas
INC253189 | Workgroup | Incident | TR: question about MFT5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-30 10:33 | 2022-03-30 15:05 | 1 | 0 |  | Anne, Oui, techniquement il n'y a aucun soucis pour du psi/bar sur les MFT5. Habituellement je t'aurais dit de juste mettre un digit Sxx/9712 et le prix du 9712 sur le MEX5. Mais vu l'attention qui est portée aux prix en ce moment, je te conseille de faire la demande a Fantinelli même si ca rend les choses plus compliquées. Best, Nicolas
INC253072 | Workgroup | Incident | TR: PRICE COMPLET SET FOR POINTER SETTING | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-29 11:18 | 2022-03-29 11:29 | 0 | 0 |  | Anne, Dans le ticket tu trouvera un document que j'avais fait pour ESP, il y a tous les codes disponnibles (les prix etaient pour ESP). Best, Nicolas
INC252787 | Workgroup | Incident | TR: transmitter of pressure bourdon haenni | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-24 11:54 | 2022-03-24 17:08 | 0 | 0 |  | Anne, Voici le detail de la config mais pour le raccord ca n'aide pas trop et je ne trouve rien d'autre. Il va falloir redefinir la connection process. pour le reste il peut partir sur un PBMH. ED 701 CAPTEUR DE PRESSION PIEZORESISTIF POUR INDUSTRIE R20 RACCORD BRIDE TYP FEDEGARI SIGNAL DE SORTIE 4...20 MA CE CONFORME 4R4 RACCORD ELECTRIQUE CONNECTEUR DIN 43650 PRECISION: ERREUR MAX. 25 MBAR ENTRE 25 ET 125C CALIBREE A 120C POUR PRECISION MAXIMUM PRESSION ABSOLUE A25 0...5 BAR 10 PARTIES MOUILLESS EN INOX LIQUIDE DE TRANSFERT - HUILE SILICONE 9007/0120 TEMPERATUR DU MEDIUM 120 DEGRES C Best, Nicolas
INC252775 | Workgroup | Incident | TR: MEP5 with SS316L case in the old design | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-24 10:15 | 2022-03-25 12:01 | 0 | 0 |  | Hi Kent, for this time, if absolutely necessary, you can go ahead with the old design, but keep in mind that it's most probably the last time we can do it. make sure at the order stage that you have the right version. best, Nicolas
INC252670 | Workgroup | Incident | DRO100/423.111 - IFG de Wit (55235195) | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-23 11:21 | 2022-03-29 12:09 | 0 | 0 |  | Hi Rudi, it seems I can't get an answer from KHW but here is the datasheet (in the ticket) so that you can have the products details and check with the customer for a replacement. best, Nicolas
INC252657 | Workgroup | Incident | TR: DIAPHRAGM SEAL | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-23 10:15 | 2022-03-24 14:12 | 0 | 0 |  | Please create a new ticket if there is more questions coming on this. Best, Nicolas
INC252552 | Workgroup | Incident | TR: MEX with 3 scale | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-22 12:27 | 2022-03-22 13:54 | 0 | 0 |  | Nathalie, oui, pas de problème il y a toutes les informations sur le plan et le code est bon (il faudra peut être ajouter des options de cadran selon la codif). il nous faudra aussi le LOGO client si ils le veulent en JPEG. Best, Nicolas
INC252505 | Workgroup | Incident | TR: Our inquiry / prijsaanvraag 22220734 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-22 09:51 | 2022-03-22 10:05 | 0 | 0 |  | Salut Celine, c'est celui-ci: 11005844 Y91333B26R/2297/2152/0409. Best, Nicolas
INC252455 | Workgroup | Incident | Drever ( 21107358 ) - RFQ J587 - MCX5... temp. +/- 300°C | Archive > BHFR Technical Support | Middle |  |  | Closed | Pickery, Rudi [piru] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-21 14:42 | 2022-03-21 15:28 | 0 | 0 |  | Hi Rudi, there will be no effet with the siphon as anyway, Nitrogen is gazeous over -180C and it'll stay gazeous. a capillary is a good idea, anyway it'll stay gazeous but it'll be good to control the temperature in the instrument (T drops anround 80C/m of capillary). Best, Nicolas
INC252441 | Workgroup | Incident | commande bonbons | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-21 14:03 | 2022-03-22 13:58 | 1 | 0 |  | Notre service n'est plus rattaché au votre pour l'approvisionnement en sucrerie, votre demande est donc rejetée. Nous sommes désolé de ne pas pouvoir plus vous aider Nicolas (Service R&D)
INC252389 | Workgroup | Incident | TR: y91 25MBAR | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-21 10:21 | 2022-03-21 10:27 | 0 | 0 |  | Bonjour Anne, Nous ne pouvons plus faire le 25mbar, le mini est 100mbar. on a rien pour remplacer a cette pression, mais on peut s'en approcher en déclassant la précision, on pourrait proposer un 60mbar (précision a valider). Best, Nicolas
INC252273 | Workgroup | Incident | TR: 21127 / QIP / Request for Quotation / F-SB-0-I-RQ-001 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-18 09:18 | 2022-03-18 10:41 | 0 | 0 |  | Hi Wesley, this mounting is not possible, the minimums for a mounting with a 2" 150 D824 are: 0...4bar range (code 44) min setp: 0.4bar Pmax: 50bar Best, Nicolas
INC252245 | Workgroup | Incident | WG: RFQ: Process Sensors ¦ Met Uys Technologies (Pty)Ltd | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-17 16:12 | 2022-03-22 09:56 | 0 | 0 |  | Hello Carlo, Here are the gauges we can offer as standard (no filling with the connection on the side (it's one or the other): MEX5D63X23/9704/9712/0751 (-100....1250 kpa /NH3 (red)) MEX5D63X24/9704/9712/0751 (-100....1600 kpa/NH3 (red)) MEX5D63X26/9704/9712/0751 (-100....2500 kpa/NH3 (red)) 9750 to be added if LOGO GEA requested. Best, Nicolas
INC252194 | Workgroup | Incident | TR: MMX1 Stainless steel 316 | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-17 11:42 | 2022-03-17 12:05 | 0 | 0 |  | Hi Anne, We don't do MMX1 in standard with 316L case. and considering the triple scale, I'd strongly suggest a MEX2 or 3: Order code: MEX2-610.B22/9713 MEX PP old Code / SAP code: EURO Nominal size: 50 mm 2 21.70 Type of mounting: Case and bezel ring AISI 316L (1.4404), stainless steel Back connection 6 11.50 Process connection: G1/8" 1 5.20 Oil filling: No oil filling 0 0.00 Pressure unit: Bar (EN standard) B 0.00 Pressure range: 0 ... 10 bar 22 0.00 OPTIONS Dial: Triple scale 9713 3.60 Purchasing-price in EURO 42.00 Best, Nicolas
INC252168 | Workgroup | Incident | WG: Richiesta di offerta FATA EPC | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-17 10:03 | 2022-03-17 10:21 | 0 | 0 |  | Hi Carlo, can't you offer our internal solutions like M21 or M23 ? Order code: M23-2P6.B15H/9941 M2 PP old Code / SAP code: EURO Nominal size: 150 mm 3 541.90 Functions of order: Maxi (8001) 1 contact CEI SJ2N 2 143.30 Type of mounting: Case and bezel ring AISI 304 (1.4301), stainless steel Bottom connection with mounting bracket for 2"pipes P 33.20 Process connection: 1/2" NPT 6 0.00 Pressure unit: Bar (EN standard) B 0.00 Pressure range: 0 ... 1 bar 15 0.00 Static pressure: 10 bar H 0.00 OPTIONS Identification: Stainless steel label and wire 9941 2.80 Purchasing-price in EURO 721.20
INC252083 | Workgroup | Incident | TR: Confirm price for differential pressure switch model RDN-8-K-3-221/SEPT | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-16 11:30 | 2022-03-22 12:02 | 1 | 0 |  | Salut Anne, c'est validé avec PPE, le RDN-8-K-3-221/SETP est OK. tu va recevoir le code/prix de Josette. Best, Nicolas
INC252071 | Workgroup | Incident | TR: MG5 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-16 10:30 | 2022-03-17 10:17 | 1 | 0 |  | Hi Thomas, MG5 and D820 can be mounted together, but the mounting depend on the pressure range and the size of the D820/D824 flange connection. a certain number already exists as you can see: [Image] Best, Nicolas
INC252000 | Workgroup | Incident | TR: RFQ 150322 BAUMER | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-15 14:48 | 2022-03-17 09:52 | 1 | 0 |  | Nathalie, Basically, the YTED is a TED6, only the ATEX differs so I would offer the same one without ATEX: [Image] Best, Nicolas
INC251887 | Workgroup | Incident | TR: Différence entre option 0808 BV et 0809 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-14 15:09 | 2022-03-21 11:40 | 0 | 1 |  | Bonjour, voir dans le ticket pour le MTTF du TED. Best, Nicolas
INC251824 | Workgroup | Incident | TR: Déssin 3D Y91-11250481 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-14 09:39 | 2022-03-14 17:12 | 0 | 0 |  | Celine, Désolé, mais la 3D n'existe pas et Pascal ne peut pas s'y mettre tout de suite. Si c'est une urgence, revient vers moi mais on devra évaluer l'urgence avec David si on doit modifier les prios de Pascal. Best, Nicolas
INC251771 | Workgroup | Incident | TR: ER11 pressure celles | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-11 15:09 | 2022-03-11 15:28 | 0 | 0 |  | Hi Zoran, Sorry for the news but we don't do those cells anymore and we don't have stock left. Best, Nicolas
INC251717 | Workgroup | Incident | TR: D824 pricing | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-11 09:06 | 2022-03-11 09:29 | 0 | 0 |  | Nathalie, Je ne donne pas de prix et les codes semblent OK.
INC251658 | Workgroup | Incident | TR: enquiry | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-10 13:54 | 2022-03-10 14:45 | 0 | 0 |  | Nathalie, Ca ressemble très fortement à des ALF de l'inde, je ne peut pas tout voir mais ce n'est pas nous qui le fournissons. Par contre, on peut bien sur remplacer par des MEX5 au besoin mais il faudra que le client vérifie les dimensions pour le montage et nous dise si c'est des raccords au dos, quels raccords etc... Nicolas
INC251534 | Workgroup | Incident | WG: 1921078 ELIQUO STULZ GmbH - Bestellung EB058938-->Sascha angeschrieben wegen Alternative/Klärung mit BHFR | Archive > BHFR Technical Support | Middle |  |  | Closed | Wendt, Sascha [wesa] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-09 15:09 | 2022-03-14 15:14 | 0 | 0 |  | Dear Sascha, I just received the confirmation from R&D and the option 0052 on MCX is no longer possible. the usage was too low and we had to stop. but there is other option of custom dial to show the target pressure or dangerous pressure level etc... what do the customer want to show with the red pointer ? Best, Nicolas
INC251456 | Workgroup | Incident | TR: TSS probe length | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-09 08:33 | 2022-03-10 09:28 | 1 | 0 |  | Nathalie, Yes, TSS is available in custom length, but for this one I don't suggest it, 5mm is not usually important on 100mm. Please consider using a sliding connection or a 100mm before created something special that is not really needed (even by the customer). Nicolas
INC251404 | Workgroup | Incident | ESSAI 1 | Archive > BHFR Technical Support | Middle |  |  | Closed | Deschamps, Josette [jau] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-08 15:00 | 2022-03-08 15:17 | 0 | 0 |  | dsfze
INC251368 | Workgroup | Incident | TR: Tri-Clamp Gauge | Archive > BHFR Technical Support | Middle |  |  | Closed | Duclos, Anne [anma] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-08 10:57 | 2022-03-08 14:11 | 0 | 0 |  | Hello Daniel, There is the DP100 with a built in clamp connection (see datasheet enclosed) but the specificity of this gauge is the transmission between the clamp and the gauge is a mechanical rod (not possibility of pollution of the process) otherwise, the best solution is the MEX+DANC including in term of pricing. Best, Nicolas
INC251351 | Workgroup | Incident | TR: Article 10346564 | Archive > BHFR Technical Support | Middle |  |  | Closed | Billault, Nathalie [nbil] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-08 09:39 | 2022-03-08 14:37 | 1 | 0 |  | Nathalie, Vu avec Sascha, on va proposer le montage ci-dessous et voir avec le client si ca convient, tu peux lui ajouter le prix et lui envoyer stp ? MEP5-131.B76(New design)/9710_9700 + AKPL2F5303 + DANC13.2L2A.01 DN40 is not possible at this range, the bare minimum is a DN50 (DIN) (or change for a range -1...5bar with DN40 I don't know who requested the 9710 and 9700 but maybe it's not necessary, I don't see the gauge on the picture so I based it on what was requested. the accuracy of the gauge should b respected according to the datasheet but as we are close to the limits, it's better to validate together with R&D once the complete mounting is validated. best
INC251129 | Workgroup | Incident | WG: Richiesta Scheda di Sicurezza Aggiornata Italiano | Archive > BHFR Technical Support | Middle |  |  | Closed | Uras, Carlo [urc] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-03-04 11:39 | 2022-03-04 15:55 | 0 | 0 |  | Dear Carlo, we talk later about what you said. but here is already the MSDS (material safety datasheet) for the BH3 that contains all information about the qualifications and norms about this product.
INC250701 | Workgroup | Incident | TR: Demande offre n/ref D220284 | Archive > BHFR Technical Support | Middle |  |  | Closed | Martin, Celine [mace] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-28 17:21 | 2022-03-01 11:49 | 0 | 0 |  | DRCE Prix
INC250596 | Workgroup | Incident | Essai David | Archive > BHFR Technical Support | Middle |  |  | Closed | Lecomte, David [leda] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Lecomte, David [leda] | 2022-02-28 10:15 | 2022-02-28 10:37 | 0 | 0 |  | reçu et clos
INC250220 | Workgroup | Incident | Test107 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-22 14:18 | 2022-02-23 10:08 | 0 | 0 |  | received, accepted and closed
INC250217 | Workgroup | Incident | Test106 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-22 14:03 | 2022-02-23 10:08 | 0 | 0 |  | received, accepted and closed
INC250213 | Workgroup | Incident | Test105 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-22 13:48 | 2022-02-23 09:06 | 1 | 0 |  | accepted and closed
INC250083 | Workgroup | Incident | Test104 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-21 11:09 | 2022-02-23 10:07 | 1 | 0 |  | received, accepted and closed
INC250082 | Workgroup | Incident | test102 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-21 11:03 | 2022-02-23 10:07 | 1 | 0 |  | received, accepted and closed
INC250080 | Workgroup | Incident | Test101 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-21 10:54 | 2022-02-23 10:07 | 1 | 0 |  | received, accepted and closed
INC250076 | Workgroup | Incident | Test100 | Archive > BHFR Technical Support | Middle |  |  | Closed | Bhend, Nadia [bna] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-21 10:39 | 2022-02-23 10:07 | 1 | 0 |  | received, accepted and closed
INC248354 | Workgroup | Incident | Ceci est un essai N°2 | Archive > BHFR Technical Support | Middle |  |  | Closed | Lecomte, David [leda] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-02-02 09:12 | 2022-02-02 09:27 | 0 | 0 |  | received, accepted and closed
INC247729 | Workgroup | Incident | Ceci est un Essai | Archive > BHFR Technical Support | Middle |  |  | Closed | Lecomte, David [leda] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Lecomte, David [leda] | 2022-01-25 10:36 | 2022-01-26 09:07 | 0 | 0 |  | Bien reçu, accepté et fermé.
INC247595 | Workgroup | Incident | Test | Archive > BHFR Technical Support | Middle |  |  | Closed | Zuberbühler, Lucca [zlu] | Birrien, Nicolas [bin] | BG - Baumer Group |  | Birrien, Nicolas [bin] | 2022-01-21 17:15 | 2022-03-01 13:57 | 0 | 0 |  | well received Nicolas Birrien